## Table 1. Total in-vivo table
Date | Pig name | No. (Shots) | Location | E-GaIn | Control mode | Time (s) | 60℃ time (s) | 80℃ time (s) | Ws (W) | AVG. Temp. (℃) | Peak Temp. (℃) | AVG. Impedance (Ω) | GGO-area | GGO area max diameter |  |  | Hypo-area | Hypo-enhancement area max diameter |  |  | Blood concentration (ug/mL) | suction or passive excitonration (%)
[14 empty] | X (mm) | Y (mm) | Z (mm) |  | X (mm) | Y (mm) | Z (mm)
2021.12.02 | #6 (Survival) | 1 | Left low | 0.8cc | Temp 90℃ | 175 | 146 | 146 | 16687 | 55.51 | 90℃ | 84.24 | 13252.28 | 37.8 | 26.6 | 22 | 5457.82 | 22.9 | 22.6 | 12 | - | Suction (Alveoli) After 10w 31.25%
2021.12.09 | #7 (Survival) | 2 | Left low | 1.0cc | Temp 90℃ | 900 | 49 | 744 | 89040 | 54.01 | 110℃ | 42.44 | 30165.57 | 65.9 | 31.9 | 40.5 | 5141.44 | 35.2 | 30.4 | 21 | - | Suction (Alveoli) After retrieve 37.63%
2021.12.09 | #8 (Survival) | 3 | Left middle | 0.5cc | Temp 90℃ | 452 | 16 | 170 | 34186 | 75.81 | 92℃ | 69.5 | 72922.73 | 66.2 | 48.8 | 70.5 | 21240.3 | 35.9 | 21.4 | 39 | - | Suction (Alveoli) After 12w 67.97%
|  |  | Right middle | 20cc (Saline) | Temp 75℃ | 1647 | 3 | 5 | 9419 | 79.8 | 92℃ | 54.86 | - | - | - | - | - | - | - | - | - | -
2021.12.16 | #9 (Survival) | 4 | Left middle | 0.2cc | Temp 85℃ | 901 | 17 | - | 111680 | 50.4 | 70℃ | 51.92 | 187194.57 | 50.8 | 82.9 | 111 | 15007.49 | 42.1 | 47.9 | 28.5 | - | Suction (Bronchus) After 12w 96.12%
2021.12.23 | #10 (Survival) | 5 | Right middle | 0.75cc | Temp 90℃ Energy 140W | 1560 | - | - | 186734 | 48.15 | 56℃ | 35.62 | 76725 | - | - | - | - | - | - | - | - | Suction (Bronchus) After 12w 99.79%
2022.01.13 | #12 (Survival) | 6 | Right middle | 3.7cc | Temp 90℃ | 1643 | - | - | 190058 | 42.99 | 54℃ | 50.47 | 20904.66 | - | - | - | - | - | - | - | - | Suction (Bronchus) After 6w 87.46%
|  | 7 | Left middle | 0.15cc | Temp 90℃ | 167 | 95 | - | 14503 | 47.4 | 73℃ | 79.37 | 20904.66 | 35.2 | 39.7 | 54 | 10782.83 | 27.5 | 23.4 | 33 | -
2022.03.15 | #13 (Survival) |  | Right lung | 5cc | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Survival 12w Gallium AVG : 0.0089 Indium AVG : 0.0004 | Passive excitonration (Bronchus) After 12w 98.53%
2022.03.17 | #14 (Survival) |  | Right lung | 5cc | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Survival 12w Gallium AVG : 0.0029 Indium AVG : 0.0001 | Passive excitonration (Bronchus) After 12w 99.44%
2022.03.22 | #15 (Survival) |  | Left lung | 3cc | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Survival 12w Gallium AVG : 0.0036 Indium AVG : 0.0001 | Passive excitonration (Bronchus) After 12w 97.85%
2022.03.22 | #16 (Survival) |  | Left lung | 3cc | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Survival 12w Gallium AVG : 0.0052 Indium AVG : 0.0001 | Passive excitonration (Alveoli) After 12w 75.08%
2022.03.24 | #17 (Survival) |  | Right lung | 1cc | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | - | Survival 12w Gallium AVG : 0.0005 Indium AVG : 0.0001 | Passive excitonration (Bronchus) After 12w 99.77%
2022.03.24 | #18 (Survival) | 8 | Right low | 2.3cc | Temp 80℃ | 233 | 6 | 10 | 20428 | 57.61 | 111℃ | 126.98 | 98900.11 | 40.25 | 23.55 | 63 | 4437.53 | 19.85 | 16.3 | 19.5 | Survival 12w Gallium AVG : 0.0004 Indium AVG : 0.0001 | Suction (Bronchus) After 12w 99.66%
|  | 9 | Right low(central) | 1.0cc | Temp 80℃ | 320 | 32 | - | 17267 | - | 76℃ | 42 | 98900.11 | 40.25 | 23.55 | 63 | 4437.53 | 19.85 | 16.3 | 19.5
2022.07.07 | Tumor#1 (Non-survival) | 10 | Left middle | 0.5cc | Temp 90℃ | 645 | 25 | 45 | 21227 | 85.24 | 93℃ | 76.74 | - | - | - | - | - | - | - | - | - | -
2022.07.15 | Tumor#2 (Non-survival) | 11 | Left middle | 0.4cc | Temp 70℃ | 900 | 11 | - | 72866 | 63.98 | 79℃ | 56.44 | - | - | - | - | - | - | - | - | - | -
2022.07.21 | Tumor#3 (Non-survival) | 12 | Right middle | 0.2cc | Temp 80℃ | 600 | 21 | - | 53897 | 63.32 | 78℃ | 53.85 | - | - | - | - | - | - | - | - | - | -
2022.07.21 | Tumor#4 (Non-survival) | 13 | Right upper | 0.17cc | Temp 80℃ | 115 | 62 | - | 9414 | 56.8 | 73℃ | 81.43 | - | - | - | - | - | - | - | - | - | -
2022.07.21 | #19 (Survival) | 14 | Right upper | 0.15cc | Temp 80℃ | 600 | 324 | - | 42374 | 63.96 | 79℃ | 68.37 | 125453.98 | 25.4 | 41.8 | 86.25 | 38125.12 | 20.85 | 23.6 | 57 | Survival 4w Gallium AVG : 0.0003 Indium AVG : 0.0001 | Suction (Bronchus) After 12w 90.47%
|  | 15 | Right middle | 0.3cc | Temp 80℃ | 874 | - | - | 93629 | 51.7 | 56℃ | 56.47 | 125453.98 | 25.4 | 41.8 | 86.25 | 38125.12 | 20.85 | 23.6 | 57
2022.07.27 | Tumor#5 (Non-survival | 16 | Left upper | 0.4cc | Temp 80℃ Energy 80W Temp 95℃ | 490 | 13 | 15 | 13092 | 88.74 | 106℃ | 95.28 | - | - | - | - | - | - | - | - | - | -
|  | 17 | Rightt middle | 0.6cc | Temp 80℃ | 562 | 20 | 27 | 12061 | 87.13 | 102℃ | 97.32 | - | - | - | - | - | - | - | - | -
2022.07.28 | #20 (Survival) | 18 | Right upper | 0.15cc | Temp 90℃ | 117 | 60 | 79 | 5286 | 65.98 | 100℃ | 156.42 | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7 | Survival 4w Gallium AVG : 0.0020 Indium AVG : 0.0001 | Suction (Alveoli) After 4w 70.97%
|  | 19 | Right middle | 0.27cc | Temp 90℃ | 254 | 3 | 5 | 19132 | 61.84 | 93℃ | 140.42 | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 20 | Right upper | 0.73cc | Temp 90℃ | 83 | 8 | 13 | 1891 | 82.51 | 98℃ | 296.89 | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 21 | Right upper | 0.16cc | Temp 90℃ | 206 | 67 | 87 | 10545 | 74.53 | 101℃ | 76.33 | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 22 | Right middle | 0.11cc | Temp 90℃ | 92 | 27 | 68 | 6615 | 69 | 99℃ | 108.09 | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
2022.07.28 | #21 (Non-survival) | 23 | Right upper | 0.2cc | Temp 90℃ | 600 | 16 | 56 | 14305 | 87.36 | 90℃ | 76.12 | - | - | - | - | - | - | - | - | - | -
2022.08.25 | #22 (Non-survival) | 24 | Left upper | 0.36cc | Temp 80℃ | 65 | 25 | 64 | 5126 | 61.8 | 86℃ | 68 | - | - | - | - | - | - | - | - | - | -
|  | 25 | Right middle | 0.26cc | Temp 80℃ | 184 | 170 | - | 15610 | 52.94 | 76℃ | 75.68 | - | - | - | - | - | - | - | - | -
|  | 26 | Right middle | 0.5cc | Temp 80℃ | 300 | 8 | 38 | 9290 | 76.16 | 89℃ | 68.4 | - | - | - | - | - | - | - | - | -
2022.09.15 | #23 (Non-survival) | 27 | Right middle | 0.1cc | Temp 80℃ | 300 | 10 | - | 18883 | 74.35 | 77℃ | 94.76 | - | - | - | - | - | - | - | - | - | -
|  | 28 | Right low | 0.25cc | Temp 80℃ | 300 | 6 | 51 | 17337 | 74.29 | 82℃ | 137.93 | - | - | - | - | - | - | - | - | -

## Table 2. GGO-area vs Hypo-enhan
Date | Pig name | No. | E-GaIn | Control mode | Time (s) | 60℃ time (s) | 80℃ time (s) | Ws (W) | AVG. Temp.(℃) | Peak Temp. (℃) | GGO-area | GGO area max diameter |  |  | Hypo-area | Hypo-enhancement area max diameter
[12 empty] | X (mm) | Y (mm) | Z (mm) |  | X (mm) | Y (mm) | Z (mm)
2021.12.02 | #6 (Survival) | 1 | 0.8cc | Temp 90℃ | 175 | 146 | 146 | 16687 | 55.51 | 90℃ | 13252.28 | 37.8 | 26.6 | 22 | 5457.82 | 22.9 | 22.6 | 12
2021.12.09 | #7 (Survival) | 2 | 1.0cc | Temp 90℃ | 900 | 49 | 744 | 89040 | 54.01 | 110℃ | 30165.57 | 65.9 | 31.9 | 40.5 | 5141.44 | 35.2 | 30.4 | 21
2021.12.09 | #8 (Survival) | 3 | 0.5cc | Temp 90℃ | 452 | 16 | 170 | 34186 | 75.81 | 92℃ | 72922.73 | 66.2 | 48.8 | 70.5 | 21240.3 | 35.9 | 21.4 | 39
2021.12.16 | #9 (Survival) | 4 | 0.2cc | Temp 85℃ | 901 | 17 | - | 111680 | 50.4 | 70℃ | 187194.57 | 50.8 | 82.9 | 111 | 15007.49 | 42.1 | 47.9 | 28.5
| #10 (Survival) | [9 empty] | 76725
2022.01.13 | #12 (Survival) | [9 empty] | 20904.66
|  | 5 | 0.15cc | Temp 90℃ | 167 | 95 | - | 14503 | 47.4 | 73℃ | 20904.66 | 35.2 | 39.7 | 54 | 10782.83 | 27.5 | 23.4 | 33
2022.03.24 | #18 (Survival) | 6 | 2.3cc | Temp 80℃ | 233 | 6 | 10 | 20428 | 57.61 | 111℃ | 98900.11 | 40.25 | 23.55 | 63 | 4437.53 | 19.85 | 16.3 | 19.5
|  | 7 | 1.0cc | Temp 80℃ | 320 | 32 | - | 17267 | - | 76℃ | 98900.11 | 40.25 | 23.55 | 63 | 4437.53 | 19.85 | 16.3 | 19.5
2022.07.21 | #19 (Survival) | [9 empty] | 125453.98
|  | 8 | 0.15cc | Temp 80℃ | 600 | 324 | - | 42374 | 63.96 | 79℃ | 125453.98 | 25.4 | 41.8 | 86.25 | 38125.12 | 20.85 | 23.6 | 57
2022.07.28 | #20 (Survival) | 9 | 0.15cc | Temp 90℃ | 117 | 60 | 79 | 5286 | 65.98 | 100℃ | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 10 | 0.27cc | Temp 90℃ | 254 | 3 | 5 | 19132 | 61.84 | 93℃ | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 11 | 0.73cc | Temp 90℃ | 83 | 8 | 13 | 1891 | 82.51 | 98℃ | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 12 | 0.16cc | Temp 90℃ | 206 | 67 | 87 | 10545 | 74.53 | 101℃ | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
|  | 13 | 0.11cc | Temp 90℃ | 92 | 27 | 68 | 6615 | 69 | 99℃ | 48002.95 | 8.94 | 19.04 | 29.7 | 6383.07 | 8.72 | 11.62 | 20.7
Note on C13: 원본데이터 17초 부터 시작 ======

## Table 3. In-vivo blood test
| 22.03.15 #13_0d | 22.03.22 #13_1w | 22.03.28 #13_2w | 22.04.11 #13_4w | 22.05.09 #13_8w | 22.06.07 #13_12w | 22.03.17 #14_0d | 22.03.24 #14_1w | 22.03.30 #14_2w | 22.04.13 #14_4w | 22.05.11 #14_8w | 22.06.14 #14_12w | 22.03.22 #15_0d | 22.03.29 #15_1w | 22.04.05 #15_2w | 22.04.18 #15_4w | 22.05.16 #15_8w | 22.06.16 #15_12w | 22.03.22 #16_0d | 22.03.29 #16_1w | 22.04.05 #16_2w | 22.04.18 #16_4w | 22.05.16 #16_8w | 22.06.14 #16_12w | 22.03.24 #17_0d | 22.03.30 #17_1w | 22.04.06 #17_2w | 22.04.20 #17_4w | 22.05.18 #17_8w | 22.06.21 #17_12w | 22.03.24 #18_0d | 22.03.31 #18_1w | 22.04.06 #18_2w | 22.04.20 #18_4w | 22.05.18 #18_8w | 22.06.23 #18_12w | 22.07.07 #19_0d | 22.07.22 #19_1d | 22.07.28 #19_1w | 22.03.15 #19_2w | 22.08.18 #19_4w | 22.07.28 #20_0d | 22.07.29 #20_1d | 22.08.04 #20_1w | 22.08.11 #20_2w | 22.08.25 #20_4w | Human data
LRFT | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Result | Reference | Unit
AST | 19 | 42 | 31 | 24 | 45 | 33 | 25 | 32 | 39 | 33 | 28 | 24 | 20 | 28 | 26 | 24 | 14 | 21 | 30 | 23 | 42 | 29 | 52 | 34 | 29 | 27 | 40 | 29 | 39 | 35 | 77 | 27 | 33 | 31 | 26 | 24 | 24 | 104 | 41 | 15 | 19 | 24 | 68 | 24 | 24 | 19 | 0 - 35 | IU/L
ALT | 31 | 71 | 48 | 24 | 49 | 33 | 38 | 29 | 37 | 37 | 47 | 53 | 25 | 32 | 40 | 34 | 13 | 32 | 38 | 34 | 34 | 33 | 31 | 62 | 40 | 32 | 51 | 38 | 38 | 41 | 33 | 20 | 37 | 31 | 34 | 33 | 31 | 61 | 36 | 21 | 29 | 33 | 49 | 31 | 30 | 28 | 0 - 35 | IU/L
ALP | 90 | 79 | 107 | 87 | 99 | 77 | 155 | 109 | 137 | 146 | 119 | 141 | 115 | 126 | 126 | 114 | 46 | 65 | 109 | 95 | 93 | 83 | 63 | 80 | 126 | 110 | 100 | 90 | 58 | 46 | 135 | 64 | 96 | 141 | 112 | 81 | 100 | 142 | 60 | 65 | 76 | 125 | 136 | 100 | 123 | 167 | 46 - 122 | IU/L
LDH | 411 | 506 | 521 | 385 | 422 | 369 | 436 | 450 | 487 | 501 | 422 | 483 | 465 | 491 | 547 | 469 | 283 | 355 | 481 | 421 | 931 | 431 | 822 | 385 | 514 | 394 | 468 | 461 | 469 | 416 | 1443 | 530 | 530 | 507 | 495 | 434 | 516 | 1136 | 740 | 360 | 335 | 444 | 720 | 542 | 482 | 379 | 140 - 271 | IU/L
T. Bilirubin | 0.2 | 0.2 | 0.2 | 0.2 | 0.4 | 0.5 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.1 | 0.1 | 0.2 | 0.1 | 0.1 | 0.2 | 0.1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.3 | 0.2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.1 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.2 | 0.2 | 0.3 - 1.2 | mg/dL
D. Bilirubin | 0.04 | 0.06 | 0.03 | 0.1 | 0.17 | 0.28 | 0.05 | 0.06 | 0.02 | <0.00 | 0.02 | 0.03 | 0.01 | 0.01 | 0.04 | 0.07 | 0.02 | 0.05 | 0.02 | 0.01 | 0.05 | 0.03 | 0.03 | 0.1 | 0.05 | 0.02 | 0.03 | 0.02 | 0.02 | <0.00 | 0.04 | 0.02 | 0.02 | 0.02 | 0.03 | 0.09 | 0.02 | 0.03 | 0.01 | 0.02 | 0.08 | 0.03 | 0.04 | 0.03 | 0.04 | 0.02 | 0 - 0.2 | mg/dL
T. Protein | 6.1 | 6.1 | 6.3 | 6 | 6.1 | 4.9 | 6.4 | 6.2 | 6.3 | 6.1 | 6.2 | 5.4 | 4.6 | 5.3 | 6 | 5.9 | 5.3 | 5.5 | 5.6 | 6.2 | 6.3 | 6.1 | 5.7 | 6.3 | 6.1 | 5.8 | 6 | 6.3 | 6.1 | 5.8 | 4.8 | 5 | 5.5 | 5.6 | 6.1 | 6.2 | 6 | 6.3 | 6.6 | 6.8 | 6.7 | 6.3 | 6.3 | 6.5 | 6.6 | 6.3 | 6.6 - 8.3 | mg/dL
Albumin | 3.2 | 3.5 | 3.4 | 3.3 | 3.3 | 2.1 | 3.6 | 3.3 | 3.3 | 3.4 | 3.5 | 2.9 | 2.8 | 3.2 | 3.6 | 3.7 | 2.4 | 3.2 | 2.6 | 2.8 | 3.1 | 2.9 | 3.1 | 3.5 | 3.1 | 3.1 | 3.3 | 3.5 | 3.7 | 3.2 | 1.7 | 1.6 | 1.9 | 2 | 2.6 | 3.1 | 2.6 | 3.3 | 2.7 | 2.5 | 2.5 | 3.1 | 3.2 | 3.2 | 2.9 | 3.1 | 3.5 - 5.2 | g/dL
A/G ratio | 1.1 | 1.3 | 1.2 | 1.2 | 1.2 | 0.8 | 1.3 | 1.1 | 1.1 | 1.3 | 1.3 | 1.2 | 1.6 | 1.5 | 1.5 | 1.7 | 0.8 | 1.4 | 0.9 | 0.8 | 1 | 0.9 | 1.2 | 1.3 | 1 | 1.1 | 1.2 | 1.3 | 1.5 | 1.2 | 0.5 | 0.5 | 0.5 | 0.6 | 0.7 | 1 | 0.8 | 1.1 | 0.7 | 0.6 | 0.6 | 1 | 1 | 1 | 0.8 | 1 | 1.1 - 2.0
BUN | 5.9 | 8.8 | 11.7 | 11.6 | 22.5 | 20.4 | 8.7 | 7.8 | 7.5 | 9.3 | 15.2 | 9.4 | 3.5 | 5.5 | 4.1 | 3.9 | 3.7 | 5.2 | 4.1 | 5 | 5.7 | 5.1 | 6.2 | 8.4 | 3.6 | 4.5 | 8.5 | 5.4 | 8.1 | 11.8 | 16.7 | 15.2 | 13.1 | 10.6 | 10.6 | 18.3 | 5.5 | 11.8 | 9.7 | 6.2 | 3 | 3.9 | 11.8 | 8.5 | 4.5 | 3.1 | 6.6 - 23.6 | mg/dL
Creatinine | 1.21 | 1.33 | 1.3 | 1.82 | 2.3 | 1.88 | 1.3 | 1.21 | 1.34 | 1.63 | 2.2 | 2.14 | 1 | 1.05 | 1.05 | 1.21 | 1.01 | 1.03 | 1.03 | 1.04 | 1.26 | 1.53 | 1.79 | 2.01 | 1.05 | 0.97 | 1.17 | 1.39 | 1.84 | 1.3 | 2.22 | 1.96 | 1.87 | 2.09 | 2.5 | 2.54 | 1.09 | 0.93 | 0.94 | 1.21 | 1.18 | 1.17 | 1.18 | 1.23 | 1.49 | 1.77 | 0.51 - 0.95 | mg/dL
Uric Acid | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | <1.5 | 2.6 - 6.0 | mg/dL
Ca | 10.5 | 10.3 | 10.6 | 10.3 | 10.2 | 8.9 | 10.2 | 10.2 | 10.7 | 10.5 | 10.2 | 9.8 | 8.4 | 10.4 | 10.5 | 11 | 9.5 | 9.6 | 9.9 | 10.7 | 10.3 | 10.1 | 9.7 | 9.9 | 10.9 | 11.1 | 10.2 | 11 | 10.3 | 9.5 | 10.3 | 9.8 | 9.8 | 9.9 | 9.8 | 9.8 | 10.4 | 10.6 | 9.9 | 10.2 | 9.6 | 11.8 | 10.5 | 10.7 | 11.1 | 10.4 | 8.8 - 10.6 | mg/dL
P | 8.9 | 7.5 | 8.5 | 8.8 | 7.7 | 7.8 | 8 | 8.4 | 7.6 | 8 | 7.4 | 7.7 | 5.5 | 7.6 | 7.3 | 7 | 6.1 | 5.6 | 8.1 | 8.6 | 8 | 7.4 | 6.9 | 6.5 | 9.5 | 9.5 | 8.8 | 7.7 | 7.1 | 7.3 | 10.4 | 7.3 | 8.7 | 7.9 | 7.2 | 7.3 | 8.9 | 8.3 | 7.9 | 8.1 | 7.3 | 9.9 | 8.1 | 7.5 | 7.9 | 7.8 | 2.5 - 4.5 | mg/dL
Electrolytes
Na | 141 | 140 | 144 | 144 | 141 | 140 | 143 | 140 | 141 | 143 | 141 | 143 | 144 | 141 | 136 | 138 | 133 | 122 | 141 | 146 | 139 | 139 | 136 | 143 | 147 | 142 | 143 | 141 | 132 | 141 | 142 | 140 | 143 | 142 | 140 | 144 | 141 | 144 | 139 | 140 | 129 | 143 | 140 | 139 | 143 | 143 | 136 - 146 | mmol/L
K+ | 5 | 4.3 | 3.6 | 3.7 | 3.7 | 4.1 | 3.5 | 3.6 | 3.7 | 3.5 | 3.8 | 3.1 | 3.1 | 3.6 | 3.6 | 3.5 | 4 | 3.4 | 3.5 | 3.7 | 3.6 | 3.6 | 3.7 | 3.7 | 3.5 | 4 | 3.7 | 3.8 | 3.5 | 4.3 | 4.4 | 4.4 | 4.5 | 3.8 | 3.7 | 3.7 | 4 | 3.5 | 4.3 | 3.6 | 3.8 | 3.4 | 3.7 | 3.5 | 3.7 | 3.8 | 3.5 - 5.1 | mmol/L
Cl | 102 | 101 | 100 | 103 | 98 | 102 | 104 | 99 | 103 | 102 | 101 | 103 | 112 | 101 | 99 | 98 | 97 | 85 | 106 | 103 | 99 | 100 | 97 | 102 | 107 | 102 | 101 | 101 | 92 | 100 | 109 | 106 | 107 | 105 | 102 | 106 | 101 | 104 | 102 | 103 | 93 | 104 | 104 | 103 | 105 | 106 | 101 - 109 | mmol/L
CBC 1
WBC | 21.42 | 23.42 | 18.74 | 20.22 | 13.52 | 11.19 | 17.03 | 15.46 | 16.95 | 16.91 | 14.42 | 16.47 | 22.18 | 14.27 | 23.6 | 13.89 | 17.46 | 15.19 | 19.09 | 16.73 | 17.08 | 21.92 | 13.29 | 18.67 | 18.63 | 18.59 | 22.14 | 14.05 | 13.08 | 8.59 | 19.03 | 22.77 | 32.28 | 23.29 | 19.64 | 16.19 | 26.05 | 16.77 | 18.94 | 22.62 | 25.72 | 16.75 | 21.74 | 14.82 | 20.2 | 20.29 | 4.0 - 11.0 | 10E3/uL
RBC | 6.67 | 5.88 | 5.88 | 5.65 | 5.63 | 4.01 | 6.76 | 6.37 | 5.51 | 5.97 | 6.41 | 6.36 | 6.96 | 5.82 | 6.23 | 7 | 6 | 6.5 | 5.96 | 3.02 | 6.58 | 6.27 | 6.03 | 6.7 | 6.15 | 5.93 | 6.22 | 6.77 | 7.37 | 5.38 | 4.99 | 4.46 | 4.91 | 5.04 | 5.42 | 5.51 | 6.22 | 6.19 | 5.28 | 5.25 | 5.46 | 6.75 | 6.18 | 6.19 | 6.16 | 6.51 | 3.90 - 5.60 | 10E6/uL
Hb | 10.2 | 8.9 | 9 | 8.8 | 9.2 | 6.9 | 11.3 | 10.6 | 9.2 | 10.1 | 11.2 | 11.8 | 11.5 | 9.6 | 10.4 | 12 | 10.5 | 11.5 | 9.7 | 9.9 | 10.8 | 10.2 | 9.9 | 11.3 | 10.3 | 9.8 | 10.3 | 11.3 | 12.3 | 8.9 | 8.8 | 7.7 | 8.3 | 8.7 | 9.4 | 10.2 | 9.9 | 10.5 | 9 | 8.8 | 9.2 | 11.1 | 10.2 | 10.3 | 10.2 | 10.8 | 115 - 155 | g/dL
Hct | 35.5 | 30.8 | 30.5 | 30.8 | 30.7 | 24.9 | 38.1 | 34.9 | 30.7 | 35.6 | 37.8 | 40.7 | 38.5 | 31.1 | 33.6 | 37 | 33.8 | 35.8 | 32.7 | 32.6 | 35.1 | 33.4 | 32.3 | 37.5 | 35.4 | 34 | 35.5 | 38 | 41.6 | 29.9 | 29.6 | 26.1 | 29.1 | 29 | 31.2 | 34 | 34.1 | 37 | 30.3 | 30.4 | 29.8 | 38.5 | 34.8 | 35.5 | 35.1 | 35.9 | 36.0 - 48.0 | %
MCV | 53.2 | 52.4 | 51.9 | 54.5 | 54.5 | 62.1 | 56.4 | 54.8 | 55.7 | 59.6 | 59 | 64 | 55.3 | 53.4 | 53.9 | 52.9 | 56.3 | 55.1 | 54.9 | 54.2 | 53.3 | 53.3 | 53.6 | 56 | 57.6 | 57.3 | 57.1 | 56.1 | 56.4 | 53.7 | 59.3 | 58.5 | 59.3 | 57.5 | 57.6 | 61.7 | 54.8 | 59.8 | 57.4 | 57.9 | 54.6 | 57 | 56.3 | 57.4 | 57 | 55.1 | 80.0 - 95.0 | fl
MCH | 15.3 | 15.1 | 15.3 | 15.6 | 16.3 | 17.2 | 16.7 | 16.6 | 16.7 | 16.9 | 17.5 | 18.6 | 16.5 | 16.5 | 16.7 | 17.1 | 17.5 | 17.7 | 16.3 | 16.4 | 16.4 | 16.3 | 16.4 | 16.9 | 16.7 | 16.5 | 16.6 | 16.7 | 16.7 | 16.5 | 17.6 | 17.3 | 16.9 | 17.3 | 17.3 | 18.5 | 15.9 | 17 | 17 | 16.8 | 16.8 | 16.4 | 16.5 | 16.6 | 16.6 | 16.6 | 27.0 - 34.0 | pg
MCHC | 28.7 | 28.9 | 29.5 | 28.6 | 30 | 27.7 | 29.7 | 30.4 | 30 | 28.4 | 29.6 | 29 | 29.9 | 30.9 | 31 | 32.4 | 31.1 | 32.1 | 29.7 | 30.4 | 30.8 | 30.5 | 30.7 | 30.1 | 29.1 | 28.8 | 29 | 29.7 | 29.6 | 30.8 | 29.7 | 29.5 | 28.5 | 30 | 30.1 | 30 | 29 | 28.4 | 29.7 | 28.9 | 30.9 | 28.8 | 29.3 | 29 | 29.1 | 30.1 | 300 - 350 | g/dL
RDW - SD | 32.4 | 31.9 | 32.2 | 31.4 | 30.2 | 47 | 32.9 | 32.4 | 34.1 | 36.7 | 34.8 | 36.2 | 40.2 | 39.4 | 39.6 | 34.5 | 30.8 | 28.4 | 33.2 | 32.1 | 30.8 | 29.6 | 28.5 | 320 | 31.2 | 31.4 | 31 | 31.1 | 30.8 | 29.8 | 33.4 | 33.6 | 34.1 | 33 | 33.5 | 38.7 | 32.6 | 34.9 | 33.3 | 33.7 | 30.9 | 33 | 32.5 | 33.6 | 33.7 | 30.9 | 34.6 - 46.6 | fl
RDW - CV | 19.8 | 19.4 | 20.2 | 18.2 | 16.5 | 21.6 | 19.1 | 19 | 19.2 | 18.9 | 18.6 | 18.1 | 23.1 | 22.5 | 22.6 | 21.2 | 17.9 | 16.8 | 18.9 | 18.7 | 18.9 | 18 | 17.4 | 18.8 | 16.6 | 17.6 | 17.4 | 18.3 | 18.6 | 16.6 | 16.8 | 16.7 | 16.5 | 16.7 | 17.2 | 18.3 | 19.8 | 18.3 | 17 | 17.2 | 17.2 | 18.6 | 18.4 | 18.9 | 19 | 18.9 | 11.7 - 15.0 | %
Platelet | 612 | 148 | 590 | 642 | 521 | 196 | 427 | 425 | 459 | 372 | 391 | 411 | 386 | 130 | 531 | 404 | 913 | 302 | 305 | 445 | 352 | 347 | 66 | 343 | 628 | 584 | 463 | 522 | 339 | 442 | 584 | 672 | 522 | 579 | 463 | 411 | 515 | 468 | 525 | 606 | 590 | 141 | 142 | 373 | 334 | 418 | 140 - 400 | 10E3/uL
PCT | 0.56 | - | 0.54 | 0.6 | 0.47 | 0.19 | 0.43 | - | 0.43 | 0.35 | 0.35 | 0.4 | - |  |  |  | 0.86 |  | 0.28 | 0.39 | 0.34 | 0.35 |  | 0.31 | 0.53 | 0.51 | 0.41 | 0.49 | 0.32 | 0.41 | 0.59 | 0.67 | 0.5 | 0.54 | 0.45 | 0.42 |  | 0.4 | 0.51 | 0.59 | 0.5 |  |  | 0.38 |  | 0.43 | 0.19 - 0.39 | %
MPV | 9.2 | - | 9.4 | 9.3 | 9 | 9.5 | 10.1 | - | 9.4 | 10 | 9 | 9.7 | - |  |  |  | 9.4 |  | 9.1 | 8.8 | 9.4 | 9.6 |  | 9 | 8.5 | 8.8 | 8.8 | 9.4 | 9.4 | 10 | 10 | 9.9 | 9.6 | 9.3 | 9.7 | 10.3 |  | 8.6 | 9.8 | 9.8 | 8.4 |  |  | 10 |  | 9.8 | 9.2 - 12.0 | fl
PDW | 11 | - | 12.7 | 10.6 | 9.7 | 11.2 | 14.3 | - | 11.1 | 14.3 | 10.3 | 12.6 | - |  |  |  | 11 |  | 11.4 | 10.1 | 11.8 | 12.1 |  | 10.4 | 9.7 | 10 | 9.7 | 11 | 11.5 | 12.6 | 12.4 | 11.4 | 11.8 | 10.7 | 11.9 | 12.6 |  | 9.5 | 12.3 | 11.6 | 8.5 |  |  | 12.3 |  | 11.5 | 9.6 - 15.2 | %
Diff
Blasts | [46 empty] | 0 - 0 | %
Seg neutrophil | 41.4 | 46.6 | 43.2 | 39.5 | 36.6 | 43.3 | 35.9 | 41.5 | 32.5 | 28.8 | 22.4 | 32 | 25.7 | 28 | 32.1 | 32.2 | 24 | 26.9 | 29.9 | 47.8 | 24.7 | 37.2 | 28.4 | 34 | 46.7 | 31.8 | 41.8 | 28.3 | 33.3 | 42.4 | 30.3 | 33 | 38.9 | 33.3 | 32 | 21.9 | 50 | 15.1 | 39.6 | 44.8 | 48.3 | 31.8 | 55.6 | 36.5 | 33.4 | 38.5 | 40 - 73 | %
Lymphocyte | 56.7 | 49.4 | 54.1 | 57.5 | 61.5 | 53.7 | 61.7 | 53 | 64.2 | 67.8 | 74.8 | 65.3 | 67.2 | 62.5 | 65.1 | 65.9 | 63.7 | 71 | 66.6 | 50.1 | 71.7 | 60.4 | 69.2 | 62.7 | 50.8 | 58 | 54.2 | 67.9 | 63.5 | 55.8 | 66.4 | 60 | 51.2 | 56.6 | 65.6 | 71.3 | 40 | 68.3 | 55.4 | 45.4 | 45.4 | 56.2 | 37.2 | 57.9 | 61.2 | 58.2 | 20 - 44 | %
Monocyte | 1.4 | 2.3 | 1.2 | 2 | 1.3 | 2.5 | 1.3 | 3.6 | 2.3 | 2 | 1.6 | 1.5 | 6.1 | 7.7 | 1.7 | 1.2 | 11.7 | 1.3 | 2.6 | 1.7 | 2.5 | 1.7 | 1.4 | 2.5 | 1.8 | 8.6 | 2.3 | 3 | 2.1 | 1.4 | 2.2 | 4 | 8.8 | 8.6 | 1.5 | 6 | 6 | 13.1 | 3.9 | 8.8 | 5.4 | 10.7 | 4.6 | 4 | 4.1 | 2.6 | 3.0 - 13 | %
Eosinophil | 0.1 | 1.3 | 1.2 | 0.7 | 0.2 | 0.4 | 0.3 | 0.9 | 0.4 | 0.4 | 0.3 | 0.5 | 0.3 | 1 | 0.4 | 0.1 | 0.2 | 0.5 | 0.5 | 0 | 0.3 | 0.1 | 0.7 | 0.2 | 0.3 | 1.2 | 1.4 | 0.4 | 0.8 | 0.1 | 0.4 | 2 | 0.4 | 0.6 | 0.2 | 0.1 | 4 | 3.3 | 0.4 | 0.5 | 0.7 | 0.7 | 2.4 | 1.3 | 0.8 | 0.2 | 0 - 6 | %
Basophil | 0.4 | 0.4 | 0.3 | 0.3 | 0.4 | 0.1 | 0.8 | 1 | 0.6 | 1 | 0.9 | 0.7 | 0.7 | 0.8 | 0.7 | 0.6 | 0.4 | 0.3 | 0.4 | 0.4 | 0.8 | 0.6 | 0.3 | 0.6 | 0.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.3 | 0.7 | 1 | 0.7 | 0.9 | 0.7 | 0.7 | 0 | 0.2 | 0.7 | 0.5 | 0.2 | 0.6 | 0.2 | 0.3 | 0.5 | 0.5 | 0 - 3 | %
Band neutrophil | [46 empty] | 0 - 5 | %
Metamyelocytes | [46 empty] | 0 - 1 | %
Myelocytes | [46 empty] | 0 - 0 | %
promyelocytes | [46 empty] | 0 - 0 | %
Nucleated RBC | [46 empty] | 0.0 - 0.0 | COI
Immature cells | [46 empty] | 0 - 0 | %
Atypical lymph | [46 empty] | 0 - 0 | %
Neutrophil # | 8.86 | 10.91 | 8.09 | 7.99 | 4.96 | 4.84 | 6.11 | 6.43 | 5.51 | 4.87 | 3.22 | 5.27 | 5.7 | 4 | 7.58 | 4.47 | 4.2 | 4.09 | 5.71 | 8 | 4.22 | 8.15 | 3.77 | 6.35 | 8.68 | 5.91 | 9.27 | 3.98 | 4.35 | 3.64 | 5.75 | 7.51 | 12.55 | 7.74 | 6.29 | 3.55 | 13.02 | 2.55 | 7.5 | 10.14 | 12.43 | 5.31 | 12.08 | 5.4 | 6.75 | 7.79 | 1.7 - 7.0 | 10E3/uL
Lymphocyte # | 12.14 | 12.14 | 10.14 | 11.62 | 8.31 | 6.01 | 10.51 | 8.19 | 10.88 | 11.47 | 10.79 | 10.76 | 14.9 | 8.92 | 15.37 | 9.15 | 11.12 | 10.78 | 12.72 | 8.39 | 12.24 | 13.23 | 9.2 | 11.71 | 9.47 | 10.79 | 11.99 | 9.54 | 8.31 | 4.79 | 12.64 | 13.66 | 16.53 | 13.19 | 12.89 | 11.55 | 10.42 | 11.45 | 10.49 | 10.26 | 11.68 | 9.42 | 8.09 | 8.58 | 12.36 | 11.81 | 1.0 - 4.0 | 10E3/uL
Monocyte # | 0.3 | 0.53 | 0.23 | 0.4 | 0.17 | 0.28 | 0.22 | 0.55 | 0.39 | 0.33 | 0.23 | 0.25 | 1.35 | 1.1 | 0.4 | 0.17 | 2.04 | 0.2 | 0.49 | 0.28 | 0.43 | 0.37 | 0.19 | 0.47 | 0.34 | 1.6 | 0.51 | 0.42 | 0.27 | 0.12 | 0.42 | 0.91 | 2.84 | 2.01 | 0.29 | 0.97 | 1.56 | 2.19 | 0.74 | 1.98 | 1.4 | 1.8 | 1.01 | 0.6 | 0.82 | 0.53 | 0.2 - 0.9 | 10E3/uL
Eosinophil # | 0.03 | 0.3 | 0.22 | 0.15 | 0.03 | 0.05 | 0.05 | 0.14 | 0.06 | 0.07 | 0.05 | 0.08 | 0.07 | 0.14 | 0.09 | 0.01 | 0.03 | 0.08 | 0.1 | 0 | 0.05 | 0.03 | 0.09 | 0.03 | 0.06 | 0.22 | 0.31 | 0.05 | 0.11 | 0.01 | 0.08 | 0.46 | 0.13 | 0.13 | 0.03 | 0.01 | 1.04 | 0.55 | 0.08 | 0.12 | 0.17 | 0.12 | 0.52 | 0.2 | 0.16 | 0.05 | 0.0 - 0.5 | 10E3/uL
Basophil # | 0.09 | 0.1 | 0.06 | 0.06 | 0.05 | 0.01 | 0.14 | 0.15 | 0.11 | 0.17 | 0.13 | 0.11 | 0.16 | 0.11 | 0.16 | 0.09 | 0.07 | 0.4 | 0.07 | 0.06 | 0.14 | 0.14 | 0.04 | 0.11 | 0.08 | 0.07 | 0.06 | 0.06 | 0.04 | 0.03 | 0.14 | 0.23 | 0.23 | 0.22 | 0.14 | 0.11 | 0 | 0.03 | 0.13 | 0.12 | 0.04 | 0.1 | 0.04 | 0.04 | 0.11 | 0.11 | 0.0 - 0.2 | 10E3/uL
Immature Granulocyte | 0.3 | 0.3 | 0.3 | 0.3 | 0.4 | 0.2 | 0.1 | 0.2 | 0.2 | 0.1 | 0.1 | 0.2 | 0.2 | 0.1 | 0.3 | 0.2 | 0.1 | 0.3 | 0.2 | 0.2 | 0.2 | 0.4 | 0.1 | 0.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.3 | 0.6 | 0.3 | - | 0.3 | 0.2 | 0.2 | 0.2 |  | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.1 | 0.3 | 0.4 |  | %
Immature Granulocyte # | 0.06 | 0.07 | 0.05 | 0.06 | 0.05 | 0.02 | 0.02 | 0.03 | 0.03 | 0.02 | 0.02 | 0.03 | 0.05 | 0.02 | 0.06 | 0.03 | 0.02 | 0.04 | 0.03 | 0.03 | 0.03 | 0.08 | 0.01 | 0.04 | 0.05 | 0.03 | 0.05 | 0.03 | 0.04 | 0.05 | 0.05 | - | 0.11 | 0.05 | 0.04 | 0.03 |  | 0.02 | 0.04 | 0.06 | 0.11 | 0.09 | 0.12 | 0.02 | 0.06 | 0.08 |  | 10E3/uL
CBC FLAG

## Table 4. In-vivo blood concentr
|  |  | 0d | 1d | 1w | 2w | 4w | 8w | 12w
#13 (Bronchus / 5 cc) | <0.00010 (ug/mL) | Galium | 7e-05 | 0.012 | 0.03165 | 0.0081 | 0.0046 | 0.0032 | 0.0025
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0005 | 0.0016 | 0.0003 | 0.0002 | 0.0001 | 0.0001
#14 (Bronchus / 5 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.0123 | 0.0037 | 0.0017 | 0.0011 | 0.0007 | 0.0007
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
#15 (Bronchus / 3 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.0093 | 0.003 | 0.0018 | 0.0026 | 0.0021 | 0.0065
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002
#16 (Alveolus / 3 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.00861 | 0.0039 | 0.0049 | 0.005 | 0.0072 | 0.0066
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0002 | 0.0002
#17 (Bronchus / 1 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.0026 | 0.0003 | 0.0002 | 0.0001 | 0.0001 | 0.0001
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
#18 (Ablation / retrieve / 3.3 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.0003 | 0.0002 | 0.0002 | 0.0004 | 0.0008 | 0.0008
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
#19 (Ablation / retrieve / 0.45 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.0003 | 0.0001 | 0.00012 | 0.00067
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001
#20 (Ablation / remain / 1.42 cc) | <0.00010 (ug/mL) | Galium | 5e-05 | 0.00174 | 0.00379 | 0.00417 | 5e-05
| <0.0012 (ug/mL) | Indium | 0.0001 | 0.0001 | 0.0001 | 0.0001 | 0.0001

## Table 5. In-vivo suction or pas
| E-GaIn (cc) | Post procedure volume | After retrieve volume | After 1d volume | After 1w volume | After 2w volume | After 4w volume | After 6w volume | After 10w volume | After 12w volume
#6 (Ablation, Alveoli) Suction X | 13.8 | 4.822 |  |  | 3.840 (20.36%) |  |  | 3.662 (24.06%) | 3.315 (31.25%)
#7 (Ablation, Alveoli) Suction X | 4.5 | 8.608 | 5.369 (37.63%)
#8 (Ablation, Alveoli) Suction X | 0.5 | 3.849 | 2.025 (47.39%) |  | 1.967 (48.90%) |  |  | 1.734 (54.95%) |  | 1.233 (67.97%)
#9 (Ablation, Bronchus) Suction O | 0.2 | 2.551 | 0.653 (74.40%) |  |  | 0.459 (82.01%) | 0.382 (85.03%) |  |  | 0.099 (96.12%)
#10 (Ablation, Bronchus) Suction O | 0.75 | 4.699 | 0.328 (93.02%) |  |  | 0.117 (97.51%) |  | 0.016 (99.66%) |  | 0.01 (99.79%)
#12 (Ablation, Bronchus) Suction O | 3.85 | 40.304 | 20.219 (49.83%) |  | 7.082 (82.43%) |  |  | 5.055 (87.46%)
#13 (Non-ablation, Bronchus) Suction X | 5 | 30.297 |  | 7.882 (73.98%) | 2.766 (90.87%) |  |  |  |  | 0.444 (98.53%)
#14 (Non-ablation, Bronchus) Suction X | 5 | 16.525 |  | 3.059 (81.49%) | 0.531 (96.79%) |  |  |  |  | 0.092 (99.44%)
#15 (Non-ablation, Bronchus) Suction X | 3 | 14.923 |  | 8.073 (45.90%) | [5 empty] | 0.321 (97.85%)
#16 (Non-ablation, Alveoli) Suction X | 3 | 13.047 |  | 5.338 (59.09%) | [5 empty] | 3.251 (75.08%)
#17 (Non-ablation, Bronchus) Suction X | 1 | 7.946 | [7 empty] | 0.018 (99.77%)
#18 (Ablation, Bronchus) Suction O | 3.3 | 7.082 | 1.292 (81.76%) | [6 empty] | 0.024 (99.66%)
#19 (Ablation, Bronchus) Suction O | 0.45 | 3.242 | 0.902 (72.18%) | 0.668 (79.40%) | 0.488 (84.95%) | 0.376 (88.40%) | 0.324 (90.01%) |  |  | 0.309 (90.47%)
#20 (Ablation, Alveoli) Suction O | 1.42 | 4.689 |  | 1.954 (58.33%) | 1.691 (63.94%) | 1.525 (67.48%) | 1.361 (70.97%)

## Table 6. Ex-vivo
Date | No. | Category | Part | E-GaIn(cc) | Controlled mode | Cumulative time(s) | 60 ℃ time(s) | 80 ℃ time(s) | Ws | Impedance AVG.(Ω) | Peak temp.(℃) | E-GaIn temp. AVG.(℃) | Proximal temp. AVG.(℃) | Distal temp. AVG.(℃) | Tumor temp. AVG.(℃) | Pleura damage depth(mm) | Note
2022.08.10 | #2 | N | left middle | 0.7 | Temp 80℃ | 254 | 97 |  | 17253 | 86.17 | 79 | 63.26 | 59.91 | 75.58 |  |  | Ground connection to lung
2022.08.10 | #3 | N | left upper | 0.4 | Temp 80℃ | 600 | 44 | 143 | 12405 | 163.84 | 80 | 76.8 | 68.4 | 57.6 |  |  | Ground connection to lung
2022.08.10 | #4 | N | right middle | 0.5 | Temp 80℃ | 435 | 72 | 207 | 17785 | 101.48 | 80 | 72.73 | 80.27 | 64.33 |  |  | Ground connection to lung
2022.08.10 | #6 | N | right middle | 0.5 | Temp 80℃ | 600 | 188 |  | 48595 | 86.37 | 77 | 65.33 | 70.27 | 61.78 |  |  | Ground connection to lung
2022.08.10 | #7 | N |  | 0.4 | Temp 80℃ | 600 | 49 | 459 | 37915 | 79.64 | 80 | 74.69 | 63.87 | 51.98 |  |  | Ground connection to lung
2022.08.10 | #8 | N | left upper | 0.4 | Temp 80℃ | 143 | 54 | 118 | 9783 | 102.92 | 80 | 65.37 | 79.31 | 85.63 |  |  | Ground connection to lung
2022.08.10 | #9 | N | right lower | 0.8 | Temp 80℃ | 600 | 43 | 118 | 20413 | 66.89 | 80 | 76.93 | 65.27 | 65.65 |  |  | Ground connection to lung
2022.08.10 | #10 | N | right upper | 0.4 | Temp 80℃ | 420 | 84 |  | 21572 | 100.71 | 79 | 69.75 | 71.58 | 52.64 |  |  | Ground connection to lung
2022.08.19 | #1 | N | left lower | 0.5 | Temp 80℃ | 300 | 60 | 123 | 18914 | 257.71 | 85 | 68.11 | 66.1 | 80.26 |  |  | Ground connection to lung
2022.08.19 | #2 | N | right lower | 0.5 | Temp 80℃ | 48 | 25 |  | 4267 | 123.79 | 77 | 57.69 | 46.63 | 41.02 |  |  | Ground connection to lung
2022.08.19 | #3 | N | right middle | 0.5 | Temp 80℃ | 418 | 165 | 357 | 30212 | 101.61 | 86 | 61.4 | 60.82 | 74.53 |  |  | Ground connection to lung
2022.08.19 | #4 | N | left middle | 0.5 | Temp 80℃ | 300 | 44 | 89 | 9549 | 150.09 | 80 | 72.49 | 68.67 | 48.68 |  |  | Ground connection to lung
2022.08.25 | #1 | N, D | left lower | 0.8 | Temp 80℃ | 300 | 55 | 110 | 20852 | 47.02 | 81 | 71.68 | 58.78 | 72.11 |  | 2.8
2022.08.25 | #2 | N, D | right lower | 0.5 | Temp 80℃ | 240 | 79 | 98 | 10298 | 58.17 | 81 | 49.45 | 46.25 | 60.21 |  | 4.8
2022.08.26 | #1 | N, D | right lower | 0.55 | Temp 65℃ | 300 | 22 |  | 8504 | 58.93 | 66 | 63.18 | 69.6 |  |  | 4.4
2022.08.26 | #2 | N, D | left lower | 0.6 | Temp 80℃ | 600 | 42 | 104 | 32730 | 47.26 | 80 | 77.06 | 79.76 | 79.39 |  | 16.9
2022.08.26 | #3 | T, D | left middle | 0.4 | Temp 80℃ | 300 | 56 | 97 | 17257 | 59.68 | 81 | 71.56 |  |  | 96.38 | 3.8
2022.08.26 | #4 | D | left middle | 0.5 | Temp 80℃ | 255 | 30 | 81 | 20970 | 48.54 | 81 | 74.41 | 63.57 |  |  | 5
2022.09.01 | #2 | T, D | right upper | 0.8 | Temp 80℃ | 300 | 208 |  | 34295 | 65.72 | 71 | 51.64 | 36.49 |  | 63.96 | 6.6
2022.09.01 | #3 | T, D | right lower | 0.1 | Temp 80℃ | 300 | 20 | 37 | 9601 | 57.69 | 81 | 77.1 |  |  | 90.93 | 2.2
2022.09.02 | #1 | T | left middle | 1 | Temp 80℃ | 300 | 116 |  | 34041 | 46.09 | 73 | 60.52 |  |  | 66.66
2022.09.08 | #1 | T | left upper | 0.6 | Temp 80℃ | 300 | 33 | 48 | 20583 | 60.48 | 81 | 74.81 |  | 65.49 | 91.43
2020.10.20 | #3 | D | left middle | 0.45 | Temp 80℃ | 600 | 12 | 18 | 19703 | 59.47 | 90 | 78.51 |  |  |  | 12.4
2020.10.21 | #2 | D |  | 0.4 | Temp 80℃ | 600 | 13 | 94 | 16118 | 50.68 | 81 | 78.91 |  |  |  | 10.2
2020.10.21 | #3 | D |  | 0.5 | Temp 80℃ | 600 | 17 | 26 | 14875 | 132.26 | 93 | 78.61 |  |  |  | 10.4
n=25
N=16 | Normal Lung
T=5 | Tumor model
D=11 | pleura depth

## Table 7. Tissue temp
No. | E-GaIn temp. AVG.(℃) | Proximal temp. AVG.(℃) | Distal temp. AVG.(℃)
#1 | 63.26 | 59.91 | 75.58
#2 | 76.8 | 68.4 | 57.6
#3 | 72.73 | 80.27 | 64.33
#4 | 65.33 | 70.27 | 61.78
#5 | 74.69 | 63.87 | 51.98
#6 | 65.37 | 79.31 | 85.63
#7 | 76.93 | 65.27 | 65.65
#8 | 69.75 | 71.58 | 52.64
#9 | 68.11 | 66.1 | 80.26
#10 | 57.69 | 46.63 | 41.02
#11 | 61.4 | 60.82 | 74.53
#12 | 72.49 | 68.67 | 48.68
#13 | 71.68 | 58.78 | 72.11
#14 | 49.45 | 46.25 | 60.21
#15 | 63.18 | 69.6 | Missing data
#16 | 77.06 | 79.76 | 79.39

## Table 8. Tumor temp
No. | E-GaIn temp. AVG.(℃) | Tumor temp. AVG.(℃)
#1 | 71.56 | 96.38
#2 | 51.64 | 63.96
#3 | 77.1 | 90.93
#4 | 60.52 | 66.66
#5 | 74.81 | 91.43

## Table 9. Depth test
No. | Cumulative time (s) | Chest wall damage depth (mm)
#1 | 240 | 4.8
#2 | 255 | 5
#4 | 300 | 4.4
#3 | 300 | 2.8
#5 | 300 | 6.6
#6 | 300 | 2.2
#7 | 300 | 3.8
#8 | 600 | 16.9
#9 | 600 | 10.4
#10 | 600 | 10.2
#11 | 600 | 12.4