Fig S1: Chromatograms for Method B from a different laboratory. From top to bottom: IA and IB (zoom) -radiochromatogram of the test sample of [177Lu]Lu-PSMA I&T; II- UV 200 nm chromatogram of the radioactive test sample;  III- mixture of PSMA I&T and  natLu-PSMA I&T UV 200 nm

[Radio]

I A

[Radio]

I B

[UV 200nm]

II

[UV 200 nm]

III

[Radio]

I A

[Radio]

I B

[UV 200nm]

II

[UV 200 nm]

III

Validation data for Method B from a different laboratory

| Criteria | Validation parameter | Method | No of repetitions | Aceptance Criteria | Found Method B Lab No. 2 |
| --- | --- | --- | --- | --- | --- |
| Identity | Accuracy | Retention time [177Lu]Lu-PSMA I&T | 6 | sd (%)<1% | 0.08% (Rt 8.44min) |
| Identity | Accuracy | Retention time natLu-PSMA I&T | 6 | sd (%)<1% | 0.08% (Rt 8.34min) |
| Identity | Accuracy | Retention time PSMA I&T | 6 | sd (%)<1% | 0.08% (Rt 7.61min) |
| Identity | Specificity | Retention time [177Lu]Lu-PSMA I&T vs natLu-PSMA I&T | 6 | deviation < 0.5 min | 0.1 min |
| Radiochemical purity | Specificity | Resolution PSMA I&T vs natLu-PSMA I&T | 6 | Resolution >2 | 3.1 |
| Radiochemical purity | Precision | RSD of peak areas of [177Lu]Lu-PSMA I&T | 6 | sd (%)≤2% | 1.6% |
| Chemical purity - content | Precision | Sum of the peak areas of PSMA-I&T + Lu-PSMA-I&T | 6 | sd (%)≤2% | 1.8% |

Fig S2: Chromatograms for PSMA I&T. From top to bottom: I-Radiochromatogram of a test sample of [177Lu]Lu-PSMA I&T with a radiochemical purity of about 96% 4 hr post preparation ; II- natLu-PSMA I&T UV 220 nm 50µl/mL irradiated with 500Gy;  I- natLu-PSMA I&T UV 220 nm 50µl/withourt irradiation
Peak labels: 1: main radiochemical impurity; 2: [177Lu]Lu-PSMA I&T; 3:  impurity of  natLu-PSMA I&T collected for MALDI-TOF MS; 4: natLu-PSMA I&T.

Fig S3: MALDI-TOF MS of impurity collected after irradiation of natLu-PSMA I&T at 500 Gy (Peak 3 according to figure S2).

Figure S4: Representative TLC Chromatograms for [177Lu]Lu-Chloride solution, [177Lu]Lu-colloid (neutralized [177Lu]Lu-Chloride solution), [177Lu]Lu-DTPA and the test solution [177Lu]Lu-PSMA I&T in TLC system 1 (left) and system 2 (right).

[177Lu]Lu-Chloride solution

[177Lu]Lu-colloid (neutralized [177Lu]Lu-Chloride solution)

[177Lu]Lu-DTPA

test solution [177Lu]Lu-PSMA I&T

Figure S5: UV/Vis absorption spectrum of PSMA I&T.