Supplement 2. Acoustic intensity output according to waveform generator settings

In the main text of the paper, the cycle count and amplitude of the tone burst generated from the waveform generator were set to 35 cycles and 500 mVpp, respectively.

The number of cycles of the tone burst signal was determined such that the spatial-peak pulse-average intensity (Isppa) would converge. Fig. S2-1 shows the acoustic pressure and intensity at the focal point of the VZPUT when the peak-to-peak voltage output from the waveform generator (Vsgpp) is 40 mVpp. As shown in Fig. S2-1, the peak positive pressure (p+), peak negative pressure (p-), and Isppa converge at approximately 35 cycles.

Figure S2-1. Acoustic pressure and intensity at the focal point with a tone burst signal of 1.1 MHz center frequency and Vsgpp = 40 mVpp: (a) the pressure waveform, (b) p+ and p-, (c) Isppa.

Typical diagnostic ultrasound transducers1, FUS transducers2, and High-intensity FUS (HIFU) transducers1 deliver ultrasound with Isppa of approximately 0.1-100 mW/cm2, 3-54 W/cm2, and 100-10,000 mW/cm2 to the focal point, respectively.

Fig. S2-2 shows the acoustic pressure and intensity at the focal point of the VZPUT when the tone burst is 35 cycles. As shown in Fig. S2-2, the VZPUT shows a wide range of intensity output from the FUS region to the HIFU region in the water tank experiment.

Figure S2-2. Acoustic pressure and intensity at the focal point with a tone burst signal of 35 cycles and 1.1 MHz center frequency: (a) the pressure waveform, (b) p+ and p-, (c) Isppa.

Supplementary References

- Dubinsky, T. J., Cuevas, C., Dighe, M. K., Kolokythas, O. & Hwang, J. H. High-intensity focused ultrasound: current potential and oncologic applications. AJR Am. J. Roentgenol. -NEW SERIES- 190, 191 (2008), DOI: https://doi.org/10.2214/AJR.07.2671

- Lee, W., Weisholtz, D. S., Strangman, G. E., & Yoo, S. S. Safety review and perspectives of transcranial focused ultrasound brain stimulation. Brain Neurorehabil, 14. (2020), DOI: https://doi.org/10.12786/bn.2021.14.e4