Table1: The technic details of MEDICAL ELECTRONIC NEEDLE Author; Institution; Year; Material structure/Key component; Needle size; Ablation/measurement frequency range; Approaches /Theories applied; Usage in medical treatment; Features. Julian Moore et al. [41]; The University of Georgia Athens, GA, USA; 2018; Ablation catheter; LC circuit; two coils; Inside stylet :0.85mm Outsit sheath stylet:1.75 mm; 20 mm spheric space; Law of Biot-Stavart; Electromagnetic induction; Magnetic resonance; Thermal ablation for tumour/lesion tissues; The ablation catheter is powered by an RL circuit and two coils wirelessly, and the power efficiency is about 49%. The temperature range of the catheter is 64 °C to 102°C.. Julian Moore et al. [42]; The University of Georgia Athens, GA, USA; 2020; Ablation catheter; RX tank circuit; TX coil; RX capacitor; RX coil; 6.5 gage; 12 cm in length;; Width of 12 mm; length of 21 mm; Law of Biot-Stavart; Inductive power transfer theory;; Thermal ablation for tumour/lesion tissues; The ablation catheter is powered wirelessly by a transmission circuit and coil and a receive coil. It has an average maximum received power of 15W, Average maximum efficiency of 63.27%. Maximum heat 85°C. Rongzhou Lin et al. [37]; Institute for Health Innovation and Technology, National University of Singapore; 2022; Eight PEDOT-MWCNT microelectrodes integrated (diameter 50um) on a medical-approval needle (Guage 18, BD) by using epoxy adhesives(MED-301, EPO-TEK); Guage 18; 1Hz to 1MHz; Electric properties change in Extravasation tissue; Early detection of extravasation; The needle can detect impedance differences in different tissues New fabrication technic for integrating the microelectrodes. Jaeho Park et al [31]; Department of Mechanical Engineering, Korea Advanced Institute of Science and Technology (KAIST); 2018; The biopsy needle consists of two insulations, one PSA coating, an Electrode attachment, and a Needle hub. Measurement tool: LCR meter(E4980a, Keysight, USA) , multiplexer(34904a,Keysignt,USA), and control PC; Diameter 1.2mm Length 15cm; 1KHz to 1MHz; Tissue property changes caused by lesions; Electrical impedance spectroscopy; Not suffer from errors induced by the electrode polarization. Sensing range in several mm Low accuracy of tissue discrimination in heterogeneous tissue structure. A. Kim et al [49]; Department of Electrical and Computer Engineering, Temple University, Philadelphia; 2020; Multiple Implantable EA microprobes, an ultrasonic transmitter, a soft tissue, and a receiver; N/A; N/A; Electrolytic ablation; ultrasound wirelessly power transfer.; Electrolytic ablation for lesion tissue; Ablation microprobes powered wirelessly by ultrasound Ablating lesion tissue by changing Ph in vivo. Abbasi Muhammad Aitzaz et al [50]; Department of Electronics Engineering, Gachon University; 2019; Configurable impedance network analyzer(ADuCM350, analogue devices, Norwood, Massachusettes US), a Monopolar needle working Electrode, a control PC with GUI; N/A; N/A; Monopolar injection, discrete Fourier transfer; Bio-impedance sensing for patient’s tissue; Monopolar insertion measurement method Applied discrete Fourier transfer-based network analyzer. Ultrasonic imaging is utilized for measurement. Jaeho Park et al [36]; Department of Mechanical Engineering, Korea Advanced Institute of Science and Technology (KAIST); 2020; Biopsy needle with multiple sensors (RF Medical, Korean), a heat-shrink tube made of polyethene terephthalate(PET), a multiplexer, an LCR meter, a Source meter, an Electrometer, a Control PC; Diameter 1.2 mm 18 gauge Length of 10.5 cm; N/A; The differences in pH, electrical conductivity, and glucose concentration between normal tissue and cancer tissue were reported; Bio-impedance, electrical conductivity, pH, glucose concentration detection for cancer diagnosis.; A new fabrication method is utilized for multiple-sensor microfabrication Using differences in the values of Bio-impedance, electrical conductivity, pH, and glucose concentration between normal and cancer tissue to detect cancer early.. Jaeho Park et al [48]; Department of Mechanical Engineering Korea Advanced Institute of Science and Technology (KAIST); 2021; An RF ablation needle, a temperature sensor, a pressure sensor,; Diameter 1.5mm. Length 15cm.; N/A; Micro-sensor; Reducing the steam pop during the RF thermal ablation in tissue; The needle can detect 0 – 500 KPa pressure and measure the temperature in vivo. V. Mishra et al. [26]; Dartmouth College; 2012; A standard biopsy needle, a bipolar RCA plug(CP3-1005-ND,CUI,Inc,Tualatin, OR), a two-shielded coaxial cable, a impedance analyzer(Agilent Technologies, Santa Clara, CA), and a laptop; 18 gauge. Inner element diameter 0.88 mm. Outer element diameter 1.09mm; N/A; Impedance changes in cancerous solid-organ; Diagnosis of solid-organ cancer by measuring its electrical property.; A complex electrical impedance Z=R+JX represents the ratio of the voltage to current as the result of measurement. Injecting 1-mA current between two electrically isolated elements for measurement.. ZHUOQI CHENG et al. [38]; Mechanical Engineering Department, Imperial College London; 2018; A concentric electrode needle(CEN); a Smart Venous Entry Indicator(SVEI) consists of a control unit(Atmega328p, Atmel Co., USA) and two LEDs.; N/A; 10kHz to 100kHz; The differences in bio-impedance between skin tissue, vessel wall tissue, and blood; Assassinating peripheral intravenous catheterization; A bioimpedance sensing system used for detecting cancerous tissue. H Kwon et al. [23]; Department of Neurology, Beth Israel Deaconess Medical Center, Harvard Medical School, Boston, MA 02215-5491, United States of America; 2017; A needle electrode (Bowdoin, ME, USA), a Terason t3000 system, impedance analyzer SFB7 (Impedimed, Inc, Brisbane, Australia); 24 gauge; Measuring impedance at frequencies from 10kHz to 1MHz; Electrical impedance myography technique; Measuring the skeletal muscle; Utilized a needle electrode for novel use in Electrical Impedance Myography.. Axel R. Sauter et al. [51]; Department of Clinical and Biomedical Engineering, Oslo University Hospital Rikshospitalet, Oslo, Norway; 2014; An ATL HDL 5000ultrasound unit (ALT, Bolthell, USA), A insulated needle (Stimuplex, A; B.Braun, Melsungen,Germany), a nerve stimulator(Stimuplex HNS 12; B.Braun, Melsumgen, Germany), a reference electrode (Blue Sensor, Q-00-A, Ag/AgCl, Ambu Medicotest A/S, Denmark), a Solartron complex impedance measurement system(SI 1260 and SI 1294, Solartron Group PLC, Hampshire, UK); 21-gauge Length 100mm; 1.26 Hz to 398kHz; A novel algorithm based on complex impedance measurement; Discriminating the nerve tissue from other tissue types; A novel algorithm is applied for bio-impedance of nerve tissue measurement.. Jinhwan Kim et al [30]; School of Mechanical Engineering, Gwangju Institute of Science and Technology (GIST); 2020; Prototype of EoN with improved IDEs; an impedance analyzer (Famry reference 600; Gamy instrument, USA); a laptop; 720 μm; 100Hz to 1MHz; A new fabrication method for improved IDE sensors; Discriminating the nerve tissue from other tissue types; An improved IDEs fabricated on a needle. Joho Yun et al. [28]; Department of Biomedical Science and Engineering, Gwangju Institute of Science and Technology (GIST), South Korea; 2015; A hypodermic needle (100 Sterican, BBran,700ym in diameter); A Impedance analyzer(Reference 600, Garmy Instruments); a laptop;; 22 gauge; Measured over the frequency range of 100Hz- 1MHz; Electrochemical impedance spectroscopy; Interdigitated electrodes (IDE); Measuring the bio-impedance of hypodermic tissues; Utilizing electrochemical impedance spectroscopy on a needle(EoN) with bipolar interdigitated electrodes. Zhuoqi Cheng et al [22]; Department of Advanced Robotics, Istituto Italiano di Tecnologia, Genova, Italy; 2020; A power insulator (Linear Technology,US); SmartProb; Arduino DUE board (Arduino Inc.Italy); Impedance converter based on AD5933(Analog Devices, Norwood,MA,USA); N/A; 10KHz to 100kHz with a step of 10kHz; Electrical bioimpedance (EBI) sensing system based on a concentric needle electrode; Detecting bioimpedance of human tissue; Utilized a GUI to display the magnitude and phase angle in different measuring frequencies.. Reihane Mahdavi et al [16]; Nano Bioelectronics Devices Lab, Cancer Electronics Research Group, School of Electrical and Computer Engineering, Faculty of Engineering, University of Tehran, Tehran, P.O. Box 14395/515, Iran; 2021; Two medical-grade stainless steel needles(G16 I.V cannula and G21 disposable spinal cannula); impedance spectroscopy; Needle cannula: 16 gauge Spinal cannula with 2-electrode: 21 gauge; Impedance magnitude in 1kHz and impedance phase slop in the frequency range of 100kHz-500kHz; Biology theory/mechanism of distinction between normal LNs and cancer LNs.; Breast cancer diagnosis; Created a new system for breast cancer diagnosis.. Jack Sharp et al. [20]; Wolfson School of Mechanical and Manufacturing Engineering, Loughborough University, Loughborough, Leicestershire LE11 3TU; 2017; Stimuplex Needle; impedance measurement devices; conductive wires; 20-gauge; Measuring with 86kHz, and a constant current of 0.5 mA; Bipolar and monopolar setups for bioimpedance measurement; Measuring Bioimpedance of Nerve tissue; Compared two bipolar setup and monopolar setups in bioimpedance measurement. Laurent Schoevaer-dts et al [39]; Dept. of Mechanical Engineering, KU Leuven - University of Leuven, Leuven, Belgium; 2018; Prototype of micro sensing needle; AD5933(Analog Devices,Norwood, Massachusetts); 30um; 1kHz to 100kHz; Micro bipolar needle with electrodes; Measuring the impedance of eye vein; A new prototype of a micro-sensing device for eye surgery. V. Mishra et al. [10]; Thayer School of Engineering, Dartmouth College, Hanover, New Hampshire; 2013; EIS-Bx; Diameter 1.09 mm; 1 to 100kHz; Utilized EIS-Bx device in pathological tissue detection; Pathological tissue detection; Utilized EIS in medical treatment. H Kwon et al. [23]; Department of Neurology, Beth Israel Deaconess Medical Center, Harvard Medical School, Boston, MA 02215-5491, United States of America; 2017; Stainless-steel needle manufactured by FHC, Inc (Bowdoin ME, USA); 19-gauge; Measuring 10kHz to 1MHz.; Electrical impedance myography (EIM); A technique to measure the electrical impedance of skeletal muscle; A system measures skeletal muscle with different electrodes angle.. Jannis Dickow et al [44]; Translational Interventional Electrophysiology Laboratory, Mayo Clinic, 1216 2nd St. SW, Rochester, MN 55905, USA; 2020; SERF catheter (Durable, Thermedical, Waltham, MA, USA), 64-element 5-10Mhz 2D ICE catheter (ACUSON AcuNavTM Ultrasound Catheter, BiosenseWebster Inc., Mountain View, CA, USA); 25-gauge; N/A; Saline-enhance radiofrequency needle tip ablation; Ablate the lesion tissue; An RF ablation system enhanced by heated saline.. Joho Yun et al. [8]; Department of Biomedical Science and Engineering, Gwangju Institute of Science and Technology (GIST), South Korea; 2017; EIS needle: Hypodermic needle integrated with bipolar interdigitated electrodes (IDEs); A PCB;; N/A; 1Hz-1MHz; Measuring by 20mVrms; Electrochemical impedance spectroscopy applied in cancer tissue discrimination; Cancer tissue detection; A bioimpedance sensing needle fabricated with micro-electrical impedance spectroscopy sensors. Jinhwan Kim et al [29]; School of Mechanical Engineering, Gwangju Institute of Science and Technology (GIST), South Korea; 2018; A hypodermic needle integrated with IDEs; an impedance analyzer (Gamry reference 600, Gamry Instrument, USA); 22-gauge; 100Hz-1MHz; 20mV; Incremental compensation method applied in EoN to increase the measuring accuracy; Pathological tissue detection; Utilized incremental compensation method to increase the measurement accuracy.. Muhammad Aitzaz Abbasi et al[27]; Department of Electronics Engineering, Gachon University; 2020; AD5933 chip (Analog Devices, Norwood, MA, USA); ARM Cortex-M4 microcontroller; a monopolar needle electrode(Gilgry, CA, USA); ultrasound equipment (SONON300C, Healcerion Inc, Seoul, Korea); 23 gauge; 10Hz to 100KHz; Monopolar bioimpedance measurement, novel algorithm.; Monitor the needle position in the intra-articular space; Used a novel algorithm based on impedance data to monitor the needle in the intra-articular space. Maria Moris et al[43]; Division of Gastroenterology, Mayo Clinic; 2016; EUS RFA needle prototype; a standard electrosurgical unit(ERBE VIO 300,Erbe USA, Marietta,GA,USA); 22 gauge; 20°C–100°C; endoscopic ultrasound (EUS) guiding, radiofrequency ablation.; Tissue radiofrequency ablation; Application of endoscopic ultrasound (EUS) in RFA needle. Joho Yun et al. [9]; Department of Biomedical Science and Engineering, Gwangju Institute of Science and Technology (GIST); 2016; A prototype of EoN; an impedance analyzer( Reference 600,Germry Instruments);; 22 gauge; 100Hz to IMHz; EIS and improved interdigitated electrode (IDE); Cancer tissue detection and discrimination; Fabricated and tested an improved IDE for the sensing needle. Camilla Carpano Maglioli et al [7]; Camilla Carpano Maglioli is with Department of Advanced Robotics; 2017; AD5933 converter (Analog Devices, Norwood, MA, USA); a sterile concentric electrode needle(EMG needle electrodes, FIAB SpA, Firenze, Italy); An Arduino DUE (Arduino Inc. Italy); 0.45 mm diameter; 10kHz to 100kHz; Impedance measurement by electrode needle; Pathological tissue detection; Tested the new needle-based bioimpedance sensing system. H. Kalvøy et al [24]; Department of Clinical and Biomedical Engineering, Oslo University Hospital Rikshospitalet; 2016; An ATL HDI 5000ultrasound unit with a L12-5 transducer; A unsulated needle(Stimuplex A; B. Braun, Melsungen, Germany); A nerve simulator(Stimuple-x HNS 12; B. Braun, Melsungen, Germany);a gel electrode (Blue Sensor, Q-00-A, Ag/AgCl, Ambu Medicotest A/S, Denmark); a Solarton complex impedance stimulator (SI 1260 and SI 1294, Solartron Group PLC, Hampshire, UK);; 21-gauge; 1.26 to 398kHz; Novel algorithm to analyse the complex impedance data; Nerve tissue discrimination; Analyse the complex impedance data in complex plane. Created a new algorithm..