7.5 mm ID

ETconnector

Silicone

tubing

23 cm

Table 1. Insufflating and expiratory volumes, I:E ratios, respiration rates, and minute volumes with i.v. catheters at lung compliances of 25, 50, and 100 mL/cmH2O in the trachea-lung model

___________________________________________________________________________

Oxygen flow: 12 L/min

Lung compliance (mL/cmH2O)	   25		   50 		  100

14 G	Insufflating volume (mL/s)	178 ± 2	189 ± 2	196 ± 2

Expiratory volume (mL/s)	131 ± 4	131 ± 4	131 ± 4

I:E ratio				1:1.37		1:1.45		1:1.50

RR(breaths/min)		12.68		12.25		11.99

MV (mL/min)			4,523 ± 87	4,633 ± 91	4,703 ± 94

2-mm ID TTC

Insufflating volume (mL/s)	178 ± 3	188 ± 3	192 ± 3

Expiratory volume (mL/s)	179 ± 7	179 ± 7	179 ± 7

I:E ratio				1:0.99		1:1.05		1:1.07

RR(breaths/min)		15.06		14.63		14.49

MV (mL/min)			5,350 ±122	5,507 ± 128	5,557 ± 130

Oxygen flow: 15 L/min

14 G	Insufflating volume (mL/s)	216 ± 6 	226 ± 6	237 ± 7

Expiratory volume (mL/s)	145 ± 5	148 ± 5	151 ± 5

I:E ratio				1:1.49		1:1.53		1:1.58

RR(breaths/min)		12.06		11.85		11.66

MV (mL/min)			5,196 ± 100	5,362 ± 105	5,525 ± 109

2-mm ID TTC

Insufflating volume (mL/s)	216 ±5		227 ± 5	238 ± 5

Expiratory volume (mL/s)	197 ± 5	198 ± 5	201 ± 6

I:E ratio				1:1.10		1:1.15		1:1.18

RR(breaths/min)		14.31		13.99		13.75

MV (mL/min)			6,183 ± 95	6,346 ± 106	6,534 ± 103

____________________________________________________________________________

Values are expressed as mean ± SD. N = 10 at each gauge, with a compliance of 25 mL/cmH2O. G = gauge. ID = internal diameter. TTC = transtracheal catheter. I:E = inspiratory time: expiratory time. RR = respiratory rate per minute. MV = minute volume (mL/min).

Table 2. Insufflating, expiratory, and minute volumes in the test lung and detailed lung simulation devices

_________________________________________________________________________				Insufflating Vol.	     Expiratory Vol.	   Minute Vol.

(mL/s)		 (mL/s)	   	    (mL/min)

_________________________________________________________________________

15 L/min

Test lung		   238 ± 5	       201 ± 6	     6,530

Schmidt et al.(#5) 	   242 ± 9	       210 ± 8	     6,720

Hamaekers et al.(#14)  256 ± 5	       218 ± 3	     7,100

12 L/min

Test lung		   192 ± 3	       179 ± 7	     5,557

Schmidt et al.(#5) 	   195 ± 6	       185 ± 6	     5,707

Hamaekers et al.(#14)  202 ± 2	       196 ± 2	     6,000

_________________________________________________________________________

Values are expressed as mean ± SD. These data are values at a lung compliance of 100 mL/ cmH2O with a 2-mm internal diameter (ID) transtracheal catheter (TTC) at oxygen flow rates of 15 and 12 L/min. An ASL 5000 (reference #5) or LS800 lung simulator (reference #14) was used as the lung simulation device. The numbers measured with test lungs, ASL 5000, and LS800 were 10, 5, and 4, respectively. Test lung indicates the test lung in our trachea-lung model. Vol. = volume.