Harmonization of brain PET images in multi-center PET studies using Hoffman phantom scan

Mahnaz Shekari1-3, Eline. E. Verwer4, Maqsood Yaqub4, Marcel Daamen5, Christopher Buckley6, Giovanni B. Frisoni7,8, Pieter Jelle Visser9-12, Gill Farrar6, Frederik Barkhof13,14, Juan Domingo Gispert1-3,15*, Ronald Boellaard4*, on behalf of the AMYPAD Consortium

1Barcelonaβeta Brain Research Center (BBRC), Pasqual Maragall Foundation. Barcelona, Spain

2IMIM (Hospital del Mar Medical Research Institute), Barcelona, Spain

3Universitat Pompeu Fabra, Barcelona, Spain

4Department of Radiology & Nuclear Medicine, Amsterdam, University Medical Centers, location VUmc, De Boelelaan 1117, 1081 HV Amsterdam, The Netherlands

5German Center for Neurodegenerative Diseases (DZNE), Bonn, Germany

6GE Healthcare, Amersham, United Kingdom

7Laboratory of Neuroimaging of Aging (LANVIE), University of Geneva, Geneva, Switzerland

8Memory Center, Department of Rehabilitation and Geriatrics, University Hospitals and University of Geneva Geneva, Switzerland

9Alzheimer Center Amsterdam, Neurology, Vrije Universiteit Amsterdam, Amsterdam UMC

location VUmc, Amsterdam, The Netherlands

10Amsterdam Neuroscience, Neurodegeneration, Amsterdam, The Netherlands

11Alzheimer Center Limburg, School for Mental Health and Neuroscience, Maastricht

University, Maastricht, the Netherlands

12Department of Neurobiology, Care Sciences and Society, Division of Neurogeriatrics, Karolinska Institutet, Stockholm Sweden

13Amsterdam UMC, Vrije Universiteit Amsterdam, Department of Radiology and Nuclear Medicine, De Boelelaan 1117, Amsterdam, Netherlands

14Institute of Neurology and Centre for Medical Image Computing, University College London, UK

15Centro de Investigación Biomédica en Red Bioingeniería, Biomateriales y Nanomedicina, (CIBER-BBN), Barcelona, Spain

*Corresponding authors

Figure_S1. Digital reference object (DRO) provided by toolbox (A) Site-specific DRO after realignment with PET scan (B) Example of PET phantom image with the overlaid contour of site-specific DRO

Figure_S2. Sagittal view of (A) digital reference object (B) smoothed DRO with Gaussian filter with an FWHM of 8 mm (S8-DRO), and (C) Hoffman phantom PET image. Blue counters in image (A) and (B) shows voxels with intensity above 0.98 and voxels included for calculating image-derived activity concentration, respectively

Figure_S3. Schematic of processing steps for estimating effective image resolution for each acquired phantom PET image

Figure_S4 (A) shows the global mean absolute difference (across all voxels) between the acquired 3D Hoffman phantom and the DRO for smoothing kernels ranging from 0.1 to 10 mm FWHM as an example of the behavior of the method for estimating the EIR of a Siemens Biograph mCT 128 PET/CT scanner. The mean absolute global difference converged to a minimum for a Gaussian filter of 4 mm FWHM, which was determined as effective image resolution (EIR). The phantom PET image converged to the target effective image resolution while applying a post-smoothing Gaussian filter of 6.6 mm FWHM, where the mean global absolute difference reached the minimum value (Figure_S4 (B)).

Figure_S4. Mean absolute global difference (±95% confidence interval) between (A) PET and site-specific DRO    smoothed for estimating EIR (B) 8 mm smoothed DRO and phantom PET image for estimating harmonization kernel as a function of FWHM of the Gaussian filters. Phantom PET images were acquired for four frames of 5 min. It was reconstructed with ordered subset expectation maximization (OSEM) algorithms, using four iterations and 24 subsets, without any post-smoothing filter.

A subset of PET images (21 scans) with EIR between 4 and 6 mm was selected for evaluating the possibility of harmonizing PET scans with sharper EIR. Figure_S5 (panels A, B, C, D, and E) shows the visual comparison between 5 different scans with EIRs ranging from 4.3 to 5.9mm. A COV% ranging from 6% to 23% is observed in the images between harmonization, and 3 scans have COV%>15%, exceeding the upper acceptable limit of COV% for image quality. Panel A’, B’, C’, and D’ represent PET scans after harmonizing to a target resolution of 6 mm, producing COV% ranging from 5.50% to 9.63%.

Figure_S5. Visual comparison of phantom PET images acquired in 5 different sites before (A, B, C, D, and E) and after (A´, B´, C´, D´, and E´) harmonization. It should be noted scans with similar EIR can represent different levels of COV% before harmonization, whereas harmonizing images resulted in providing similar image qualities and uniformity without significant signal degradation.

Table_S1 shows image quality and quantitative metrics calculated for scans before and after harmonization. Matching the EIR of the images to 6mm resulted in reducing the variability of COV% from 5.35 to 1.97, and producing images with COV%<15%. However, harmonization did not affect the variabilities of other quantitative metrics (Contrast, GMRC), and all represented similar standard deviations pre and post-harmonization.

| Quantification metrics | Non-Harmonized Mean±SD(range) | Harmonized Mean±SD(range) | Difference Mean±SD(range) |
| --- | --- | --- | --- |
| COV% | 18.21±5.35 (6 to 28.20) | 8.98±1.97 (5.50 to 12.10) | 9.22±4.18 (0.50 to 18.54) |
| GMRC | 0.79±0.01 (0.75 to 0.80) | 0.75±0.01 (0.72 to 0.78) | 0.04±0.006 (0.03 to 0.05) |
| Contrast | 1.85±0.06 (1.66 to 1.92) | 1.66±0.05 (1.52 to 1.74) | 0.18±0.03 (0.014 to 0.25) |
| GMRCerod | 0.97±0.03 (0.92 to 1.04) | 0.96±0.03 (0.89 to 1.02) | 0.02±0.002 (0.01 to 0.02) |
| Contrasterod | 3.61±0.24 (3 to 4.03) | 3.41±0.23 (2.82 to 3.81) | 0.20±0.03 (0.15 to 0.30) |
| Left-to-right GMRC ratio | 1.02±0.008 (0.99 to 1.02) | 1.02±0.009 (0.99 to 1.03) | 0.003±0.002 (-0.006 to 0.002) |
| Cold-spot RC | 0.04±0.01 (0.02 to 0.08) | 0.05±0.01 (0.03 to 0.08) | 0.002±0.001 (-0.004 to 0.006) |

Table_S1. Image quality metrics are measured by toolbox before and after harmonization. This analysis corresponds to 21 scans with EIR ranging from 4 to 5.9 mm, and they were harmonized to result in 6mm EIR. The fourth column represents differences between quantitative metrics before vs after harmonization

Figure_S6 shows the impact of harmonization on GMRC and contrast for both eroded and non-eroded VOI as a function of COV%. Figure_S6 (A) shows that before harmonization, even though PET scans had similar GMRC, high variabilities were observed in the COV% levels, and only 7 PET scans (33.33%) had COV%<15%. Also, GMRC and contrast of eroded VOI were higher than upper acceptable limits for 14.28% (3 scans) and 8.33% (2 scans) respectively. Left-to-right RC ratio and cold-spot RC were minimally affected by harmonization.

Figure_S6. Comparing GMRC and contrast for (A & C) non-eroded and (B & D) eroded VOIs as a function of COV% for harmonized vs non-harmonized PET images. Panel E and F display left-to-right hemisphere RC ratios and cold-spot RC of different reconstructions for harmonized vs non-harmonized images