Materials and methods

Patients

This study was carried out retrospectively on a sample of 49 patients enrolled in a clinical trial between 2016 and 2017 (EUDRACT 2015-005546-63), which received local institutional ethics committee approval at the Azienda USL-IRCCS of Reggio Emilia hospital.

All patients were affected by NETs and were treated with PRRT. According to the trial design, each patient underwent several 177Lu- and 90Y-DOTATOC administrations. Dosimetry was conducted at the first cycle of therapy after a therapeutic injection of 177Lu-DOTATOC. A mean value of 4.2 ± 0.9 GBq of 177Lu-DOTATOC was administered to patients. A maximum of 5 lesions was analysed for each patient, with a total of 154. The clinical trial was conducted before the Lutathera approval by EMA and FDA.

Imaging

All examinations were performed using a hybrid Symbia T2 SPECT/CT (Siemens Healthineers, Germany). The SPECT gamma camera was equipped with a medium-energy general purpose collimator (MEGP). The energy windows (EW) of 177Lu photopeaks were set at 113 keV ± 7.5% and 208.4 keV ± 7.5%. Images were acquired in step and shoot mode, with 32x2 views at 30 s per view. SPECT projections were reconstructed using an iterative algorithm with compensations for attenuation from CT images, scatter and full collimator-detector response in the Siemens E-Soft workstation (Syngo, MI Application version 32B, Siemens Medical Solution, Germany) with Flash 3D iterative algorithm (10 iterations; 8 subsets; Gaussian filter with 4.8 mm cut-off). The TEW scatter correction for the lower EW and the DEW scatter correction for the higher EW were employed.  The reconstructed SPECT images were sampled on a matrix of 128x128x112 elements, resulting in a 4.8 mm cubic voxel size.

The FWHM of the system was measured by Grassi et al [25] and the result was 10.4 ± 0.7 mm.

The imaging protocol consisted of four sequential SPECT/CT scans of the abdomen typically at 1, 24, 44, 72 h p.i. (post injection). If necessary, also the thorax area at 1, 24 and 72 h p.i. was scanned.

A total of 141 lesions in the abdomen and a total of 13 lesions in the thorax were analysed.

Dosimetry workflow

At the first cycle a complete dosimetric evaluation of the selected tumours was performed based on SPECT/CT acquisitions.  The SPECT/CT system was previously calibrated using a cylindrical Jasczcak phantom (Data Spectrum Corporation, USA) filled with a homogenous 177Lu radioactive solution. A calibration factor (CF = 36.5 cps/MBq) was determined by the ratio between the known activity and the measured total counts, following the procedure described by Grassi et al [26]. A series of sequential multiple acquisitions of the phantom was performed. The standard uncertainty from repeating activity measurements was taken.

Subsequent to each SPECT acquisition, a CT image was acquired for attenuation correction. For radiation protection of patients, low resolution CT scans were acquired (90 mAs at the first scan and 30 mAs at the following acquisitions) and no contrast medium was used. As a consequence, most of the lesions were not visible on the CT image. For that reason, all tumours were manually segmented on the SPECT image. Contouring was performed in the Velocity Workstation (Varian Medical System, USA) using a variable threshold (as a percentage of the maximum value) defined by a nuclear medicine physician. To minimize mis-registration errors, contours were outlined on the fused SPECT/CT image acquired at 24h p.i., transferred automatically to all co-registered images and manually translated by the user, in case of need, to adapt them to the lesion site.

Activities were corrected for partial volume effects using recovery coefficients (RCs) previously determined based on phantoms with spherical inserts [27]. RCs as a function of insert volume were fitted with the following exponential curve:

RC(v)=α∙exp-β∙v+γ           (1)

where α, β and γ are the fitting parameters and v (volume) is the independent variable.

Two different exponential curves were used to fit the time-activity points:

f1t=A0∙exp⁡(-λ1∙t)           (2)

f2t=A0∙exp-λ1∙t1-exp -λ2∙t           (3)

Where A0, λ1 and λ2 are the fitting parameters and t (time) is the independent variable. Eq. 2 was used in case of monotonically decreasing data-points or if only 3 time-points were available. Otherwise, Eq. 3 was used.

Time-integrated activities (TIAs) were calculated by solving the integral of the exponential functions, based on the fitting parameters.

Tumour absorbed doses were calculated using the OLINDA1.1 sphere model. S-factors derived from OLINDA1.1 were fitted against mass using a power function, as shown in Fig. 1. In this study, the absorbed dose at the first therapy cycle and the relative uncertainty were calculated.

Uncertainty analysis

In this section, it is briefly described how uncertainty associated with each parameter within the dosimetry workflow was calculated. Please refer to the EANM published guidelines [24] for more details.

- Volume uncertainty u(v) was calculated using the analytical expression:
uvv2=3u(d)d2           (4) 
where d is the equivalent diameter of the outlined lesion, with uncertainty:
u2d=a26+FWHM24ln2           (5) 
where a is the voxel size and FWHM is the resolution of the imaging system.

- Referring to Eq. 1 and assuming the vector b=α, β, γ,vT, the squared standard uncertainty associated with RC was calculated as:
u2RC=gbTVbgb          (6) 
where gb is the vector containing the partial derivatives of first order of RC with respect to b and Vb is the covariance matrix extended by one element, namely the partial derivative of first order of RC(v) with respect to v.

- Uncertainty associated with the number of counts (C) within the VOI was calculated by assuming a Gaussian profile of the counts with standard deviation σ and by propagating the volume, CF and RC uncertainties:
uC=C2∙RC∙uvverf2rσ2-2σr2π1-e-2r2σ2           (7)
where r is the equivalent radius, erf is the error function and σ=FWHM22ln2 .

- Calibration factor uncertainty was determined by applying the LPU of measured activity (nominal accuracy of dose calibrator) and counts (standard deviation from multiple measurements).

- Uncertainties associated with the measured activities were determined by error propagation of CF, RC and C (uncertainty of administered activity was assumed to be negligible).

- Uncertainty associated with the fitting parameters (Eq. 2 and 3) were assumed to be the main source of uncertainty of the time-activity curve fitting.

- Uncertainty associated with the TIA (denoted as A) included both the uncertainty of the fitting parameters and the uncertainty of the activities:
u2A= gpTVpgp+uAAA2           8
where gp is the gradient matrix of A with respect to the vector p containing the fitting parameters, Vp is the covariance matrix and u(A)A is the fraction standard uncertainty associated with the measured activities.

- Uncertainty associated with S-factors, uS, was derived by propagation of the volume error (errors associated with the S-factors against volume fitting parameters were assumed to be negligible).

- Absorbed dose uncertainty was determined by applying the LPU:
uDD2=uAA2+uSS2+2uA,SA∙S           9
where uA,S is the covariance between A and S.

Data analysis and statistics

All absorbed dose calculations and statistical analyses were performed in MATLAB R2019a (The MathWorks Inc., USA). A MATLAB script was developed and used to automatically calculate the uncertainty associated with each parameter within the absorbed dose calculation chain.

Box-plots were used to visualize the distribution of standard uncertainty of each variable included in this analysis. Association between each variable and the absorbed dose uncertainties were qualitatively assessed graphically.

As discussed in the EANM guidance, uncertainty in the absorbed dose is expected to largely depend on the precision with which the lesion volume can be estimated.  An absorbed dose uncertainty (AD uncertainty) curve against lesion volume (v) was determined by least squared fitting. A Power function of Eq. 6 was used to fit the empirical data points:

AD uncertainty v= A∙vB           (10)

where A and B are the fitting parameters and v is the independent variable.

In this study, we further evaluated the expected improvement in absorbed dose uncertainty attainable with potential improvement in accuracy of the volume estimation. This was achieved by repeating all uncertainty calculations assuming a range of system spatial resolutions, such as those typical of 68Ga PET/CT and CT imaging.