Determination of Essential and Non-Essential Elements in Medicinal Herbs from Brazil Using ICP OES techniques

João B. Pereira Junior*, Vanessa S. Carvalho, William Q. Ferreira, Rennan G. O. Araujo, Sérgio L.C. Ferreira

Universidade Federal da Bahia, Instituto de Química, Departamento de Química Analítica, Campus Universitário de Ondina, 40170-270, Salvador, Bahia, Brazil

Instituto Nacional de Ciência e Tecnologia, INCT, de Energia e Ambiente, Universidade Federal da Bahia, 40170-115, Salvador, Bahia, Brazil

_______________________________________

*Corresponding author.

E-mail addresses: joaobatista.junior@gmail.com (J. B. Pereira Junior), rennan@ufba.br (R.G.O. Araújo), slcf@ufba.br (S.L.C. Ferreira).

Table S1. Classification of medicinal herbs studied by their common names, scientific names, anacronym, and medicinal properties.

| Common names | Scientific name | Part used | Anacronym | Medicinal property |
| --- | --- | --- | --- | --- |
| Moringa | Moringa oleifera | Leaves | MOR | Used to supply deficiencies such as osteoporosis, and high blood pressure and also in repairing boné fractures caused by hypocalcemia. |
| Chá preto | Camellia sinensis | Leaves | CP | It helps improve digestion, control diabetes, favor weight loss, and prevent diseases such as cancer and heart attack. |
| Chá Branco | Camellia sinensis | Leaves | CB | It is an antioxidant that helps to accelerate metabolism, deflate the body, and reduce the level of lower cholesterol (LDL). |
| Hibisco | Hibiscus sabdariffa | Flowers | HB | It has flavonoids, vitamin C, and organic acids, which help control blood pressure and promote weight loss. |
| Ipê-roxo | Handroanthus impetiginosus | Barks | IR | It has great anti-inflammatory, antibiotic, antitumor, and analgesic action. |
| Dente de Leão | Taraxacum officinale | Leaves | DL | It has antioxidant, anti-inflammatory, and diuretic action. |
| Chapéu-de-couro | Echinodorus grandiflorus | Leaves | CC | It has anti-inflammatory, anti-rheumatic, diuretic, and hypotensive properties, widely used for treating arthritis, urinary infection, or skin problems such as dermatitis. |
| Oliveira | Olea europaea | Leaves | OL | It improves blood sugar levels, protects against cancer, and improves blood pressure. |
| Cavalinha | Equisetum arvense | Leaves | CV | It has diuretic, anti-inflammatory, and digestive action and is indicated for treating anxiety. |
| Camomila | Matricaria chamomilla | Flowers | CM | It has anti-inflammatory and antispasmodic action, controlling the production of acids in the stomach, helping fight nausea, poor digestion, and gas, and preventing gastritis and gastric ulcers. |
| Oolong | Camellia sinensis | Leaves | OG | It contains antioxidants that prevent cancer, diabetes, circulatory problems, coronary heart disease, and degenerative brain disease. |
| Alecrim | Salvia rosmarinus | Leaves and stems | AL | It improves digestion, relieves headaches, and combats fatigue. |
| Amora | Rubus sellowii | Leaves | FA | It has antioxidants that help reduce wrinkles, strengthen collagen, reduce hypertension, help in the treatment of cataracts, as well as helping to reduce anxiety and stress. |
| Carqueja | Baccharis trimera | Leaves | CQ | It regulates blood pressure and blood sugar levels, as well as helps to strengthen the immune system. |
| Pata de vaca | Bauhinia forticata | Leaves | PV | It helps in diuresis and the treatment of diseases of the urinary system. |
| Chá verde | Camellia sinensis | Leaves | CE | It helps to prevent various types of cancer, heart disease, and diabetes, promoting weight loss and improving physical and mental disposition. |
| Fucus | Fucus vesiculosus L. | Seaweed | FC | Protects the brain and improves cognitive functions, to maintain concentration, focus, and reasoning. |

Table S2. Analysis of Variance (ANOVA) for the matrix of 24-1 fractional factorial design used to optimize the acid decomposition procedure

| Parameter | SS | df | MS | F-value | p-value |
| --- | --- | --- | --- | --- | --- |
| (1) [HNO3] (mol L-1) | 3.42 × 10-5 | 1 | 3.42 × 10-5 | 8.32 × 10-4 | 0.979 |
| (2) [H2O2] (% m m-1) | 3.34 × 10-1 | 1 | 3.34 × 10-1 | 8.16 | 0.064 |
| (3) Temperature (°C) | 7.08 × 10-1 | 1 | 7.08 × 10-1 | 17.25 | 0.025 |
| (4) Time (min) | 4.42 × 10-2 | 1 | 4.42 × 10-2 | 1.08 | 0.375 |
| Lack of Fit | 2.05 × 10-1 | 4 | 5.12× 10-2 | 1.25 | 0.446 |
| Pure Error | 1.23 × 10-1 | 3 | 4.10× 10-1 |  |  |
| Total SS | 1.41 | 11 |  |  |  |

SS: sum of square; df: degree of freedom; MS: mean of square.

Table S3. Matrix of 24-1 fractional factorial design proposed, concentrations of RCC and RAC in the acid decomposition procedure proposed, and values of HNO3 regeneration (%) using closed digester block.

| Experiment | [HNO3] (mol L−1)a | [H2O2] (% m m−1)a | Temperature (°C) | Time (min) | RCC (mg L-1) | RAC (mol L-1) | HNO3 regeneration (%) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1.0 | 0.0 | 140 °C | 60 | 79.7 | 0.93 | 93 |
| 2 | 4.0 | 0.0 | 140 °C | 120 | 48.4 | 3.60 | 90 |
| 3 | 1.0 | 6.0 | 140 °C | 120 | 58.1 | 0.96 | 96 |
| 4 | 4.0 | 6.0 | 140 °C | 60 | 12.3 | 3.69 | 92 |
| 5 | 1.0 | 0.0 | 180 °C | 120 | 38.5 | 0.72 | 72 |
| 6 | 4.0 | 0.0 | 180 °C | 60 | 14.5 | 3.15 | 79 |
| 7 | 1.0 | 6.0 | 180 °C | 60 | 11.1 | 0.90 | 90 |
| 8 | 4.0 | 6.0 | 180 °C | 120 | 4.9 | 2.73 | 68 |
| 9 | 2.5 | 3.0 | 160 °C | 90 | 13.3 | 2.19 | 88 |
| 10 | 2.5 | 3.0 | 160 °C | 90 | 10.3 | 2.25 | 90 |
| 11 | 2.5 | 3.0 | 160 °C | 90 | 10.2 | 2.16 | 86 |
| 12 | 2.5 | 3.0 | 160 °C | 90 | 16.8 | 2.22 | 89 |

a Reaction flasks (PTFE) with a final volume of 8.0 mL // RCC = residual carbon concentration //  RAC = residual acid concentration.

Table S4. Performance parameters using the proposed digestion method and detection by ICP OES.

| Elements | Analytical curve equationa | R2 | BEC (mg L-1) | LOD (mg kg-1) | LOQ (mg kg-1) |
| --- | --- | --- | --- | --- | --- |
| As | I = 358.7 [As] – 4.957 | 0.9997 | 0.006 | 0.1 | 0.5 |
| Ca | I = 585.51 [Ca] + 999.1 | 0.9991 | 0.03 | 1.2 | 3.9 |
| Cd | I = 13511 [Cd] + 280.45 | 0.9993 | 0.001 | 0.06 | 0.2 |
| Cr | I = 10904 [Cr] - 8.3754 | 0.9992 | 0.03 | 1.2 | 3.9 |
| Cu | I = 1334.3 [Cu] + 11.201 | 0.9993 | 0.01 | 0.6 | 2.2 |
| Fe | I = 5899.9 [Fe] + 47.945 | 0.9992 | 0.02 | 0.9 | 3.0 |
| K | I = 202.38 [K] - 29.914 | 0.9991 | 0.10 | 0.8 | 2.8 |
| Mg | I = 1027.3 [Mg] + 2267.5 | 0.9985 | 0.02 | 0.9 | 3.0 |
| Mn | I = 47346 [Mn] + 844.56 | 0.9994 | 0.002 | 0.1 | 0.4 |
| Na | I = 18.072 [Na] + 21.752 | 0.9992 | 0.09 | 0.6 | 2.0 |
| P | I = 10.88 [P] + 12.66 | 0.9992 | 0.05 | 1.9 | 6.3 |
| Pb | I = 117.11 [Pb] + 0.5929 | 0.9992 | 0.01 | 0.5 | 1.7 |
| Sr | I = 8257 [Sr] + 159.19 | 0.9992 | 0.01 | 0.6 | 2.0 |
| Zn | I = 12559 [Zn] + 256.22 | 0.9998 | 0.03 | 0.4 | 1.2 |

a I: intensity of emission obtained for each chemical element; [  ] = concentration of each chemical element;

R2 = determination coefficient; BEC = the background equivalent concentration; LOD = limit of detection;

LOQ = limit of    quantification.

Table S5. Values of certified, found, RSD (%), and agreement (%) to three certified reference materials in the multi-elementary determination by ICP OES.

| CRM | Analyte | Certified values | Found values | RSD (%) | Agreement (%) |
| --- | --- | --- | --- | --- | --- |
| Agro C1003a (Tomato leaves) | As (mg kg-1) | 19.02 ± 5.74 | 22.1 ± 0.1 | 0.1 | 116 ± 1 |
|  | Ca (mg kg-1) | 27400 ± 6412 | 27265 ± 20 | 0.1 | 100.0 ± 0.1 |
|  | Cd (mg kg-1) | 26.63 ± 4.47 | 28.3 ± 0.4 | 1.4 | 106 ± 1 |
|  | Cr (mg kg-1) | 6.69 ± 2.56 | 6.48 ± 0.23 | 3.6 | 97 ± 3 |
|  | Cu (mg kg-1) | 1132 ± 237 | 1115 ± 14 | 1.2 | 98 ± 1 |
|  | Fe (mg kg-1) | 1125 ± 361 | 1155 ± 15 | 1.3 | 103 ± 1 |
|  | K (mg kg-1) | 43321 ± 5165 | 46653 ± 354 | 0.8 | 108 ± 1 |
|  | Mg (mg kg-1) | 4112 ± 567 | 4254 ± 100 | 2.3 | 103 ± 2 |
|  | Mn (mg kg-1) | 470.2 ± 108.5 | 456 ± 15 | 3.3 | 97 ± 3 |
|  | Na (mg kg-1) | 2714 ± 1272 | 2618 ± 76 | 2.9 | 96 ± 3 |
|  | P (mg kg-1) | 4371 ± 850 | 4768 ± 111 | 2.3 | 109 ± 3 |
|  | Pb (mg kg-1) | 3.67 ± 1.05 | 3.56 ± 0.03 | 0.9 | 97 ± 1 |
|  | Sr (mg kg-1) | 154 ± 59 | 135 ± 2 | 1.2 | 88 ± 1 |
|  | Zn (mg kg-1) | 37.53 ± 8.29 | 40.6 ± 0.2 | 0.6 | 108 ± 1 |
| Agro C1005a (Sugar cane leaves) | Ca (mg kg-1) | 4378 ± 150 | 4099 ± 89 | 2.2 | 94 ± 2 |
|  | Cu (mg kg-1) | 7.79 ± 1.02 | 7.44 ± 0.27 | 3.6 | 96 ± 3 |
|  | Fe (mg kg-1) | 220 ± 9 | 228 ± 10 | 4.3 | 104 ± 4 |
|  | K (mg kg-1) | 16190 ± 720 | 16202 ± 305 | 2.2 | 100 ± 2 |
|  | Mg (mg kg-1) | 1517 ± 75 | 1319 ± 1 | 0.1 | 86 ± 2 |
|  | Mn (mg kg-1) | 126 ± 6 | 135 ± 5 | 3.4 | 107 ± 4 |
|  | P (mg kg-1) | 2220 ± 98 | 1890 ± 19 | 1.0 | 85 ± 1 |
|  | Sr (mg kg-1) | 25.7 ± 1.6 | 21.8 ± 0.4 | 2.0 | 85 ± 2 |
|  | Zn (mg kg-1) | 13.30 ± 0.87 | 12.8 ± 0.4 | 3.2 | 96 ± 3 |
| NCS DC 73351 (Tea) |  |  |  |  |  |
|  | Ca (%) | 0.43 ± 0.04 | 0.42 ± 0.02 | 5.2 | 96 ± 1 |
|  | Cu (mg kg-1) | 17.3 ± 1.8 | 17.3 ± 0.1 | 0.8 | 100 ± 1 |
|  | Fe (mg kg-1) | 264 ± 15 | 255 ± 5 | 2.0 | 96 ± 2 |
|  | K (%) | 1.66 ± 0.12 | 1.62 ± 0.01 | 1.1 | 98 ± 1 |
|  | Mg (%) | 0.17 ± 0.02 | 0.151 ± 0.001 | 1.1 | 89 ± 1 |
|  | Mn (mg kg-1) | 1240 ± 70 | 1301 ± 15 | 1.1 | 105 ± 1 |
|  | P (mg kg-1) | 2840 ± 90 | 2540 ± 14 | 0.6 | 89 ± 1 |
|  | Pb (mg kg-1) | 4.4 ± 0.3 | 4.1 ± 0.3 | 6.4 | 92 ± 6 |
|  | Sr (mg kg-1) | 15.2 ± 0.7 | 12.6 ± 0.2 | 1.5 | 83 ± 1 |
|  | Zn (mg kg-1) | 26.3 ± 2.0 | 29.8 ± 0.4 | 1.5 | 113 ± 2 |

Results expressed as mean ± standard deviation

Table S6. Concentrations (intraday and interday), intraday RSD, and interday RSD obtained in the multielement determination by ICP OES (n = 3; Ftab = 19.0; df = 2).

| Analyte | Samplesa | Mean intraday (mg kg-1) | Intraday RSD (%) | Mean interday (mg kg-1) | Interday RSD (%) | F-test |
| --- | --- | --- | --- | --- | --- | --- |
| As | A B C D | 22.10 ± 0.17 ND ND 2.74 ± 0.04 | 0.8 ND ND 1.4 | 18.27 ± 0.07 ND ND 2.78 ± 0.11 | 0.4 ND ND 3.8 | 5.90 ND ND 7.56 |
| Ca | A B C D | 27265 ± 20 16464 ± 168 5679 ± 214 13937 ± 254 | 0.1 1.0 3.8 1.8 | 24384 ± 15 15746 ± 123 5606 ± 208 13020 ± 154 | 0.1 0.8 3.7 1.2 | 1.78 1.86 1.06 2.72 |
| Cd | A B C D | 28.35 ± 0.41 ND ND ND | 1.4 ND ND ND | 27.94 ± 0.43 ND ND ND | 1.5 ND ND ND | 1.10 ND ND ND |
| Cr | A B C D | 6.48 ± 0.23 ND ND ND | 3.6 ND ND ND | 6.80 ± 0.38 ND ND ND | 5.5 ND ND ND | 2.73 ND ND ND |
| Cu | A B C D | 1115 ± 14 4.78 ± 0.22 20.34 ± 0.91 8.56 ± 0.36 | 1.2 4.7 4.6 4.3 | 1042 ± 12 4.72 ± 0.22 19.90 ± 0.32 8.04 ± 0.31 | 1.1 4.6 1.6 3.8 | 1.36 1.00 8.09 1.35 |
| Fe | A B C D | 1115 ± 14 323 ± 11 357 ± 19 387 ± 4 | 1.3 3.4 5.4 1.0 | 1099 ± 16 329 ± 16 343 ± 13 368 ± 5 | 1.5 4.8 3.9 1.3 | 1.31 2.11 2.14 1.56 |
| K | A B C D | 46653 ± 354 14120 ± 445 17534 ± 200 9775 ± 33 | 0.8 3.2 1.1 0.3 | 44922 ± 397 14213 ± 234 16244 ± 327 9702 ± 45 | 0.9 1.6 2.0 0.5 | 1.26 3.62 2.67 1.86 |
| Mg | A B C D | 4254 ± 100 8047 ± 203 1626 ± 82 2272 ± 48 | 2.3 2.5 5.1 2.1 | 4174 ± 138 8039 ± 143 1612 ± 27 2247 ± 64 | 3.3 1.8 1.7 2.8 | 1.90 2.01 9.22 1.78 |
| Mn | A B C D | 456 ± 15 38.17 ± 1.96 1228 ± 80 144.9 ± 2.8 | 3.3 5.1 6.5 1.9 | 428 ± 8 35.36 ± 0.85 1133 ± 42 139.9 ± 3.4 | 1.9 2.4 3.7 2.4 | 3.51 5.33 3.63 1.47 |
| Na | A B C D | 2618 ± 76 ND ND ND | 2.9 ND ND ND | 2492 ± 71 ND ND ND | 2.9 ND ND ND | 1.14 ND ND ND |
| P | A B C D | 4768 ± 111 2501 ± 14 2470 ± 145 2754 ± 10 | 2.3 0.6 5.9 0.4 | 4616 ± 115 2561 ± 18 2464 ± 68 2785 ± 25 | 2.5 0.7 2.8 0.9 | 1.07 1.65 4.55 6.25 |
| Pb | A B C D | 3.56 ± 0.03 ND ND ND | 0.9 ND ND ND | 3.46 ± 0.08 ND ND ND | 2.3 ND ND ND | 7.11 ND ND ND |
| Sr | A B C D | 134.9 ± 1.6 104.9 ± 1.3 20.2 ± 0.4 86.7 ± 0.8 | 1.2 1.2 2.0 0.9 | 137.5 ± 2.0 112.3 ± 2.5 19.1 ± 0.3 87.0 ± 0.3 | 1.5 2.2 1.5 0.4 | 1.56 3.70 1.78 7.11 |
| Zn | A B C D | 40.6 ± 0.2 16.4 ± 0.6 32.5 ± 0.8 24.9 ± 1.3 | 0.6 3.6 2.4 5.3 | 39.3 ± 0.5 15.1 ± 0.3 33.5 ± 0.4 23.7 ± 0.9 | 1.4 1.8 1.2 3.7 | 6.25 4.00 4.00 2.09 |

a Sample A: CRM-Agro C1003a; Sample B: MOR; Sample C: CP; Sample D; CB; ND: Not determined.

Results expressed as average ± standard deviation

Table S7. Values of loadings of the variables for the first four principal components.

| Variables | PC1 | PC2 | PC3 | PC4 |
| --- | --- | --- | --- | --- |
| As. | -0.254 | 0.187 | 0.417 | -0.789 |
| Ca. | -0.334 | 0.749 | 0.055 | 0.070 |
| Cu. | 0.534 | -0.455 | 0.053 | -0.259 |
| Fe. | -0.396 | -0.051 | 0.532 | -0.134 |
| K. | -0.569 | -0.594 | 0.343 | 0.116 |
| Mg. | -0.563 | -0.438 | 0.010 | 0.430 |
| Mn. | 0.454 | -0.674 | -0.317 | -0.000 |
| Na. | -0.833 | -0.270 | -0.335 | -0.178 |
| P. | -0.001 | -0.165 | 0.764 | 0.400 |
| Sr. | -0.836 | -0.137 | -0.413 | -0.162 |
| Zn. | 0.171 | -0.570 | 0.167 | -0.392 |
| Variance total (%) | 26.26 | 20.44 | 14.32 | 11.66 |
| Variance cumulative (%) | 26.26 | 46.70 | 61.02 | 72.68 |

Values in bold represent correlations greater than (0.5).