| Parameters | Methods | Instrument | Reference | QA/QC |
| --- | --- | --- | --- | --- |
| Electrical Conductivity | Conductometric Method | Conductivity Meter (Model: Eutech, Con 700) | IS 3025 : 1964 | Samples filtered by 0.45 µm membrane filter; Instrument calibrated with 1413 µS cm-1 KCL solution; Temperature compensation inbuilt in instrument and results reported at 25C |
| pH | Electrometric | pH Meter (Model: Eutech, pH 700) | IS : 3025 (Part II)-1983 (reaffirmed 2002) | Samples filtered by 0.45 µm membrane filter; Instrument calibrated with buffers of 4.0, 7.0 and 9.2 |
| Total Dissolved Solid | Gravimetric Method | Balance with 0.0001 g readability (d) | IS: 3025 (Part 16), 1984 | Desiccated GF/C filter used for filtration; Weighing Balance calibrated with external ASTM Class 1 weights |
| Sodium | Flame Photometry | Flame Photometer (Model: Systronics, Model 128) | IS: 9497-1980 and IS: 8651-1977 (reaffirmed 1999) | Samples filtered by 0.45 µm membrane filter; Instrument calibrated with standards of 0-100 ppm Na in 10 ppm graduation prepared from AR Grade anhydrous NaCl (Merck) dissolved in de-ionized water |
| Potassium | Flame Photometry | Flame Photometer (Model: Systronics, Model 128) | IS: 9497-1980 and IS: 8651-1977 (reaffirmed 1999) | Samples filtered by 0.45 µm membrane filter; Instrument was calibrated with standards of 0-100 ppm K in 10 ppm graduation, prepared from AR Grade anhydrous KCl (Merck) dissolved in de-ionized water |
| Total hardness | Titrimetric Method (EDTA) | Titration Assembly | IS : 3025 (Part 21) | Standard Ca solutions were prepared from dry analytical grade CaCO3 dissolved in deionized water with required addition of HCl for CaCO3 dissolution |
| Calcium | EDTA Titrimetric method | Titration Assembly | ASTM Method D 1126 – 02 | Standard Ca solutions were prepared from dry analytical grade CaCO3 dissolved in deionized water with required addition of HCl for CaCO3dissolution; pH adjusted to 12 to 13 to precipitate Mg before titration for Ca |
| Magnesium | EDTA Titrimetric method | Titration Assembly | ASTM Method D 1126 – 02 | Standard Mg solutions were prepared from dry analytical grade MgO dissolved in 1 N HCl |
| Alkalinity | Titrimetric method | Titration Assembly | IS : 3025 (Part 23) | Sodium thiosulphate was added to remove interference of free available residual chlorine; Samples filtered by 0.45 µm membrane filter |
| Sulphate | Spectrophotometric method | UV-Vis Spectrophotometer (Model: Thermo Orion, Aquamate8000) | USEPA Method 9038 | Samples filtered by 0.45 µm membrane filter; Standards (0-10 ppm SO42-) prepared by dissolved in de-ionized water |
| Nitrate | UV-Spectrophotometric method | UV-Vis Spectrophotometer (Thermo Orion, Aquamate 8000) | USEPA Method 353.2, Revision 2.0 | Samples filtered by 0.45 µm membrane filter; Spectrophotometer calibrated with 0-10 ppm of NO3- solution in 1 ppm graduations, prepared from anhydrous KNO3 |
| Chloride | Argentometric method | Titration Assembly | USEPA METHOD 9253 | Samples filtered by 0.45 µm membrane filter before analysis |
| Iron | 1, 10, Phenanthroline method | UV-Vis Spectrophotometer (Thermo Orion, Aquamate 8000) | IS : 3025 (Part 53) | Samples filtered by 0.45 µm membrane filter and then acidified to <2 pH just after sampling; Spectrophotometer calibrated with 0-1 ppm of Fe solution in 0.1 ppm graduations, prepared from Fe (NH4)2SO4.H2O |
| Fluoride | Ion-selective Electrode Method | Ion meter (Thermo Orion); Fluoride ion-sensitive electrode (Orion 9609BNWP) | USEPA Method 9214 | Samples filtered by 0.45 µm membrane filter; Instrument was calibrated with standard fluoride (TISAB II) solution. |
| Arsenic | Silver Diethyl Dithiocarbamate Method using spectrophotometer | UV-Vis Spectrophotometer (Thermo Orion, Aquamate 8000) | IS : 3025 ( Part 37 ) | Samples filtered by 0.45 µm membrane filter; Procedure followed as per standard SOP of arsenic test kit (Merck MQuant 1.17927.0001) |

Supplementary table 1: Analytical technique and QA/QC