Supplemental Table 4. Information about SDC2 methylation kits.

| Manufacturer | Number of Participants | CpG Sites | LOD（ug/ml） | Judgment Principle |
| --- | --- | --- | --- | --- |
| Creative Biosciences | 109 | 54-62 | 0.02 | CtSDC2 value≤38 and CtACTB values≤36 were methylation positive; CtSDC2 value>38 and CtACTB values≤36 were methylation negative. |
| Life Technology | 15 | 66-76 | 0.33 | CtSDC2 or CtTFPI2 ≤38 were methylation positive; CtSDC2 or CtTFPI2> 38 or Unde were methylation negative. |
| Genomics | 3 | 91-106 | 2.00 | CtSDC2 value≤38, CtACTB values≤37and ΔCT≤9 were methylation positive; both (1) CtSDC2>38 and (2)CtSDC2 value≤38, ΔCT>9 were methylation negative. |
| Realbio Technology | 2 | 43-55 | 0.0022 | Human SFRP 2 and SDC 2 gene methylation combined test kit (fluorescent PCR method) positive judgment value analysis software is used to determine the Score value, methylation positive when Score value is <20; methylationnegative when Score value is 20 |
| AmoyDx | 3 | 53-70 | 90.00 | The △Ct Cut-off value of reaction solution A is 10.0 and the △Ct Cut-off value of reaction solution B is 10.5. The reaction solution A or reaction solution B of any 1 tube △Ct value <△Ct Cut-off value were hypermethylated, otherwise unmethylated. |
| Anhui Targene Medical Technology | 2 | 104-116 | 5 | The risk value is calculated according to the risk value calculation formula determined by statistics, and the cut-off value was set at 5. △Ct < 5 is negative, and △Ct> 5 is positive. |
| Nanjing Vazyme Biotech | 1 | 54-62 | 0.1 | △ Ct (CtSDC2 value - CtReference gene value) less than 6 is hypermethylation, △ Ct >6 is unmethylated. |
| Hangzhou Heyi Gene Technology | 1 | 53-55 | 10 | CtSDC2 value≤40 and CtACTB values≤36 were methylation positive; CtSDC2 value>40 and CtACTB values≤36 were methylation negative. |
| Wuhan YZY MEDICAL Science and Technology | 1 | 1-10 | 2 | CtSDC2 value≤38 and CtACTB values≤36 were methylation positive; CtSDC2 value>38 and CtACTB values≤36 were methylation negative; CtACTB> 36 or no Ct value were considered the sample was not qualified. |
| Renhe Future Biotechnology | 1 | 45-64 | 0.5 | CtACTB values＜38, CtSDC2 value≤38 or CtNDRG4 values≤36 were methylation positive; CtACTB values＜38, CtNDRG4 value >39 or no Ct value and CtSDC2 value> 39 or no Ct value were methylation negative.When the ACTB gene CtACTB>38 or no Ct, the sample was not qualified or detection failed. |

Supplemental Table 5. Compliance rate of different kits for each EQA sample.

| Kit manufacturer | EQA sample 1 | EQA sample 2 | EQA sample 3 | EQA sample 4 | EQA sample 5 | EQA sample 6 | EQA sample 7 | EQA sample 8 | EQA sample 9 | EQA sample 10 | Positive conformance rate | Negative conformance rate |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Creative Biosciences | 109/109 | 109/109 | 109/109 | 109/109 | 109/109 | 109/109 | 109/109 | 107/109 | 109/109 | 109/109 | 870/872(99.77%) | 218/218(100%) |
| Life Technology | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 15/15 | 120/120(100%) | 30/30(100%) |
| Genomics | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 24/24(100%) | 6/6(100%) |
| AmoyDx | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 3/3 | 24/24(100%) | 6/6(100%) |
| Anhui Targene Medical Technology | 2/2 | 1/2 | 2/2 | 2/2 | 2/2 | 2/2 | 1/2 | 1/2 | 1/2 | 2/2 | 13/16(81.25%) | 3/4(75%) |
| Realbio Technology | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 2/2 | 16/16(100%) | 4/4(100%) |
| CEYUPUHUI | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 0/1 | 8/8(100%) | 1/2(50%) |
| Hangzhou Heyi Gene Technology | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 0/1 | 1/1 | 8/8(100%) | 1/2(50%) |
| Jiaxing Yunying Medical Inspection | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 8/8(100%) | 2/2(100%) |
| Nanjing Vazyme Biotech | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 8/8(100%) | 2/2(100%) |
| Wuhan YZY MEDICAL Science and Technology | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 8/8(100%) | 2/2(100%) |
| Renhe Future Biotechnology | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 1/1 | 8/8(100%) | 2/2(100%) |

Supplemental Table 6. The SD and CV of the Ct values for the three replicate samples from each laboratory.

| Laboratory number | CtSDC2 of EQA sample 6 | CtSDC2 of EQA sample 5 | CtSDC2 of EQA sample 4 | Mean of Ct | SD | CV |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 22.18 | 21.41 | 21.06 | 21.55 | 0.57 | 2.66 |
| 2 | 34.36 | 33.81 | 33.89 | 34.02 | 0.29 | 0.86 |
| 3 | 30.84 | 30.97 | 31.35 | 31.05 | 0.27 | 0.85 |
| 4 | 17.95 | 17.81 | 15.25 | 17.00 | 1.52 | 8.94 |
| 5 | 35.44 | 35.45 | 33.83 | 34.91 | 0.93 | 2.67 |
| 6 | 35.81 | 36.05 | 35.26 | 35.71 | 0.41 | 1.13 |
| 7 | 36.06 | 36.09 | 36.10 | 36.08 | 0.02 | 0.06 |
| 8 | 37.05 | 35.73 | 35.52 | 36.10 | 0.83 | 2.30 |
| 9 | 34.77 | 35.02 | 35.59 | 35.13 | 0.42 | 1.20 |
| 10 | 34.82 | 35.07 | 35.26 | 35.05 | 0.22 | 0.63 |
| 11 | 35.23 | 35.35 | 35.66 | 35.41 | 0.22 | 0.63 |
| 12 | 36.65 | 35.70 | 35.72 | 36.02 | 0.54 | 1.51 |
| 13 | 36.35 | 35.96 | 35.36 | 35.89 | 0.50 | 1.39 |
| 14 | 35.40 | 36.05 | 35.10 | 35.52 | 0.49 | 1.37 |
| 15 | 34.29 | 34.57 | 33.72 | 34.19 | 0.43 | 1.27 |
| 16 | 34.29 | 34.57 | 33.72 | 34.19 | 0.43 | 1.27 |
| 17 | 35.09 | 35.79 | 35.60 | 35.49 | 0.36 | 1.02 |
| 18 | 35.21 | 35.89 | 35.38 | 35.49 | 0.35 | 1.00 |
| 19 | 35.10 | 35.66 | 35.10 | 35.29 | 0.32 | 0.92 |
| 20 | 36.52 | 36.31 | 36.94 | 36.59 | 0.32 | 0.88 |
| 21 | 36.37 | 36.07 | 35.80 | 36.08 | 0.29 | 0.79 |
| 22 | 34.92 | 34.83 | 35.31 | 35.02 | 0.26 | 0.73 |
| 23 | 36.45 | 35.94 | 36.25 | 36.21 | 0.26 | 0.71 |
| 24 | 35.50 | 35.18 | 35.11 | 35.26 | 0.21 | 0.59 |
| 25 | 35.15 | 34.98 | 35.31 | 35.15 | 0.17 | 0.47 |
| 26 | 35.23 | 35.16 | 34.92 | 35.10 | 0.16 | 0.46 |
| 27 | 35.87 | 35.61 | 35.73 | 35.74 | 0.13 | 0.36 |
| 28 | 35.31 | 35.13 | 35.07 | 35.17 | 0.12 | 0.36 |
| 29 | 34.70 | 34.70 | 34.50 | 34.63 | 0.12 | 0.33 |
| 30 | 35.12 | 35.35 | 35.26 | 35.24 | 0.12 | 0.33 |
| 31 | 35.79 | 35.84 | 35.64 | 35.76 | 0.10 | 0.29 |
| 32 | 35.20 | 35.13 | 35.09 | 35.14 | 0.06 | 0.16 |
| 33 | 35.25 | 35.17 | 35.15 | 35.19 | 0.05 | 0.15 |
| 34 | 36.03 | 36.04 | 36.01 | 36.03 | 0.02 | 0.04 |
| 35 | 34.79 | 34.98 | 35.39 | 35.05 | 0.31 | 0.87 |
| 36 | 35.72 | 35.78 | 35.46 | 35.65 | 0.17 | 0.48 |
| 37 | 36.93 | 36.68 | 36.30 | 36.64 | 0.32 | 0.87 |
| 38 | 33.60 | 35.12 | 34.87 | 34.53 | 0.82 | 2.36 |
| 39 | 32.99 | 32.74 | 32.71 | 32.81 | 0.15 | 0.47 |
| 40 | 32.88 | 33.57 | 35.61 | 34.02 | 1.42 | 4.17 |
| 41 | 35.81 | 35.73 | 35.59 | 35.71 | 0.11 | 0.31 |
| 42 | 36.88 | 36.36 | 34.69 | 35.98 | 1.14 | 3.18 |
| 43 | 32.89 | 33.84 | 33.97 | 33.57 | 0.59 | 1.76 |
| 44 | 36.11 | 36.43 | 36.98 | 36.51 | 0.44 | 1.21 |
| 45 | 35.44 | 34.67 | 34.99 | 35.03 | 0.39 | 1.10 |
| 46 | 35.47 | 36.28 | 35.62 | 35.79 | 0.43 | 1.20 |
| 47 | 35.56 | 35.26 | 35.98 | 35.60 | 0.36 | 1.02 |
| 48 | 34.97 | 35.57 | 35.47 | 35.34 | 0.32 | 0.91 |
| 49 | 33.37 | 33.37 | 33.44 | 33.39 | 0.04 | 0.12 |
| 50 | 35.98 | 35.98 | 36.58 | 36.18 | 0.34 | 0.95 |
| 51 | 35.76 | 35.39 | 35.31 | 35.49 | 0.24 | 0.68 |
| 52 | 20.35 | 19.52 | 17.70 | 19.19 | 1.36 | 7.06 |
| 53 | 32.95 | 33.81 | 33.52 | 33.43 | 0.44 | 1.31 |
| 54 | 35.62 | 34.44 | 34.56 | 34.87 | 0.65 | 1.86 |
| 55 | 35.71 | 35.56 | 35.21 | 35.49 | 0.26 | 0.72 |
| 56 | 36.53 | 36.70 | 36.83 | 36.69 | 0.15 | 0.41 |
| 57 | 33.80 | 34.30 | 33.69 | 33.93 | 0.33 | 0.96 |
| 58 | 35.30 | 35.61 | 35.56 | 35.49 | 0.17 | 0.47 |
| 59 | 34.77 | 36.69 | 34.62 | 35.36 | 1.15 | 3.26 |
| 60 | 34.93 | 34.66 | 37.31 | 35.63 | 1.46 | 4.09 |
| 61 | 35.86 | 38.78 | 36.93 | 37.19 | 1.48 | 3.97 |
| 62 | 33.17 | 33.43 | 35.02 | 33.87 | 1.00 | 2.96 |
| 63 | 35.17 | 33.50 | 34.79 | 34.49 | 0.87 | 2.53 |
| 64 | 36.44 | 35.44 | 35.65 | 35.84 | 0.53 | 1.47 |
| 65 | 35.35 | 35.70 | 36.36 | 35.80 | 0.51 | 1.43 |
| 66 | 36.07 | 35.13 | 35.82 | 35.67 | 0.49 | 1.36 |
| 67 | 35.69 | 36.03 | 35.11 | 35.61 | 0.47 | 1.31 |
| 68 | 34.44 | 34.05 | 34.93 | 34.47 | 0.44 | 1.28 |
| 69 | 36.78 | 36.78 | 35.97 | 36.51 | 0.47 | 1.28 |
| 70 | 35.68 | 36.33 | 35.46 | 35.82 | 0.45 | 1.26 |
| 71 | 34.09 | 33.60 | 34.44 | 34.04 | 0.42 | 1.24 |
| 72 | 35.73 | 35.97 | 36.55 | 36.08 | 0.42 | 1.17 |
| 73 | 35.54 | 35.18 | 35.79 | 35.50 | 0.31 | 0.86 |
| 74 | 35.07 | 35.27 | 34.75 | 35.03 | 0.26 | 0.75 |
| 75 | 35.21 | 35.58 | 35.71 | 35.50 | 0.26 | 0.73 |
| 76 | 35.31 | 35.72 | 35.30 | 35.44 | 0.24 | 0.68 |
| 77 | 35.53 | 35.24 | 35.06 | 35.28 | 0.24 | 0.67 |
| 78 | 35.37 | 35.76 | 35.80 | 35.64 | 0.24 | 0.67 |
| 79 | 35.36 | 35.69 | 35.81 | 35.62 | 0.23 | 0.65 |
| 80 | 35.56 | 35.69 | 35.99 | 35.75 | 0.22 | 0.62 |
| 81 | 36.30 | 35.90 | 35.96 | 36.05 | 0.22 | 0.60 |
| 82 | 34.43 | 34.80 | 34.72 | 34.65 | 0.19 | 0.56 |
| 83 | 36.16 | 35.79 | 35.99 | 35.98 | 0.19 | 0.51 |
| 84 | 35.18 | 34.88 | 35.14 | 35.07 | 0.16 | 0.46 |
| 85 | 34.31 | 34.55 | 34.61 | 34.49 | 0.16 | 0.46 |
| 86 | 35.32 | 35.06 | 35.08 | 35.15 | 0.14 | 0.41 |
| 87 | 34.93 | 34.67 | 34.70 | 34.77 | 0.14 | 0.41 |
| 88 | 34.93 | 34.67 | 34.70 | 34.77 | 0.14 | 0.41 |
| 89 | 35.58 | 35.63 | 35.36 | 35.52 | 0.14 | 0.40 |
| 90 | 36.03 | 35.90 | 36.19 | 36.04 | 0.15 | 0.40 |
| 91 | 35.88 | 35.61 | 35.71 | 35.73 | 0.14 | 0.38 |
| 92 | 35.47 | 35.26 | 35.34 | 35.36 | 0.11 | 0.30 |
| 93 | 35.41 | 35.20 | 35.30 | 35.30 | 0.11 | 0.30 |
| 94 | 32.24 | 32.42 | 32.36 | 32.34 | 0.09 | 0.28 |
| 95 | 34.45 | 34.55 | 34.49 | 34.50 | 0.05 | 0.15 |
| 96 | 35.93 | 35.88 | 35.86 | 35.89 | 0.04 | 0.10 |
| 97 | 34.80 | 34.84 | 34.84 | 34.83 | 0.02 | 0.07 |
| 98 | 34.91 | 34.31 | 34.25 | 34.49 | 0.36 | 1.06 |
| 99 | 36.39 | 36.80 | 35.83 | 36.34 | 0.49 | 1.34 |
| 100 | 35.27 | 35.72 | 36.00 | 35.66 | 0.37 | 1.03 |
| 101 | 36.13 | 36.07 | 36.00 | 36.07 | 0.07 | 0.18 |
| 102 | 36.20 | 35.84 | 35.73 | 35.92 | 0.25 | 0.69 |
| 103 | 34.73 | 35.23 | 35.96 | 35.31 | 0.62 | 1.75 |
| 104 | 36.15 | 36.48 | 36.25 | 36.29 | 0.17 | 0.47 |
| 105 | 36.03 | 35.90 | 36.19 | 36.04 | 0.15 | 0.40 |
| 106 | 36.20 | 36.94 | 36.59 | 36.58 | 0.37 | 1.01 |
| 107 | 36.39 | 36.09 | 35.89 | 36.12 | 0.25 | 0.70 |
| 108 | 34.71 | 34.86 | 34.98 | 34.85 | 0.14 | 0.39 |
| 109 | 36.30 | 36.10 | 36.70 | 36.37 | 0.31 | 0.84 |
| 110 | 10.00 | 10.00 | 11.00 | 10.33 | 0.58 | 5.59 |
| 111 | 35.60 | 35.32 | 35.16 | 35.36 | 0.22 | 0.63 |
| 112 | 10.88 | 11.47 | 11.37 | 11.24 | 0.32 | 2.81 |
| 113 | 36.45 | 35.44 | 35.38 | 35.76 | 0.60 | 1.68 |
| 114 | 34.76 | 34.59 | 34.44 | 34.60 | 0.16 | 0.46 |
| 115 | 35.76 | 35.44 | 35.64 | 35.61 | 0.16 | 0.45 |
| 116 | 35.74 | 35.79 | 36.26 | 35.93 | 0.29 | 0.80 |
| 117 | 34.11 | 34.06 | 34.41 | 34.19 | 0.19 | 0.55 |
| 118 | 27.84 | 27.92 | 27.80 | 27.85 | 0.06 | 0.22 |
| 119 | 32.80 | 33.28 | 33.68 | 33.25 | 0.44 | 1.32 |
| 120 | 33.28 | 35.52 | 33.74 | 34.18 | 1.18 | 3.46 |
| 121 | 36.60 | 35.70 | 36.00 | 36.10 | 0.46 | 1.27 |
| 122 | 35.56 | 35.03 | 34.82 | 35.14 | 0.38 | 1.09 |
| 123 | 35.16 | 35.63 | 35.10 | 35.30 | 0.29 | 0.82 |
| 124 | 33.38 | 32.89 | 33.35 | 33.21 | 0.27 | 0.83 |
| 125 | 35.44 | 35.45 | 33.83 | 34.91 | 0.93 | 2.67 |
| 126 | 35.76 | 35.15 | 35.22 | 35.38 | 0.33 | 0.94 |
| 127 | 35.05 | 35.64 | 35.03 | 35.24 | 0.35 | 0.98 |
| 128 | 35.17 | 34.82 | 34.81 | 34.93 | 0.21 | 0.59 |
| 129 | 35.17 | 34.82 | 34.81 | 34.93 | 0.21 | 0.59 |
| 130 | 31.75 | 31.16 | 33.92 | 32.28 | 1.45 | 4.50 |
| 131 | 35.52 | 36.21 | 35.44 | 35.72 | 0.42 | 1.19 |
| 132 | 35.45 | 35.48 | 35.60 | 35.51 | 0.08 | 0.22 |
| 133 | 35.82 | 35.56 | 35.47 | 35.62 | 0.18 | 0.51 |
| 134 | 33.26 | 34.06 | 33.06 | 33.46 | 0.53 | 1.58 |
| 135 | 36.77 | 36.67 | 35.92 | 36.45 | 0.46 | 1.27 |
| 136 | 35.78 | 35.61 | 35.31 | 35.57 | 0.24 | 0.67 |
| 137 | 35.55 | 36.23 | 34.95 | 35.58 | 0.64 | 1.80 |
| 138 | 35.41 | 35.36 | 34.38 | 35.05 | 0.58 | 1.66 |
| 139 | 35.82 | 35.91 | 36.29 | 36.01 | 0.25 | 0.69 |
| 140 | 34.82 | 34.65 | 34.73 | 34.73 | 0.09 | 0.24 |