# A Parallel, 2-DOF Exoskeleton for the Human Shoulder: Device Characterization and Preliminary Results on Healthy Subjects

# Supplementary Information

# Fabrication

The spine and modules are primarily composed of a Nylon sheet coated with thermoplastic polyurethane (TPU) on both sides. Nylon is the primary structural component of the parts while TPU facilitates fabrication by providing a layer that can be melted and heal-sealed. Fabrication begins by directly generating 3-D printed structures onto the sheets (Figure S 1-A). The sheets are mounted on a fused deposition modelling (FDM) 3-D printer (Lulzbot, Taz 6). Direct 3-D printing onto the sheets allows for a highly repeatable fabrication process whereas a manual attachment of these structures introduces unwanted variability of the modules. Double-sided adhesive tape is placed on the perimeter of the sheets. The print bed must be heated to a minimum temperature of 60°C in order to ensure proper adhesion of the tape. The height of the nozzle is set to 0.50mm above the surface of the Nylon sheet. The structures are then printed normally using flexible TPU-based filament (Polyflex, Polymaker). The use of TPU-based filament is necessary in order to prevent delamination of the structures from the sheets during operation since the TPU filament bonds exceptionally well to the TPU coat of the sheet. The sheets are then cut based on the desired size of the modules.

At this stage, the spines are complete and are ready for use. Meanwhile, the modules are further processed based on the technique established in previous research (25). A hole is drilled on the under-side of the module that connects the fluid path from the 3-D structures onto the modules. The sides are then heat-sealed, with the structure positioned inside as shown in Figure S 1-B. Paper-based tape is used to cover the sealing area in order to prevent the TPU on the external side from sticking to the sealer. The module is then flipped inside-out; the opening at the top is then heat-sealed to make the module fully hermetic. When deflated, the modules resemble a rectangle with a single seam on the top.

Figure S 1: The fabrication steps are shown. (A) 3-D structures are printed directly onto fabric. (B) The modules are fabricated through heat sealing. After sealing the sides, the module is flipped inside-out

# Exoskeleton Control System

Control of the exoskeleton is primarily achieved by regulating the applied pressure.  Along with the angular position of the actuator, the internal pneumatic pressure is responsible for the magnitude of its torque output. The control system is outlined in Figure S2. An air compressor supplies pressurized air into the entire system. A mass flow controller/sensor (FMA A2317, Omega Engineering) was placed at the outlet of the compressor. The flow controller provides consistency to the exoskeleton’s performance regardless of the state of the compressor. Four DC-controlled pneumatic regulators (ITV-1031, SMC Corp.) regulated compressed air into each actuator. The regulators are also equipped with electronic feedback systems that allow them to relay the magnitudes of their pressure outputs into the microcontroller, facilitating closed-loop control. A microcontroller (Teensy 3.6, PJRC.com LLC) manages the output of the pressure regulators through a digital potentiometer (MCP41010, Microchip Technology Inc.). The pneumatic control system has a bandwidth of 10 Hz, and was configured to have a pressure range of 10-150kPa, with a resolution of 1kPa. Alternatively, it can be vented to the atmosphere, effectively setting system pressure to 0kPa.

Two pushbuttons were added into the circuit. The two buttons must be engaged concurrently in order to operate the exoskeleton. The regulators’ set pressures are immediately changed to 0kPa if any of these buttons are released. During the experiment, one button is given to the subject while the other is held by the experimenter. These two buttons act as failsafe mechanisms ensuring that either the experimenter or the subject is able to safely and immediately disengage the exoskeleton.

Figure S2: A schematic diagram of the utilized control circuitry. Each actuator is supplied with its own regulator for maximum performance. Two fail safe buttons are implemented. Mechanically behaving as pushbuttons, both of these buttons must be pressed down in order for the pressure regulation system to function. Releasing any one of the two buttons immediately depowers the exoskeleton.

Meanwhile, the Teensy 3.6 microcontroller facilitates pressure regulation through a standard Proportional-Integral-Differential controller taking the form of equation [S1]. The primary objective of the controller is to track a predetermined pressure setpoint. By controlling the input pressure, the output torque out can also be varied.

Pout=Pet+Ietdt+Dddte(t)	 [S1]

P, E, and D are the proportional, integral, and differential gains while e(t) is the programmed setpoint. Pout is the output pressure as determined by the equation.

## Control during healthy subject test.

Before the start of each powered set, the experimenter placed the exoskeleton’s control system into a ready state. In this state, the depression of both failsafe buttons engages the control system and begins pressurization of the targeted actuators. The experimenter depresses their failsafe button first; engagement of the exoskeleton is then initiated by the subject by depressing their failsafe button. Exoskeleton movement coincided with the onset of arm motion for the powered phase. For slow movements, the exoskeleton’s pressure setpoint was programmed to increase linearly from 0kPa to 80kPa over a span of 3 seconds; the agonist was left open to atmosphere. This scheme was reversed after 3 seconds. A 350kPa, 500ms, impulse was sent to the exoskeleton during fast movements. While the control system was only active for 500ms, it continued to provide assistance to the subject. At the peak of each movement during the powered phase, a longer pause was implemented in order to ensure that exoskeleton activation and subject movement were concurrently initiated.

# Actuator Experimental Platforms

The actuators were placed across a fixed arm and a movable arm. The movable arm was placed on a swing plate—a device responsible for rotating the arm along the A-A’ axis. The movable arm was then mounted through a pin on the swing plate. This pin served as the center of rotation of the B-B’ angle, which enabled rotation of the arm along the B-B’ axis. Threaded indexing holes on the swing arm indexer and the swing plate locked into position of the swing plate and the movable arm, respectively. The movable arm was bolted onto the swing plate to set its orientation. By rotating the swing plate with respect to the indexer, and the movable arm with respect to the swing plate, the A-A’ and B-B’ angles, respectively, were set. The torque measurement setup was meant to simulate the 2-DOF of the human shoulder – abduction/angle of elevation (AoE) and horizontal flexion/angle of plane of elevation (PoE) (Doorenbosch, Harlaar, & Veeger, 2003). Four load cells (FX1901, TE Connectivity) were mounted and placed in custom-designed holders, between the swing plate and the movable arm; and their outputs were translated to torque by factoring in their distance from the center of rotation. In order to measure torque along the A-A’ axis, 2 load cells were mounted parallel to the swing plate (Figure S3-A). The remaining 2 were mounted perpendicularly and were designed to measure torque along the B-B’ axis (Figure S3-B). Figure S3-C shows an actual photo of the platform during an experiment. In the photo, the B-B’ angle is 30°.

A separate platform was used to quantify the actuator’s step input response and static free bending output. A rigid plastic bar was embedded into the neoprene sheet (Figure S3-E). An 800g weight was attached on the distal end in order to apply resistive torque to the actuator, while 2 IMUS (BN055, Bosch Inc.) were attached to the actuator—one on the posterior end and another on the base of the rigid bar. Measuring the relative position of these IMUs allowed us to determine the bending angle at any given moment. Based on preliminary experimental results, an 800-gram weight was chosen since it provides ~1.4N-m of resistance at 90° and also corresponds to D1’s maximum torque value at 90° and 80kPa when combined with the length of the rigid bar. The weight was removed during the free bending experiments. The actuators were then hung vertically (Figure S3-F). Static measurements were acquired by reading measurements 30s after pressurization in order to ensure steady-state conditions were reached. A 60s period square wave, with peak-to-peak values of 0-80kPa, was used to perform the dynamic experiment and to ensure that sufficient time was provided for the actuators to reach steady state, where each wave starts with a pressurization input.

Figure S3: (A-B) A schematic representation of the static blocked torque measurement platform. The parts are color-coded for easy identification. The rotational axes, A-A’ and B’B are shown. Positive angular rotations are shown. (C) A photo of the measurement platform. (D) A legend sowing the color coding for the schematic representations. (E) A plastic rigid bar (red outline) and a 800g weight is attached to the actuator for step response measurements. (F) For step response and static free response experiments, the actuators are hung vertically.

# Dynamic Response to Step Input

Figure S4-A shows the bending angle response of a weighted actuator. The curves presented have been averaged across the trials. At 0s, the square wave was fed into the actuator. The overall response of the actuators resembled a similar square wave, with an expected time delay. The presence of oscillatory movement was visible in the waveform. A variation in the behavior of the actuator during the course of the experiment was observed and primarily attributed to the industrial compressor engaging its motor due to low pressure in the compressor’s tank. While this affects the inflow into the pressure regulation system, the overall effects of these variations were minimized by taking multiple measurements and the flow sensor/controller. Figure S4-B and Figure S4-C show the averaged inflation and deflation responses, respectively. Variants D1, D2, and D3 have inflation rise times of 4.72s, 2.12s and 3.62s, respectively, and deflation rise times of 3.40s, 4.42s and 1.82s, respectively. While previous designs feature longer deflation times, the current actuator has no discernable difference between deflation and inflation times (Natividad, Del Rosario, Chen, & Yeow, 2018). This change can be attributed to the pneumatic pathing configuration wherein each module has its own pneumatic line connected directly to a regulator that was utilized. Moreover, no discernable correlation was found between module sizing and response times while the step response resembles that of a 1st-order system with no overshoots.

The first order behavior of the actuator highlights its feasibility for use in exoskeletons; higher-order behaviors such as overshoot are unnecessary risks to users since they might apply massive forces or force the arm to a position outside its safe range. Some oscillation in the behavior of the actuator was present during the course of the experiments, which was expected and can be attributed to the actuator’s tendency to buckle when in a free state. Such a problem is not expected during operation of the shoulder sleeve since the actuators are restrained from buckling. The dynamic results presented in (Natividad et al., 2018) indicate that actuator speed is inversely proportional to total volumetric capacity of the installed modules. This behavior is absent in the actuator presented in this chapter and is attributed to the improvements brought forth by the new pneumatic line routing scheme wherein each module is provided its own inlet, as compared to the hybrid actuator (Natividad, Del Rosario, Chen, & Yeow, 2017; Natividad et al., 2018) where the modules are inflated in a sequence.

Figure S4: (A) The waveform in response to a 60s square wave, with 0-80kPa peak-to-peak amplitudes. (B) The averaged step response during inflation. (C) The averaged step response during deflation.

# Data Processing

Data processing was primarily performed using MATLAB’s Sensor Fusion and Tracking Toolbox and Signal Processing Toolbox (MATLAB R2018b, Mathworks). The IMU data from the chest and the right arm acquired were transmitted in quaternion form. The IMU data was pre-processed through equation [S2].

QuatRel= QuatArm* QuatChest-1	[S2]

QuatRel refers to the relative orientation of the arm with respect to the chest. Meanwhile, QuatChest and QuatArm refer the orientation of the IMUs on the chest and on the arm, respectively. QuatRel was translated into three euler angles: angle of the plane of elevation, angle of elevation and the angle of rotation. The converted angles were then mapped using linear interpolation and normalized based on the recorded static IMU values acquired during the beginning of the test.

Raw EMG data was passed through a high pass, infinite impulse response filter with a passband frequency of 20Hz, a transition width of 3Hz with stopband attenuation of 60dB (10th order filter). A similar filter, with a stopband frequency between 48Hz to 52Hz, was used to remove power line interference. The resulting dataset was rectified and outliers, attributed to artifacts, were handled by removing values which exceed 1.5 times the value of the upper 75th quartile. The envelope was then extracted by calculating the root mean square (RMS) of a 500-sample wide moving window. IMU data was used to segment the data sets according to phase and type of movement. Loading and unloading sections of each movement were identified by locating the local maxima and minima of each set. The average data sets were extracted from each repetition in a set. The segmented data set was further analyzed by calculating the RMS values of the angle of the plane of elevation and the angle of elevation, and the EMG data of the relevant muscles. These RMS values were compared in the following pairs: powered against free, powered against unpowered, and unpowered against free. The same pairs were also statistically compared using a paired t-test with a significance level (α) of 0.05

# Anthropomorphic Measurements of Recruited Subjects

Table S 1: Anthropomorphic Measurements of Recruited Subjects

| Sex | 7 Male | 3 Female |
| --- | --- | --- |
| Age | 27±2.9 |  |
| Weight | 63.74 kg ± 10.8 kg |  |
| Height | 168 cm ± 8 cm |  |
| Shoulder Width | 44.4 cm ± 3.7 cm |  |
| Humerus Length | 31.4 cm ± 2.2 cm |  |
| Flexed Bicep Circumference | 30.4 cm ± 3.3 cm |  |
| Relaxed Bicep Circumference | 25.5 cm ± 8.1 cm |  |
| Values reported in mean ± standard deviation except for sex |  |  |

# Tables of changes in RMS of Trajectory

Table S 2: Relative Change in Trajectory RMS Between Phases: Fast

| Loading Motion |  | Angle (o) | PvFa | PvUb | UvFc |
| --- | --- | --- | --- | --- | --- |
| Abduction | Angle of Elevation |  | 0.19(0.02) | 0.14(0.21) | 0.13(0.09) |
| Horizontal Flexion | Angle of Elevation |  | 0.22(0.12) | 0.22(0.25) | 0.24(0.33) |
| Forward Flexion | Angle of Plane of Elevation |  | 0.14(0.03) | 0.09(0.33) | 0.11(0.29) |
| Unloading Motion | Angle (o) |  | PvFa | PvUb | UvFc |
| Adduction | Angle of Elevation |  | 0.14(0.01) | 0.14(0.06) | 0.05(0.12) |
| Horizontal Extension | Angle of Elevation |  | 0.34(0.07) | 0.08(0.02) | 0.29(0.46) |
| Forward Extension | Angle of Plane of Elevation |  | 0.09(0.19) | 0.07(0.00) | 0.04(0.12) |
| (p-values). Statisfically significant results in bold. P-values <0.01 reported as 0.00 a.Powered vs. Free b.Powered vs. Unpowered c.Unpowered vs. Free |  |  |  |  |  |

Table S 3: Relative Change in Trajectory RMS Between Phases: Slow

| Loading Motion |  | Angle (o) | PvFa | PvUb | UvFc |
| --- | --- | --- | --- | --- | --- |
| Abduction | Angle of Elevation |  | 0.19(0.02) | 0.10(0.68) | 0.11(0.10) |
| Horizontal Flexion | Angle of Elevation |  | 0.09(0.75) | 0.11(0.67) | 0.08(0.35) |
| Forward Flexion | Angle of Plane of Elevation |  | 0.52(0.25) | 0.15(0.00) | 0.38(0.22) |
| Unloading Motion | Angle (o) |  | PvFa | PvUb | UvFc |
| Adduction | Angle of Elevation |  | 0.07(0.35) | 0.15(0.17) | 0.09(0.16) |
| Horizontal Extension | Angle of Elevation |  | 0.13(0.38) | 0.08(0.21) | 0.07(0.94) |
| Forward Extension | Angle of Plane of Elevation |  | 0.33(0.46) | 0.23(0.00) | 0.18(0.26) |
| (p-values). Statisfically significant results in bold. P-values <0.01 reported as 0.00 |  |  |  |  |  |
| a.Powered vs. Free b.Powered vs. Unpowered c.Unpowered vs. Free |  |  |  |  |  |

# Tables of changes in RMS of Muscle Activation

Table S 4: Relative Change in Muscle Activation RMS Between Phases: Slow

| Loading Motion |  | Muscle | PvFa | PvUb |  | UvFc |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Abduction | Lateral Deltoid |  | -0.01(0.46) |  | -0.18(0.00) |  | 0.29(0.02) |  |
|  | Infraspinatus |  | -0.01(0.72) |  | -0.10(0.01) |  | 0.10(0.01) |  |
| Horizontal Flexion | Pectoralis Major |  | -0.06(0.96) |  | -0.21(0.03) |  | 0.07(0.05) |  |
|  | Anterior Deltoid |  | -0.05(0.84) |  | -0.14(0.09) |  | 0.10(0.02) |  |
| Forward Flexion | Anterior Deltoid |  | -0.18(0.01) |  | -0.19(0.00) |  | 0.01(0.40) |  |
|  | Infraspinatus |  | -0.05(0.45) |  | -0.09(0.07) |  | 0.06(0.16) |  |
|  | Lateral Deltoid |  | -0.24(0.01) |  | -0.22(0.00) |  | 0.02(0.92) |  |
| Unloading Motion | Muscle |  | PvFa |  | PvUb |  | UvFc |  |
| Adduction | Lateral Deltoid |  | -0.06(0.18) |  | -0.20(0.00) |  | 0.26(0.03) |  |
|  | Infraspinatus |  | -0.02(0.51) |  | -0.12(0.00) |  | 0.11(0.01) |  |
| Horizontal Extension | Pectoralis Major |  | -0.02(0.99) |  | -0.20(0.02) |  | 0.24(0.06) |  |
|  | Posterior Deltoid |  | -0.13(0.01) |  | -0.33(0.00) |  | 0.48(0.05) |  |
| Forward Extension | Anterior Deltoid |  | -0.21(0.01) |  | -0.21(0.20) |  | 0.01(0.19) |  |
|  | Infraspinatus |  | -0.07(0.19) |  | -0.08(0.10) |  | 0.03(0.37) |  |
|  | Lateral Deltoid |  | -0.27(0.01) |  | -0.25(0.01) |  | 0.00(0.81) |  |
| (p-values). Statisfically significant results in bold. P-values <0.01 reported as 0.00 |  |  |  |  |  |  |  |  |
| a.Powered vs. Free b.Powered vs. Unpowered c.Unpowered vs. Free |  |  |  |  |  |  |  |  |

Table S 5: Relative Change in Muscle Activation RMS Between Phases: Fast

| Loading Motion |  | Muscle | PvFa | PvUb |  | UvFc |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Abduction | Lateral Deltoid |  | -0.65(0.00) |  | -0.66(0.00) |  | 0.06(0.09) |  |
|  | Infraspinatus |  | -0.26(0.00) |  | -0.31(0.00) |  | 0.08(0.71) |  |
| Horizontal Flexion | Pectoralis Major |  | -0.34(0.02) |  | -0.45(0.01) |  | 0.29(0.03) |  |
|  | Anterior Deltoid |  | -0.26(0.08) |  | -0.31(0.01) |  | 0.10(0.03) |  |
| Forward Flexion | Anterior Deltoid |  | -0.45(0.00) |  | -0.39(0.00) |  | 0.13(0.06) |  |
|  | Infraspinatus |  | -0.12(0.06) |  | -0.06(0.08) |  | 0.05(0.14) |  |
|  | Lateral Deltoid |  | -0.25(0.25) |  | -0.28(0.10) |  | 0.04(0.64) |  |
| Unloading Motion | Muscle |  | PvFa |  | PvUb |  | UvFc |  |
| Adduction | Lateral Deltoid |  | -0.67(0.00) |  | -0.64(0.00) |  | 0.03(0.70) |  |
|  | Infraspinatus |  | -0.26(0.00) |  | -0.26(0.00) |  | 0.01(0.45) |  |
| Horizontal Extension | Pectoralis Major |  | -0.33(0.02) |  | -0.44(0.01) |  | 0.28(0.04) |  |
|  | Anterior Deltoid |  | -0.12(0.04) |  | -0.36(0.04) |  | 0.91(0.18) |  |
| Forward Extension | Anterior Deltoid |  | -0.25(0.01) |  | -0.35(0.01) |  | -0.17(0.08) |  |
|  | Infraspinatus |  | -0.10(0.08) |  | 0.00(0.11) |  | 0.01(0.27) |  |
|  | Lateral Deltoid |  | -0.22(0.56) |  | -0.21(0.54) |  | 0.00(0.90) |  |
| (p-values). Statisfically significant results in bold. P-values <0.01 reported as 0.00 |  |  |  |  |  |  |  |  |
| a.Powered vs. Free b.Powered vs. Unpowered c.Unpowered vs. Free |  |  |  |  |  |  |  |  |

# References

Doorenbosch, C. A. M., Harlaar, J., & Veeger, D. (H. E. J. . (2003). The globe system: An unambiguous description of shoulder positions in daily life movements. The Journal of Rehabilitation Research and Development, 40(2), 149. https://doi.org/10.1682/JRRD.2003.03.0149

Natividad, R. F., Del Rosario, M. R. J., Chen, P. C. Y., & Yeow, R. C. H. (2017). A Hybrid Plastic-Fabric Soft Bending Actuator with Reconfigurable Bending Profiles. In Proceedings - IEEE International Conference on Robotics and Automation (pp. 6700–6705). Singapore: IEEE. https://doi.org/10.1109/ICRA.2017.7989792

Natividad, R. F., Del Rosario, M. R. J., Chen, P. C. Y., & Yeow, R. C. H. (2018). A Reconfigurable Pneumatic Bending Actuator with Replaceable Inflation Modules. Soft Robotics, 5(3), 304–317. https://doi.org/10.1089/soro.2017.0064