# Supplementary Material

A novel lab-on-chip cartridge and AI tools for automated urinalysis

Avinash S, Srinivasan K*, Dhanu V,  Eshwarmurthy T, Vaibhav B, Sharitha M N and T. Rai Dastidar

SigTuple Technologies Pvt. Ltd, Bengaluru, Karnataka 560102, India

*srinivasan@sigtuple.com

Supplementary Table S1  Sedimentation time study for the urine cells and particles in GSC.

| S. No. | Cartridge | Time (min) | Observation (layer) | RBC | WBC | Bacteria | Cast | Crystals | Yeast | Epithelial Cells |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | WM | 0 | Single | 4+ | NS | NS | ND | ND | ND | D |
|  | GSC | 1 | Multi | NA |  |  |  |  |  |  |
|  |  | 2 | Multi | NA |  |  |  |  |  |  |
|  |  | 3 | Multi | NA |  |  |  |  |  |  |
|  |  | 4 | Multi | NA |  |  |  |  |  |  |
|  |  | 5 | Enriched | 4+ | NS | NS | ND | ND | ND | D |
|  |  | 7 | Enriched | 4+ | NS | NS | ND | ND | ND | D |
|  |  | 10 | Enriched | 4+ | NS | NS | ND | ND | ND | D |

Supplementary Table S2 Bucket-wise grading of urine cells and particles

Qualitative assessment

| Epithelial, Crystal, Cast & Yeast | Detected (D) or Not Detected (ND) |
| --- | --- |

Quantitative assessment

| Grade | RBC/WBC (cells per FOV) | Bacteria (cells per FOV) |
| --- | --- | --- |
| Non Significant (NS) | 0-5 | 0-1 |
| 1+ | 6-10 | 1-2 |
| 2+ | 11-20 | 3-5 |
| 3+ | 21-50 | 6-10 |
| 4+ | >50 | >10 |

Supplementary Table S3 Comparative study of cell count between GSC (manual count) and wet mount

| S.No. | Cartridge | RBC | WBC | Bacteria | Cast | Crystals | Yeast | Epithelial Cells |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | WM | 1+ | 2+ | 4+ | D | ND | ND | D |
|  | GSC | 1+ | 2+ | 4+ | D | ND | ND | D |
| 2 | WM | 2+ | NS | 4+ | ND | D | ND | D |
|  | GSC | 2+ | NS | 4+ | ND | D | ND | D |
| 3 | WM | 3+ | NS | 1+ | ND | ND | D | D |
|  | GSC | 3+ | NS | 1+ | ND | ND | D | D |

Supplementary Table S4 Cells and particle count of raw urine sample in GSC at a different time interval.

| Time (min) | Cartridge | Chamber | RBC | WBC | Bacteria | Cast | Crystals | Yeast | Epithelial Cells |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | WM | - | 3+ | NS | 1+ | ND | ND | D | D |
|  | GSC | 1 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 2 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 3 | 3+ | NS | 1+ | ND | ND | D | D |
| 30 | GSC | 1 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 2 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 3 | 3+ | NS | 1+ | ND | ND | D | D |
| 60 | GSC | 1 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 2 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 3 | 3+ | NS | 1+ | ND | ND | D | D |
| 90 | GSC | 1 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 2 | 3+ | NS | 1+ | ND | ND | D | D |
|  |  | 3 | 3+ | NS | 1+ | ND | ND | D | D |

Supplementary Figure S1 Images of urine cells and particles in GSC captured using AI100 device (a) Bacteria, calcium oxalate and uric acid crystals (b) RBC (c) Yeast, WBC (d) Epithelial cells

The definition of analytical parameters, namely sensitivity, specificity, precision and accuracy, are

(1)

(2)

(3)

(4)

Where TP, True positive; TN, True negative; FN, False negative; FP, False positive.

True Positives and True Negatives are instances of a match between WM and the GSC data (manual microscopy, tele reporting and AI model). For example, if both WM and GSC report epithelial as detected, it's a true positive; if both report it as not detected, then it's a true negative.

False positives and False negatives are defined as instances where there is a mismatch between the WM and the GSC data (manual microscopy, tele reporting and AI model). For example, if WM reports epithelial as not detected and GSC reports epithelial as detected, then it is a false positive, whereas if wet mount reports it as detected and GSC reports it as not detected, then it's a false negative. For quantitative classes, we consider cells >5 cells / FOV as detected and <=5 cells / FOV as not detected to calculate performance metrics.