Dr. Carlos Rossa

Editor-in-Chief,
BMC Research Notes

March 8h, 2021

Re: "The Cupola: an additional layer of protection for providers working in the oropharyngeal region" (RESN-D-21-00030)

Dear Dr. Rossa,

We are thankful for the reviewers’ thoughtful comments. We have considered them carefully and revised the manuscript accordingly. Our point-by-point responses are given below.

Please find enclosed the revised manuscript with “track changes”.  Please, let me know if you have any questions.

Thank you again for your support.

Sincerely,

Alessandro Villa

Comments for the Author:

Editor Comments:

Dear Dr Villa,

After revising your manuscript, I think it conforms to the requirements of BMC Research Notes, providing technically sound, scientifically valid information with appropriate conclusions. Moreover, the manuscript does address a relevant gap in knowledge. Please find the comments of one of the reviewer's below for your consideration. However, I do emphasize that the central aspect of the revision should be to provide further information/clarification on the procedure used to generate the aerosol. Specifically, was the high speed hand piece held close to a flat surface or was it used inside a mannequin's mouth? (on page 5, line 42 it is stated that "...the airborne particle counter was positioned pre-set at head height (50 cm above the simulate patient's head"); however one cannot see if there is a mannequin's head on Figure 1B.

Response: thank you for your comment. The high speed was held close to the dental chair (no mannequin) at a distance of 50 cm above the dental chair headrest. We clarified this in the text.

Also, in the 'limitations' sub-section, it may be interesting to make the readers' aware that, when possible, in dental procedures with a high speed hand piece, using a rubber dam isolation and/or high power vacuum suction may markedly reduce the spread of virus-contaminated particles.

Response: We added the following sentence to the limitation section as recommended: “In addition, during most aerosol generating dental procedures with a high speed hand piece, dentists typically use a rubber dam isolation and/or high power vacuum suction which may also markedly reduce the spread of virus-contaminated particles.”

Reviewer reports:

Reviewer 2: Abstract

- In general, the summary was confused and does not sufficiently demonstrate the objective and method of the study. Need to be redone and more detailed.

- The objective needs to be more direct:`` to assess the effectiveness of an additional layer of protection developed for providers working in the oropharyngeal region''.

- There is a lack of information on the type of study, on the stages of development and testing. The establishment of the comparison group. The statistical analysis employed.

Response: thank you for suggestions. We reworded the objective section as recommended and included the study type (pilot study). We did not expand on the methods in the abstract to comply with the journal’s guidelines (Abstract: structured into two sections: Objective and Results description. State the purpose of the research presented and give a brief summary of the main findings).

Introduction

- SARS-CoV-2: put it out in the first use

Response: the suggested revision was made.

- Although the aspect of COVID-19 is relevant today, there are other pathogens that need additional care, and therefore, the introduction could clarify that even in the routine of this work, professionals were and will continue to be exposed to contamination. Which makes the developed material even more relevant.

Response: thanks for the suggestions. We added a sentence in the conclusion on this (“Even after the COVID-19 pandemic, professionals will continue to be exposed to possible contamination.”)

- Reference: ``Coronavirus disease (or COVID-19) cases continue to climb in the United States and around the world.``.

Response: a reference was added and we updated the statistics.

- The objective here is closer to that suggested in the summary. Either way, they need to be the same.

Response: the suggested revision was made.

Main text

- The introduction must be inserted in here, as a section.

Response: the suggested revision was made.

- The text does not achieve the principle of reproducibility: Why did you choose this method? When was it done? Where? Who did? Who measured?

Response: we included that “All experiments were carried out by one of the investigators (A.V.) in a dental office”; the methodology was used following previous experiments as reported in references 11 and 12.

- There is a lack of information on the type of study.

Response: we added that this was a pilot study as requested by the reviewer.

- I believe that the initial part of the text should bring about the motivation stage and the development stages.

- There is a lack of clarity from a broader perspective on what is later put on as a form of evaluation.

- The analysis of the data could be a separate section.

Response: we added a subtitle as suggested

- The results need to be a separate section.

Response: we added a subtitle as suggested

-There is a lack of discussion of the data with the literature. How it contributes. What is the cost ratio? For other oropharyngeal professionals, what would adaptation be like in the chair?

- The conclusion must respond to the objective directly.

Response:  We expanded on this and included two new sentences in the discussion (“Several porotypes have been developed to include an articulating arm with a clamp for attaching to a headrest or backrest of a dental chair and operating bed.”)

A cost analysis was not performed and will be part of a secondary paper. The focus of this pilot study was to assess the effectiveness of an additional layer of protection developed for providers working in the oropharyngeal region'.