Table 1: Materials properties for different parts of prosthetic leg

| Model Reference | Properties |
| --- | --- |
| Ankle Assembly Parts |  |
|  | Name:; Ti-5Al-2.5Sn Annealed (SS). Model type:; Linear Elastic Isotropic. Default failure:; Max von Mises Stress. Yield strength:; 7.44634e+008 N/m^2. Tensile strength:; 8.61e+008 N/m^2. Compressive strength:; 8.3e+008 N/m^2. Elastic modulus:; 1.103e+011 N/m^2. Poisson's ratio:; 0.31. Mass density:; 4480 kg/m^3. Shear modulus:; 4.8e+010 N/m^2. Thermal expansion:; 9e-006 /Kelvin. |
|  | Name:; 2024 Alloy (SN). Model type:; Linear Elastic Isotropic. Default failure:; Max von Mises Stress. Yield strength:; 7.58291e+007 N/m^2. Tensile strength:; 1.86126e+008 N/m^2. Elastic modulus:; 7.3e+010 N/m^2. Poisson's ratio:; 0.33. Mass density:; 2800 kg/m^3. Shear modulus:; 2.8e+010 N/m^2. Thermal expansion:; 2.3e-005 /Kelvin. |
|  | Name:; 3003-H14, Rod (SS). Model type:; Linear Elastic Isotropic. Default failure:; Max von Mises Stress. Yield strength:; 1.45e+008 N/m^2. Tensile strength:; 1.5e+008 N/m^2. Elastic modulus:; 6.9e+010 N/m^2. Poisson's ratio:; 0.33. Mass density:; 2730 kg/m^3. Shear modulus:; 2.5e+010 N/m^2. Thermal expansion:; 2.3e-005 /Kelvin. |
| Pneumatic assembly parts |  |
|  | Name:; Alumina. Model type:; Linear Elastic Isotropic. Tensile strength:; 3e+008 N/m^2. Compressive strength:; 3e+009 N/m^2. Elastic modulus:; 3.7e+011 N/m^2. Poisson's ratio:; 0.22. Mass density:; 3960 kg/m^3. Shear modulus:; 1.5e+011 N/m^2. Thermal expansion:; 7.4e-006 /Kelvin. |
|  | Name:; Titanium Ti-8Al-1Mo-1V. Model type:; Linear Elastic Isotropic. Default failure:; Max von Mises Stress. Yield strength:; 9.1e+008 N/m^2. Tensile strength:; 9.37e+008 N/m^2. Elastic modulus:; 1.2e+011 N/m^2. Poisson's ratio:; 0.32. Mass density:; 4370 kg/m^3. Shear modulus:; 4.6e+010 N/m^2. Thermal expansion:; 8.5e-006 /Kelvin. |
| Limb shell |  |
|  | Name:; PE High Density. Model type:; Linear Elastic Isotropic. Tensile strength:; 2.21e+007 N/m^2. Elastic modulus:; 1.07e+009 N/m^2. Poisson's ratio:; 0.4101. Mass density:; 952 kg/m^3. Shear modulus:; 3.772e+008 N/m^2. |

Table 2: volumetric properties of ankle assembly parts

| Document Name and Reference | Treated As | Volumetric Properties |
| --- | --- | --- |
| Fillet6 | Solid Body | Mass:0.383111 kg Volume:8.55158e-005 m3 Density:4480 kg/m3 Weight:3.75449 N |
| Cut-Extrude7 | Solid Body | Mass:0.465899 kg Volume:0.000103995 m3 Density:4480 kg/m3 Weight:4.56581 N |
| Fillet1 | Solid Body | Mass:0.0533369 kg Volume:1.90489e-005 m3 Density:2800 kg/m3 Weight:0.522702 N |
| Revolve1 | Solid Body | Mass:0.362359 kg Volume:0.000132732 m3 Density:2730 kg/m3 Weight:3.55112 N |

Table 3: Load and fixtures for ankle part

| Fixture Name | Fixture Image | Fixture Details |
| --- | --- | --- |
| Fixed-1 |  | Entities: 1 face(S) Type: Fixed Geometry Resultant Forces: Components; X; Y; Z; Resultant. Reaction force(N); -0.113014; -200.049; 0.144142; 200.049. |

| Load Name | Load Image | Load Details |
| --- | --- | --- |
| Force-1 |  | Entities:; 2 face(s). Type:; Apply normal force. Value:; 500 N. |
| Force-2 |  | Entities:; 2 face(s). Type:; Apply normal force. Value:; 500 N. |

Table 4: Von-mises stress analysis for ankle assembly part

| Type | Min | Max |
| --- | --- | --- |
| VON: von Mises Stress | 3.776e-004N/mm2 (MPa) | 1.536e+002N/mm2 (MPa) |

Table 5: Equivalent strain analysis for ankle assembly part

| Type | Min | Max |
| --- | --- | --- |
| ESTRN: Equivalent Strain | 1.611e-008 | 8.726e-004 |

Table 6: resultant displacement analysis for ankle assembly part

| Type | Min | Max |
| --- | --- | --- |
| URES: Resultant Displacement | 0.000e+000mm | 6.904e+000mm |

Table 7: factor of safety for ankle assemble part

| Type | Min | Max |
| --- | --- | --- |
| Factor of Safety | 4.848e+000 | 1.972e+006 |

Table 8: volumetric properties of pneumatic shaft assembly

| Document Name and Reference | Treated As | Volumetric Properties |
| --- | --- | --- |
| Fillet1 | Solid Body | Mass:0.527262 kg Volume:0.000133147 m3 Density:3960 kg/m3 Weight:5.16717 N |
| Ø5.5 (5.5) Diameter Hole1 | Solid Body | Mass:0.0776634 kg Volume:1.77719e-005 m3 Density:4370 kg/m3 Weight:0.761102 N |

Table 9: load and fixtures for pneumatic shaft assembly

| Fixture Name | Fixture Image | Fixture Details |
| --- | --- | --- |
| Fixed-1 |  | Entities: 1 face(s) Type: Fixed Geometry Resultant Forces: Components; X; Y; Z; Resultant. Reaction force(N); 0.00402164; 4806.24; 0.142075; 4806.24. |
| Load Name | Load Image | Load Details |
| Pressure-1 |  | Entities:; 1 face(s). Reference:; Edge< 1 >. Type:; Along Edge. Value:; -10. Units:; N/mm2 (MPa). Phase Angle:; 0. Units:; deg. |

Table 10: von-mises stress for pneumatic shaft assembly

| Type | Min | Max |
| --- | --- | --- |
| VON: von Mises Stress | 4.123e-007N/mm2 (MPa) | 7.039e+001N/mm2 (MPa) |

Table 11: equivalent strain for pneumatic shaft assembly

| Type | Min | Max |
| --- | --- | --- |
| ESTRN: Equivalent Strain | 1.058e-012 | 4.565e-004 |

Table 12: resultant displacement for pneumatic shaft assembly

| Type | Min | Max |
| --- | --- | --- |
| URES: Resultant Displacement | 0.000e+000mm | 2.859e-002mm |

Table 13: factor of safety for pneumatic shaft assembly

| Type | Min | Max |
| --- | --- | --- |
| Factor of Safety | 1.000e+001 | 6.161e+011 |

Table 14: volumetric properties of limb shell

| Document Name and Reference | Treated As | Volumetric Properties |
| --- | --- | --- |
| Fillet4 | Solid Body | Mass:0.373722 kg Volume:0.000392565 m^3 Density:952 kg/m^3 Weight:3.66247 N |
| Shell Bodies |  |  |
| Surface-3 | Thin | Thickness:2 mm Weight:0.0120009 N Volume:1.28632e-006 m^3 Mass:0.00122458 kg Density:952kg/m^3 |

Table 15: fixture details of limb shell

| Fixture Name | Fixture Image | Fixture Details |
| --- | --- | --- |
| Fixed-1 |  | Entities: 4edge(s), 3 face(s) Type: Fixed Geometry Resultant Forces: Components; X; Y; Z; Resultant. Reaction force(N); -1.83645e-008; 3.67619; -7.71252e-009; 3.67619. |