Additional Files

File name and format: Additional File 1.docx

Title: Additional File 1. Current Biomarker Tests and Algorithms for Liver Disease

Description: Table listing the current biomarkers and algorithms and their performance for diagnosis of liver disease.

File name and format: Additional File 3.docx

Title: Additional File 2. GP73 muhuFab/CHO Conjugate Stability: 30°C X 14 Day & 37°C X 14 Day comparing IR (incorporation ratio of acridinium per conjugate antibody) range of conjugates.

Description: Table showing conjugate stability under various temperature and time conditions

File name and format: Additional File 2.docx

Title: Additional File 3. Microparticle Stability: 45°C X 3 Days

Description: Table showing microparticle stability under various temperature and time conditions.

Additional File 1. Current Biomarker Tests and Algorithms for Liver Disease

| Tests/Algorithms | Limitations | Performance |
| --- | --- | --- |
| Liver biopsy/pathology | - Invasive procedure - High risk of sampling error - Risk of cancer spread along needle path | Gold standard |
| FIB-4 (platelet count, AST, ALT, age) | - Indirect measurement of fibrosis - Not prognostic of progression | FIB-4 AUC 0.751 vs GP73 AUC 0.898 (by HotGen ELISA) for fibrosis ≥2 [15] |
| Enhanced liver fibrosis (ELF) test to rule-out patients for biopsy (hyaluronic acid, PIIINP, TIMP1) | - NICE-recommended based on Peds studies | AUC 0.811 for advanced fibrosis and cirrhosis (N=19,285 patients with considerable variability [36] |
| Aspartate Platelet Ratio Index (APRI) (AST, platelet count) | - More informative for HCV than other liver diseases | APRI AUC 0.737 vs GP73 AUC 0.898 (by HotGen ELISA) for fibrosis ≥2 [15] |
| Fibrotest (CE)/FibroSure (US) (α2-macroglobulin, haptoglobin, GGT, gamma-globulin, total bilirubin, apolipoprotein A1) | - Specificity issues (several biomarkers are not exclusive to liver) | AUC 0.74 for significant fibrosis in chronic HCV [37] Suboptimal for detection of fibrosis with chronic HBV [38] |
| Fibrometer (age, weight, platelet count, AST, ALT, ferritin, glucose) | - Indirect measurement of fibrosis - Poor prognosis of progression | AUCs 0.75-0.80 for advanced fibrosis in NAFLD [39] AUC 0.81 for autoimmune liver disease [40] |
| Fibroscan to identify need for biopsy (ultrasound imaging) | - Challenging in patients with ascites, obesity - High sampling error - High cost/low accessibility | AUC 0.85 for fibrosis >3 in NAFLD [41] |
| MRE to measure liver stiffness and assess disease stage (magnetic resonance elastography imaging) | - High cost/low accessibility | AUC 0.841 for advanced NASH and NAFLD [42] |

AUC, area under the ROC curve; ELISA, enzyme-linked immunosorbent assay; GGT, gamma-glutamyltransferase; NAFLD, non-alcoholic fatty liver disease; NASH, non-alcoholic steatohepatitis; NICE, National Institute for Health and Care Excellence; PIINP, type III procollagen peptide; TIMP1, tissue inhibitor of metalloproteinase-1.

Additional File 2. GP73 muhuFab/CHO Conjugate Stability: 30°C X 14 Day & 37°C X 14 Day comparing IR (incorporation ratio of acridinium per conjugate antibody) for a range of conjugates. GP73 conjugates (Abbott Diagnostics China R&D) showed acceptable RLU stability with various IR with less than 20% average RLU loss.

Stability is not significantly impacted by IR.

| Test Condition | Conjugate | %Conc. Difference: Stressed vs. Non-Stressed |  |  | Ave Panels %Conc. Difference |
| --- | --- | --- | --- | --- | --- |
|  |  | Panel L | Panel M | Panel H |  |
| 30°C, 14DAY | IR 0.98 | -2.7% | -0.7% | 0.1% | -1.1% |
|  | IR 1.87 | -0.5% | 0.8% | -2.6% | -0.8% |
|  | IR 2.73 | -4.7% | -1.8% | -1.7% | -2.7% |
|  | IR 3.45 | -2.5% | 0.4% | -3.3% | -1.8% |
|  | IR 4.14 | -3.7% | -0.2% | 2.8% | -0.4% |
|  | IR 5.02 | -3.2% | 3.8% | -1.5% | -0.3% |
|  | IR 5.85 | -2.6% | 1.7% | -3.5% | -1.5% |
|  | LC* IR 4.12 | -5.2% | -1.5% | 0.2% | -2.2% |
| 37°C, 14DAY | IR 0.98 | -0.2% | -2.0% | 3.2% | 0.3% |
|  | IR 1.87 | 0.2% | 0.5% | -4.6% | -1.3% |
|  | IR 2.73 | -3.3% | 0.2% | -0.6% | -1.2% |
|  | IR 3.45 | -0.6% | 3.2% | 2.7% | 1.8% |
|  | IR 4.14 | -2.5% | 0.2% | 6.4% | 1.4% |
|  | IR 5.02 | -1.5% | 1.0% | 0.5% | 0.0% |
|  | IR 5.85 | -5.7% | -0.9% | 2.5% | -1.4% |
|  | LC* IR 4.12 | 0.2% | -1.9% | 1.6% | 0.0% |

*Abbott Lake County, Abbott Park, IL, USA (LC).

Additional File 3. Microparticle Stability: 45°C X 3 Days. GP73 IgG antibody coated on magnetic microparticles is very stable with RLU loss ≤3.5% at 45°C X 3 days for both calibrators and panels.

| Sample ID (5 reps) | Target (ng/mL) | RLU %Difference 45°C/3day vs. 2-8°C | Conc. %Difference 45°C/3day vs. 2-8°C |
| --- | --- | --- | --- |
| GP73 Cal A | 0 | -13.3% | 95.2% |
| GP73 Cal B | 3.9 | -1.4% | 3.0% |
| GP73 Cal C | 15.6 | -3.9% | 1.5% |
| GP73 Cal D | 62.5 | -2.8% | 1.7% |
| GP73 Cal E | 250 | 0.1% | 3.3% |
| GP73 Cal F | 1000 | -2.3% | -3.1% |
| Ave Cal B-F RLU% Difference |  | -2.0% | 1.3% |
| Panel L | 36 | -5.9% | -1.2% |
| Panel M | 125 | -0.9% | 3.1% |
| Panel H | 500 | -3.7% | -2.5% |
| Ave Panels RLU% Difference |  | -3.5% | -0.2% |