Supplementary Materials

An UHPLC-QE-Orbitrap-MS method for accurate quantification of short-chain fatty acids in serum from an older Chinese population

Sheng Wang a, b, 1, Rui Feng b, c, 1, Li Kong d, Rui Zhou c, Fang-ting Hu d, Shu-Jing Sun d, Guan-Jun Chen a, b, 1, Fang-Biao Tao d, e, f, Kai-Yong Liu d, e, f, *

a Center for Scientific Research, Anhui Medical University, No. 81 Meishan Road, Hefei 230032, Anhui, China

b Inflammation and Immune Mediated Diseases Laboratory of Anhui Province, No. 81 Meishan Road, Hefei 230032, Anhui, China

c School of Pharmacy, Anhui Medical University, No. 81 Meishan Road, Hefei 230032, Anhui, China

d School of Public Health, Anhui Medical University, No. 81 Meishan Road, Hefei 230032, Anhui, China

e MOE Key Laboratory of Population Health Across Life Cycle, No. 81 Meishan Road, Hefei 230032, Anhui, China

f Anhui Provincial Key Laboratory of Population Health and Aristogenics, No. 81 Meishan Road, Hefei 230032, Anhui, China

1 Co-first author

* Corresponding authors: Kai-Yong Liu (kyliu@ahmu.edu.cn)

Table and figure of Contents for Supporting Information

SUPPORTING FIGURES AND TABLES

Fig. S1. The schematic of sample preparation process.

Table S1 Information of the analytes for 11 SCFAs and the internal standard.

Table S2 Elution gradient of 11 SCFAs detection.

Table S3 The raw data of calibration and validation tests for 11 SCFAs.

Table S4 Intra- and inter-batches precisions and accuracies of the 11 SCFAs targets.

Table S5 The raw data of the intra- and inter-day precisions and accuracies of 11 SCFAs.

Table S6 Stability of 11 SCFAs for method validation.

Table S7 Concentrations of 11 SCFAs in 1028 Chinese elderly serum samples.

Fig. S1. The schematic of sample preparation process.

Table S1 Information of the analytes for 11 SCFAs and the internal standard.

| NO. | Name | Forum | MW of Derivatives | Purity (%) | Band | CAS NO. |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | AA | C2H4O2 | 194.05711 | ≥98 | Dr.Ehrenstorfer GmbH | 64-19-7 |
| 2 | PA | C3H6O2 | 208.07276 | ≥98 | Dr.Ehrenstorfer GmbH | 79-09-4 |
| 3 | BHB | C4H6O2 | 220.07276 | ≥99.9 | Shanghai Macklin Biochemical Co., Ltd | 107-93-7 |
| 4 | BA | C4H8O2 | 222.08841 | ≥98 | Dr.Ehrenstorfer GmbH | 79-09-4 |
| 5 | Iso-BA | C4H8O2 | 222.08841 | ≥98 | Dr.Ehrenstorfer GmbH | 79-31-2 |
| 6 | LA | C3H6O3 | 224.06768 | ≥98 | Chengdu Huaxia Chemical Reagent Co., Ltd | 79-33-4 |
| 7 | VA | C5H10O2 | 236.10406 | ≥98 | Dr.Ehrenstorfer GmbH | 109-52-4 |
| 8 | Iso-VA | C5H10O2 | 236.10406 | ≥98 | Dr.Ehrenstorfer GmbH | 503-74-2 |
| 9 | BD | C4H8O3 | 238.08333 | ≥98 | HPC Standards Gmbh | 565-70-8 |
| 10 | CA | C6H12O2 | 250.11971 | ≥98 | Dr.Ehrenstorfer GmbH | 142-62-1 |
| 11 | Iso-CA | C6H12O2 | 250.11971 | ≥99 | shanghai Yuanye Biotechnology Co., Ltd | 646-07-1 |
| 12 | BA-d7 | C4HD7O2 | 229.13235 | ≥98 | Cayman | 73607-83-7 |

Note: lactic acid (LA); β-hydroxybutyrat (BHB); acetic acid (AA); propionic acid (PA); crotonic acid (BD-CA); isobutyric acid (ISO-BA); butyric acid (BA); isovaleric acid (ISO-VA); valeric acid (VA); isocaproic acid (ISO-CA); caproic acid (CA).

Table S2 Elution gradient of 11 SCFAs detection.

| Time（min） | Flow（mL/min） | %A | %B |
| --- | --- | --- | --- |
| 0.00 | 0.35 | 95 | 5 |
| 0.50 | 0.35 | 95 | 5 |
| 1.70 | 0.35 | 85 | 15 |
| 3.50 | 0.35 | 85 | 15 |
| 8.00 | 0.35 | 75 | 25 |
| 16.00 | 0.35 | 60 | 40 |
| 16.01 | 0.50 | 10 | 90 |
| 18.00 | 0.50 | 10 | 90 |
| 18.10 | 0.35 | 95 | 5 |
| 20.00 | 0.35 | 95 | 5 |

Table S3 The raw data of calibration and validation tests for 11 SCFAs .

| Analyte | Actual Conc. | Area | IS Area | Area Ratio | Calculated Con. | Accuracy (%) |
| --- | --- | --- | --- | --- | --- | --- |
| LA | 5.00 | 24529694 | 646655 | 37.93 | 4.08 | 81.65 |
|  | 10.00 | 48802146 | 919176 | 53.09 | 9.96 | 99.59 |
|  | 20.00 | 73969545 | 863394 | 85.67 | 22.59 | 112.9 |
|  | 50.00 | 171368795 | 965397 | 177.5 | 58.19 | 116.4 |
|  | 100.0 | 306467240 | 998097 | 307.0 | 108.4 | 108.4 |
|  | 200.0 | 490490952 | 991284 | 494.8 | 181.2 | 90.59 |
|  | 500.0 | 852677489 | 739819 | 1152 | 436.1 | 87.23 |
|  | 1000 | 1356252769 | 662459 | 2047 | 782.9 | 78.30 |
|  | 2000 | 1764715785 | 329651 | 5353 | 2064 | 103.2 |
| AA | 100.0 | 2894424 | 429008 | 6.747 | 102.2 | 102.2 |
|  | 200.0 | 2066454 | 296640 | 6.966 | 194.2 | 97.10 |
|  | 400.0 | 2264380 | 303666 | 7.457 | 399.8 | 99.96 |
|  | 1000 | 3883747 | 451825 | 8.596 | 877.3 | 87.73 |
|  | 2000 | 3930366 | 339017 | 11.59 | 2134 | 106.7 |
|  | 4000 | 6939055 | 421605 | 16.46 | 4173 | 104.3 |
|  | 10000 | 14140186 | 450788 | 31.37 | 10423 | 104.2 |
|  | 20000 | 23751280 | 434000 | 54.73 | 20215 | 101.1 |
|  | 40000 | 47282692 | 478909 | 98.73 | 38662 | 96.66 |
| BHB | 0.20 | 298774 | 646655 | 0.462 | 0.22 | 108.6 |
|  | 0.40 | 723642 | 919176 | 0.787 | 0.34 | 85.05 |
|  | 0.80 | 1719538 | 863394 | 1.992 | 0.80 | 99.46 |
|  | 2.00 | 4586590 | 965397 | 4.751 | 1.84 | 91.97 |
|  | 4.00 | 9808012 | 998097 | 9.827 | 3.76 | 93.98 |
|  | 8.00 | 20302431 | 991284 | 20.48 | 7.79 | 97.36 |
|  | 20.00 | 44208917 | 739819 | 59.76 | 22.64 | 113.2 |
|  | 40.00 | 77252029 | 662459 | 116.6 | 44.15 | 110.4 |
|  | 80.00 | 101790227 | 329651 | 308.8 | 116.8 | 146.0 |
| PA | 1.00 | 221614 | 429008 | 0.517 | 1.06 | 106.0 |
|  | 2.00 | 163239 | 296640 | 0.550 | 1.75 | 87.54 |
|  | 4.00 | 203106 | 303666 | 0.669 | 4.18 | 104.5 |
|  | 10.00 | 401896 | 451825 | 0.889 | 8.70 | 87.03 |
|  | 20.00 | 503734 | 339017 | 1.486 | 20.93 | 104.6 |
|  | 40.00 | 1070203 | 421605 | 2.538 | 42.50 | 106.3 |
|  | 100.00 | 2466697 | 450788 | 5.472 | 102.6 | 102.6 |
|  | 200.00 | 4696600 | 434000 | 10.82 | 212.3 | 106.1 |
|  | 400.00 | 9127719 | 478909 | 19.06 | 381.1 | 95.28 |
| BD | 1.00 | 908 | 429008 | 0.002 | 1.04 | 104.4 |
|  | 2.00 | 1575 | 296640 | 0.005 | 1.85 | 92.64 |
|  | 4.00 | 3993 | 303666 | 0.013 | 3.84 | 95.96 |
|  | 10.00 | 17082 | 451825 | 0.038 | 10.08 | 100.8 |
|  | 20.00 | 27581 | 339017 | 0.081 | 21.12 | 105.6 |
|  | 40.00 | 62678 | 421605 | 0.149 | 38.16 | 95.41 |
|  | 100.0 | 172588 | 450788 | 0.383 | 97.49 | 97.49 |
|  | 200.0 | 351370 | 434000 | 0.810 | 205.6 | 102.8 |
|  | 400.0 | 792024 | 478909 | 1.654 | 419.4 | 104.9 |
| iso-BA | 1.00 | 1404395 | 1285826 | 1.092 | 1.02 | 102.2 |
|  | 2.00 | 1520402 | 1367840 | 1.112 | 2.07 | 103.4 |
|  | 4.00 | 1945135 | 1709833 | 1.138 | 3.48 | 86.99 |
|  | 10.00 | 1959902 | 1581113 | 1.240 | 9.00 | 89.98 |
|  | 20.00 | 2039422 | 1397722 | 1.459 | 20.88 | 104.4 |
|  | 40.00 | 2582022 | 1396008 | 1.850 | 42.02 | 105.1 |
|  | 100.0 | 5272013 | 1798727 | 2.931 | 100.6 | 100.6 |
|  | 200.0 | 6653869 | 1355020 | 4.911 | 207.7 | 103.9 |
| BA | 1.00 | 128140 | 429008 | 0.299 | 0.97 | 96.61 |
|  | 2.00 | 131413 | 296640 | 0.443 | 2.23 | 111.4 |
|  | 4.00 | 183631 | 303666 | 0.605 | 3.64 | 91.03 |
|  | 10.00 | 584116 | 451825 | 1.293 | 9.66 | 96.56 |
|  | 20.00 | 872927 | 339017 | 2.575 | 20.86 | 104.3 |
|  | 40.00 | 2068502 | 421605 | 4.906 | 41.24 | 103.1 |
|  | 100.0 | 5276822 | 450788 | 11.71 | 100.7 | 100.7 |
|  | 200.0 | 10111852 | 434000 | 23.29 | 202.0 | 101.0 |
|  | 400.0 | 20984986 | 478909 | 43.82 | 381.4 | 95.34 |
| iso-VA | 1.00 | 281239 | 1285826 | 0.219 | 1.02 | 102.2 |
|  | 2.00 | 307223 | 1367840 | 0.225 | 1.83 | 91.35 |
|  | 4.00 | 413883 | 1709833 | 0.242 | 4.22 | 105.4 |
|  | 10.00 | 462215 | 1581113 | 0.292 | 11.10 | 111.0 |
|  | 20.00 | 476538 | 1397722 | 0.341 | 17.75 | 88.74 |
|  | 40.00 | 759304 | 1396008 | 0.544 | 45.53 | 113.8 |
|  | 100.0 | 1601991 | 1798727 | 0.891 | 92.98 | 92.98 |
|  | 200.0 | 2144933 | 1355020 | 1.583 | 187.7 | 93.86 |
|  | 400.0 | 5401085 | 1712263 | 3.154 | 402.8 | 100.7 |
| VA | 0.100 | 64855 | 429008 | 0.151 | 0.10 | 104.4 |
|  | 0.200 | 47593 | 296640 | 0.160 | 0.18 | 91.06 |
|  | 0.400 | 57628 | 303666 | 0.190 | 0.43 | 107.0 |
|  | 1.00 | 108802 | 451825 | 0.241 | 0.86 | 85.58 |
|  | 2.00 | 123524 | 339017 | 0.364 | 1.89 | 94.58 |
|  | 4.00 | 251764 | 421605 | 0.597 | 3.84 | 96.08 |
|  | 10.00 | 634274 | 450788 | 1.407 | 10.63 | 106.3 |
|  | 20.00 | 1150944 | 434000 | 2.652 | 21.07 | 105.3 |
|  | 40.00 | 2571342 | 478909 | 5.369 | 43.85 | 109.6 |
| iso-CA | 0.10 | 87393 | 429008 | 0.204 | 0.09 | 92.64 |
|  | 0.20 | 131174 | 296640 | 0.442 | 0.22 | 112.2 |
|  | 0.40 | 240459 | 303666 | 0.792 | 0.42 | 104.4 |
|  | 1.00 | 823222 | 451825 | 1.822 | 0.99 | 98.64 |
|  | 2.00 | 1329429 | 339017 | 3.921 | 2.15 | 107.3 |
|  | 4.00 | 3064629 | 421605 | 7.269 | 4.00 | 99.87 |
|  | 10.00 | 8175411 | 450788 | 18.14 | 10.00 | 99.96 |
|  | 20.00 | 15328208 | 434000 | 35.32 | 19.49 | 97.43 |
|  | 40.00 | 30414224 | 478909 | 63.51 | 35.05 | 87.63 |
| CA | 1.00 | 270972 | 429008 | 0.632 | 1.03 | 103.1 |
|  | 2.00 | 201320 | 296640 | 0.679 | 2.11 | 105.6 |
|  | 4.00 | 219092 | 303666 | 0.721 | 3.10 | 77.42 |
|  | 10.00 | 447135 | 451825 | 0.990 | 9.26 | 92.60 |
|  | 20.00 | 515044 | 339017 | 1.519 | 21.43 | 107.8 |
|  | 40.00 | 994864 | 421605 | 2.360 | 40.75 | 101.9 |
|  | 100.0 | 2240907 | 450788 | 4.971 | 100.8 | 100.8 |
|  | 200.0 | 4333393 | 434000 | 9.985 | 216.0 | 108.0 |
|  | 400.0 | 8899966 | 478909 | 18.58 | 413.7 | 103.4 |

Note: lactic acid (LA); β-hydroxybutyrat (BHB); acetic acid (AA); propionic acid (PA); crotonic acid (BD-CA); isobutyric acid (ISO-BA); butyric acid (BA); isovaleric acid (ISO-VA); valeric acid (VA); isocaproic acid (ISO-CA); caproic acid (CA).

Table S4 Intra- and inter-batches precisions and accuracies of the 11 SCFAs targets.

| Analyte | Conc. (ng/mL) | Intra-batches (n=6) |  |  | Inter-batches (n=18) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Mean ± SD | Precision (RSD, %) | Accuracy (RE, %) | Mean ± SD | Precision (RSD, %) | Accuracy (RE, %) |
| LA a | 25.0 | 33.21±1.23 | 3.70 | 132.8 | 38.87±5.20 | 13.4 | 155.5 |
|  | 625 | 474.0±61.12 | 12.9 | 75.83 | 460.9±47.43 | 10.3 | 73.74 |
|  | 1250 | 1167±82.47 | 7.06 | 93.42 | 1111±177.4 | 15.6 | 88.92 |
| AA a | 0.50 | 0.45±0.10 | 2.22 | 90.00 | 0.55±0.12 | 21.3 | 109.7 |
|  | 12.5 | 14.79±0.36 | 2.43 | 118.3 | 13.62±1.10 | 8.1 | 109.0 |
|  | 25.0 | 28.55±0.65 | 2.29 | 114.2 | 25.08±2.71 | 10.8 | 100.3 |
| BHB a | 1.00 | 1.38±0.10 | 6.94 | 138.0 | 1.41±0.09 | 6.2 | 140.6 |
|  | 25.0 | 27.33±2.65 | 9.70 | 109.3 | 28.75±2.78 | 9.7 | 115.0 |
|  | 50.0 | 47.07±7.21 | 15.32 | 94.14 | 50.37±7.85 | 15.6 | 100.7 |
| PA b | 5.00 | 5.90±0.34 | 5.84 | 118.0 | 5.46±1.03 | 18.8 | 109.2 |
|  | 125 | 147.2±2.22 | 1.51 | 117.8 | 134.7±9.57 | 7.1 | 107.7 |
|  | 250 | 293.6±8.78 | 2.99 | 117.4 | 268.5±22.53 | 8.4 | 107.4 |
| BD b | 5.00 | 6.24±0.15 | 2.46 | 124.8 | 5.69±0.56 | 9.8 | 113.9 |
|  | 125 | 125.0±1.22 | 0.98 | 100.0 | 129.3±4.84 | 3.7 | 103.5 |
|  | 250 | 245.9±2.33 | 0.95 | 98.36 | 260.1±11.99 | 4.6 | 104.0 |
| iso-BA b | 5.00 | 5.16±0.29 | 5.70 | 103.2 | 5.70±0.61 | 10.3 | 114.0 |
|  | 125 | 119.1±5.95 | 5.00 | 95.25 | 125.9±15.83 | 10.8 | 100.8 |
|  | 250 | 242.6±2.55 | 1.05 | 97.04 | 255.3±26.37 | 12.6 | 102.1 |
| BA b | 5.00 | 5.15±0.12 | 2.39 | 103.0 | 5.17±0.39 | 7.6 | 103.4 |
|  | 125 | 124.3±1.98 | 1.59 | 99.48 | 125.07±1.48 | 1.2 | 100.1 |
|  | 250 | 246.3±3.17 | 1.29 | 98.50 | 247.0±3.49 | 1.4 | 98.80 |
| iso-VA b | 5.00 | 6.38±0.04 | 0.69 | 127.6 | 6.40±0.99 | 15.5 | 127.9 |
|  | 125 | 101.5±2.91 | 2.86 | 81.20 | 119.6±14.07 | 11.8 | 95.70 |
|  | 250 | 193.9±4.69 | 2.42 | 77.58 | 237.6±33.94 | 14.3 | 95.03 |
| VA b | 0.50 | 0.44±0.03 | 5.88 | 88.00 | 0.46±0.07 | 15.8 | 91.60 |
|  | 12.5 | 10.94±0.38 | 3.44 | 87.52 | 11.81±0.68 | 5.8 | 94.47 |
|  | 25.0 | 21.96±0.31 | 1.40 | 87.84 | 24.05±1.56 | 6.5 | 96.21 |
| iso-CA b | 0.50 | 0.50±0.02 | 4.81 | 100.0 | 0.52±0.03 | 5.8 | 104.2 |
|  | 12.5 | 11.33±0.49 | 4.37 | 90.64 | 11.76±0.51 | 4.3 | 94.11 |
|  | 25.0 | 21.47±0.30 | 1.40 | 85.88 | 22.92±1.65 | 7.2 | 91.67 |
| CA b | 5.00 | 4.74±0.69 | 14.6 | 94.7 | 4.74±0.69 | 21.0 | 90.04 |
|  | 125 | 128.4±0.96 | 0.75 | 102.7 | 123.9±4.01 | 3.2 | 99.11 |
|  | 250 | 254.6±2.68 | 1.05 | 101.8 | 252.8±3.93 | 1.6 | 101.1 |

Note: Tryptophan (TRP); serotonin (5-HT); kynurenine (KYN); kynurenic acid (KA); xanthurenic acid (XA); indole-3-acetic acid (IAA); indole-3-propionic acid (IPA); tryptamine (TA), a: (g/mL); b: (ng/mL);

Table S5 The raw data of the intra- and inter-day precisions and accuracies of 11 SCFAs.

| Analyte | batch | Conc. | 1 | 2 | 3 | 4 | 5 | 6 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| LA a | 1 | 25.0 | 33.67 | 31.43 | 35.20 | 32.94 | 33.16 | 32.84 |
|  |  | 625 | 411.5 | 531.3 | 568.02 | 445.5 | 441.8 | 445.8 |
|  |  | 1250 | 1239 | 1176 | 1236 | 1060 | 1071 | 1224 |
|  | 2 | 25.0 | 42.13 | 39.08 | 39.65 | 37.48 | 36.47 | 37.15 |
|  |  | 625 | 426.4 | 468.1 | 511.6 | 404.7 | 405.6 | 439.2 |
|  |  | 1250 | 1001 | 1507 | 1272 | 1273 | 1191 | 1226 |
|  | 3 | 25.0 | 48.07 | 40.64 | 43.84 | 45.74 | 44.82 | 45.37 |
|  |  | 625 | 428.4 | 518.6 | 511.8 | 429.4 | 454.2 | 454.2 |
|  |  | 1250 | 940.7 | 916.3 | 942.1 | 918.5 | 904.9 | 905.8 |
| BHB a | 1 | 1.00 | 1.23 | 1.38 | 1.51 | 1.40 | 1.35 | 1.44 |
|  |  | 25.0 | 25.50 | 30.24 | 30.91 | 25.79 | 24.49 | 27.04 |
|  |  | 50.0 | 1.26 | 1.40 | 1.51 | 1.38 | 1.41 | 1.47 |
|  | 2 | 1.00 | 41.22 | 46.68 | 41.48 | 52.14 | 58.99 | 41.92 |
|  |  | 25.0 | 27.63 | 32.05 | 32.81 | 27.68 | 26.51 | 29.24 |
|  |  | 50.0 | 44.68 | 48.81 | 43.22 | 56.31 | 69.51 | 45.24 |
|  | 3 | 1.00 | 1.31 | 1.35 | 1.54 | 1.42 | 1.45 | 1.51 |
|  |  | 25.0 | 27.19 | 33.14 | 32.83 | 27.78 | 26.16 | 30.58 |
|  |  | 50.0 | 49.98 | 51.90 | 46.96 | 56.84 | 62.55 | 48.28 |
| AA a | 1 | 0.50 | 0.4528 | 0.4582 | 0.4387 | 0.4488 | 0.4506 | 0.4689 |
|  |  | 12.5 | 14.97 | 14.23 | 14.74 | 14.89 | 15.30 | 14.63 |
|  |  | 25.0 | 28.70 | 28.58 | 27.36 | 29.33 | 28.85 | 28.50 |
|  | 2 | 0.50 | 0.6365 | 0.6220 | 0.6406 | 0.6306 | 0.6256 | 0.6476 |
|  |  | 12.5 | 13.78 | 14.91 | 13.37 | 12.68 | 12.73 | 14.06 |
|  |  | 25.0 | 24.96 | 23.43 | 23.60 | 23.22 | 25.55 | 24.37 |
|  | 3 | 0.50 | 0.4733 | 0.4513 | 0.4098 | 0.4907 | 0.8031 | 0.7192 |
|  |  | 12.5 | 12.09 | 12.90 | 12.64 | 12.40 | 12.25 | 12.67 |
|  |  | 25.0 | 22.42 | 21.84 | 23.07 | 22.72 | 22.50 | 22.39 |
| PA b | 1 | 5.00 | 5.56 | 6.45 | 6.20 | 5.67 | 5.81 | 5.72 |
|  |  | 125 | 147.5 | 144.0 | 145.0 | 149.6 | 148.3 | 148.9 |
|  |  | 250 | 300.7 | 296.2 | 276.9 | 293.8 | 293.2 | 300.5 |
|  | 2 | 5.00 | 5.89 | 4.47 | 5.52 | 5.95 | 4.45 | 6.11 |
|  |  | 125 | 130.6 | 133.0 | 130.2 | 130.5 | 130.6 | 131.8 |
|  |  | 250 | 279.6 | 264.6 | 269.5 | 266.4 | 271.5 | 267.3 |
|  | 3 | 5.00 | 1.87 | 5.70 | 5.71 | 5.77 | 5.52 | 5.85 |
|  |  | 125 | 122.6 | 125.8 | 127.8 | 126.7 | 125.3 | 125.6 |
|  |  | 250 | 244.8 | 237.7 | 244.3 | 236.0 | 247.3 | 241.7 |
| BD-CA b | 1 | 5.00 | 6.18 | 6.37 | 6.44 | 6.172 | 6.01 | 6.27 |
|  |  | 125 | 125.2 | 123.3 | 126.0 | 123.7 | 125.8 | 126.2 |
|  |  | 250 | 247.3 | 246.0 | 243.8 | 247.2 | 248.6 | 242.5 |
|  | 2 | 5.00 | 5.21 | 4.99 | 4.44 | 5.81 | 5.94 | 5.32 |
|  |  | 125 | 131.2 | 134.4 | 137.1 | 130.9 | 135.2 | 141.2 |
|  |  | 250 | 274.0 | 276.8 | 274.9 | 258.5 | 277.8 | 269.5 |
|  | 3 | 5.00 | 5.4 | 5.23 | 5.3 | 5.96 | 5.33 | 6.14 |
|  |  | 125 | 128.6 | 127.7 | 128.8 | 128.5 | 127.5 | 127.0 |
|  |  | 250 | 261.9 | 258.9 | 266.7 | 267.2 | 262.2 | 257.5 |
| iso-BA b | 1 | 5.00 | 4.80 | 5.24 | 5.17 | 5.64 | 4.90 | 5.18 |
|  |  | 125 | 120.8 | 123.5 | 113.8 | 119.6 | 126.3 | 110.4 |
|  |  | 250 | 240.0 | 243.1 | 240.6 | 242.3 | 247.2 | 242.5 |
|  | 2 | 5.00 | 5.61 | 5.194 | 5.439 | 6.877 | 5.89 | 5.198 |
|  |  | 125 | 150 | 167.61 | 129.3 | 128.19 | 128.41 | 154.14 |
|  |  | 250 | 290.2 | 258.3 | 286.2 | 271.9 | 322.9 | 294.4 |
|  | 3 | 5.00 | 6.44 | 6.37 | 5.71 | 6.46 | 6.41 | 6.03 |
|  |  | 125 | 111.2 | 119.5 | 112.4 | 117.9 | 116.9 | 117.2 |
|  |  | 250 | 238.6 | 232.8 | 233.3 | 235.3 | 241.0 | 235.6 |
| BA b | 1 | 5.00 | 5.00 | 5.10 | 5.16 | 5.09 | 5.18 | 5.37 |
|  |  | 125 | 125.9 | 122.0 | 121.7 | 126.1 | 125.5 | 124.9 |
|  |  | 250 | 247.1 | 248.9 | 240.0 | 247.6 | 247.7 | 246.2 |
|  | 2 | 5.00 | 4.92 | 4.72 | 5.01 | 4.99 | 4.64 | 5.02 |
|  |  | 125 | 125.5 | 127.9 | 124.8 | 126.3 | 123.9 | 126.2 |
|  |  | 250 | 249.7 | 249.0 | 247.7 | 243.7 | 254.3 | 248.8 |
|  | 3 | 5.00 | 5.29 | 5.6 | 4.51 | 5.86 | 5.74 | 5.86 |
|  |  | 125 | 124.6 | 126.2 | 124.8 | 125.0 | 124.6 | 125.5 |
|  |  | 250 | 242.9 | 240.3 | 248.6 | 247.0 | 250.0 | 246.7 |
| ISO-VA b | 1 | 5.00 | 6.37 | 6.45 | 6.34 | 6.33 | 6.39 | 6.39 |
|  |  | 125 | 99.90 | 99.18 | 99.29 | 106.55 | 103.33 | 100.76 |
|  |  | 250 | 196.0 | 192.7 | 189.2 | 198.7 | 198.9 | 188.1 |
|  | 2 | 5.00 | 6.47 | 6.01 | 6.39 | 6.03 | 5.92 | 6.23 |
|  |  | 125 | 130.6 | 132.2 | 136.0 | 134.6 | 132.8 | 137.4 |
|  |  | 250 | 271.5 | 275.5 | 270.7 | 271.8 | 273.7 | 275.6 |
|  | 3 | 5.00 | 6.05 | 7.49 | 3.76 | 6.43 | 9.14 | 6.95 |
|  |  | 125 | 123.4 | 123.1 | 125.3 | 123.2 | 123.8 | 121.8 |
|  |  | 250 | 245.4 | 243.5 | 250.7 | 246.1 | 246.5 | 241.5 |
| VA b | 1 | 0.50 | 0.41 | 0.45 | 0.45 | 0.42 | 0.48 | 0.42 |
|  |  | 12.5 | 10.85 | 10.62 | 10.86 | 11.32 | 11.47 | 10.52 |
|  |  | 25.0 | 21.91 | 21.76 | 21.78 | 22.56 | 21.78 | 21.98 |
|  | 2 | 0.50 | 0.48 | 0.42 | 0.37 | 0.38 | 0.37 | 0.61 |
|  |  | 12.5 | 11.98 | 12.63 | 12.36 | 12.39 | 12.30 | 12.12 |
|  |  | 25.0 | 24.77 | 25.32 | 24.78 | 24.56 | 25.13 | 24.72 |
|  | 3 | 0.50 | 0.48 | 0.51 | 0.37 | 0.59 | 0.5 | 0.54 |
|  |  | 12.5 | 12.03 | 12.08 | 12.53 | 12.29 | 12.11 | 12.09 |
|  |  | 25.0 | 25.76 | 24.72 | 25.57 | 25.26 | 25.38 | 25.23 |
| ISO-CA b | 1 | 0.50 | 0.50 | 0.52 | 0.53 | 0.50 | 0.46 | 0.48 |
|  |  | 12.5 | 11.80 | 11.07 | 10.84 | 11.79 | 11.73 | 10.76 |
|  |  | 25.0 | 21.78 | 21.57 | 21.48 | 21.75 | 21.06 | 21.17 |
|  | 2 | 0.50 | 0.53 | 0.52 | 0.52 | 0.48 | 0.51 | 0.51 |
|  |  | 12.5 | 12.21 | 12.49 | 11.51 | 11.39 | 11.39 | 11.92 |
|  |  | 25.0 | 22.75 | 21.673 | 21.59 | 23.15 | 22.75 | 21.74 |
|  | 3 | 0.50 | 0.570 | 0.580 | 0.530 | 0.550 | 0.550 | 0.520 |
|  |  | 12.5 | 12.07 | 12.1 | 12.22 | 12.31 | 12.13 | 12.01 |
|  |  | 25.0 | 25.05 | 24.00 | 24.87 | 24.99 | 25.89 | 25.28 |
| CA b | 1 | 5.00 | 4.590 | 5.300 | 3.660 | 4.900 | 4.370 | 5.590 |
|  |  | 125 | 128.5 | 126.9 | 129.6 | 128.6 | 127.8 | 128.9 |
|  |  | 250 | 254.3 | 254.2 | 252.3 | 258.4 | 256.9 | 251.2 |
|  | 2 | 5.00 | 4.860 | 2.640 | 4.000 | 3.870 | 4.230 | 5.580 |
|  |  | 125 | 122.0 | 127.0 | 125.4 | 122.6 | 121.0 | 124.2 |
|  |  | 250 | 250.4 | 248.0 | 247.5 | 244.1 | 255.6 | 249.8 |
|  | 3 | 5.00 | 4.490 | 4.260 | 2.510 | 5.870 | 5.520 | 4.790 |
|  |  | 125 | 120.3 | 118.7 | 118.4 | 118.4 | 122.2 | 119.5 |
|  |  | 250 | 254.9 | 253.9 | 257.4 | 251.7 | 251.8 | 258.1 |

Note: lactic acid (LA); β-hydroxybutyrat (BHB); acetic acid (AA); propionic acid (PA); crotonic acid (BD-CA); isobutyric acid (ISO-BA); butyric acid (BA); isovaleric acid (ISO-VA); valeric acid (VA); isocaproic acid (ISO-CA); caproic acid (CA). a: (g/mL); b: (ng/mL);

Table S6 Stability of 11 SCFAs for method validation.

| Analyte | Conc. (ng/ml) | Stability (RSD, %) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Freeze–thaw (-20℃) 3 cycles | Autosampler (4℃, 36 h) | Working solution (-20℃, 5 day) | Working solution (4℃, 5 day) |
| LA a | 25.0 | 84.30 (3.9) | 108.0 (4.4) | 105.8 (6.9) | 108.4 (2.8) |
|  | 625 | 91.46 (2.0) | 95.73 (9.9) | 89.53 (8.6) | 91.51 (9.2) |
| AA a | 0.50 | 93.88 (6.2) | 105.0 (4.6) | 118.1 (5.6) | 119.8 (6.0) |
|  | 25.0 | 97.51 (3.3) | 99.48 (4.5) | 105.2 (3.6) | 99.95 (5.5) |
| BHB a | 1.00 | 100.6 (5.0) | 106.2 (6.7) | 111.8 (1.9) | 116.8 (1.0) |
|  | 50.0 | 101.4 (6.3) | 101.6 (11.2) | 83.00 (3.2) | 83.32 (4.0) |
| PA b | 5.00 | 107.9 (4.4) | 108.0 (4.1) | 96.80 (4.1) | 87.90 (14.5) |
|  | 250 | 86.40 (2.4) | 86.41 (2.7) | 86.02 (2.3) | 82.05 (3.5) |
| BD-CA b | 5.00 | 92.65 (6.4) | 101.1 (7.6) | 94.91 (12.6) | 94.20 (9.2) |
|  | 250 | 91.71 (1.9) | 104.8 (2.7) | 84.81 (1.9) | 86.82 (3.1) |
| iso-BA b | 5.00 | 89.30 (7.5) | 110.3 (7.4) | 115.3 (3.6) | 105.7 (15.1) |
|  | 250 | 92.89 (4.2) | 98.15 (5.9) | 104.9 (6.8) | 96.59 (7.9) |
| BA b | 5.00 | 102.1 (3.4) | 99.67 (4.0) | 105.0 (3.6) | 101.4 (7.8) |
|  | 250 | 98.27 (1.7) | 102.9 (2.6) | 97.16 (2.3) | 94.32 (1.5) |
| iso-VA b | 5.00 | 74.17 (11.8) | 116.8 (15.2) | 85.78 (11.7) | 82.32 (12.5) |
|  | 250 | 106.6 (3.3) | 104.6 (4.5) | 116.3 (4.2) | 116.7 (18.4) |
| VA b | 0.50 | 113.7 (5.6) | 107.3 (14.2) | 100.5 (0.4) | 99.58 (0.9) |
|  | 25.0 | 105.0 (2.0) | 101.9 (2.4) | 109.4 (1.4) | 105.3 (1.1) |
| iso-CA b | 0.50 | 111.2 (3.3) | 102.0 (2.9) | 102.7 (3.6) | 105.76 (3.1) |
|  | 25.0 | 107.0 (2.2) | 99.82 (1.0) | 112.8 (4.0) | 107.3 (5.9) |
| CA b | 5.00 | 119.4 (5.8) | 98.54 (7.2) | 125.0 (12.7) | 128.5 (15.8) |
|  | 250 | 108.6 (2.8) | 100.3 (2.0) | 110.2 (2.3) | 107.2 (3.6) |

Note: lactic acid (LA); β-hydroxybutyrat (BHB); acetic acid (AA); propionic acid (PA); crotonic acid (BD-CA); isobutyric acid (ISO-BA); butyric acid (BA); isovaleric acid (ISO-VA); valeric acid (VA); isocaproic acid (ISO-CA); caproic acid (CA). a: (g/mL); b: (ng/mL);

Table S7 Concentrations of 11 SCFAs in 1028 Chinese elderly serum samples

| Analyte | %>LOD | Geometric mean (95% CI) | Percentile |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5th | 25th | 50th | 75th | 95th |
| LA a | 100 | 331.1 (356.3, 377.7) | 156.6 | 243.5 | 328.0 | 445.3 | 751.4 |
| BHB a | 100 | 9.80 (10.96, 11.86) | 4.100 | 7.020 | 9.620 | 13.49 | 24.33 |
| AA b | 100 | 1580 (2239, 2501) | 191.7 | 1045 | 1862 | 2869 | 6600 |
| PA b | 100 | 18.14 (19.69, 21.12) | 8.650 | 13.30 | 17.87 | 24.18 | 41.22 |
| BD-CA b | 67.9 | 0.67 (0.84, 0.92) | 0.25 | 0.36 | 0.78 | 1.08 | 2.30 |
| iso-BA b | 100 | 28.43 (41.60, 47.71) | 5.49 | 17.48 | 30.20 | 49.97 | 151.9 |
| BA b | 100 | 5.56 (6.05, 6.54) | 2.92 | 4.11 | 5.24 | 7.15 | 13.01 |
| iso-VA b | 100 | 47.52 (66.90, 74.89) | 12.21 | 33.03 | 53.55 | 88.67 | 183.1 |
| VA b | 100 | 0.87 (1.31, 1.51) | 0.14 | 0.48 | 0.95 | 1.74 | 3.95 |
| iso-CA b | 100 | 0.102 (0.115, 0.124) | 0.04 | 0.07 | 0.10 | 0.16 | 0.25 |
| CA b | 100 | 81.65 (95.72, 103.2) | 32.57 | 63.35 | 87.30 | 121.7 | 208.5 |

Note: CI: confidence interval; lactic acid (LA); β-hydroxybutyrat (BHB); acetic acid (AA); propionic acid (PA); crotonic acid (BD-CA); isobutyric acid (ISO-BA); butyric acid (BA); isovaleric acid (ISO-VA); valeric acid (VA); isocaproic acid (ISO-CA); caproic acid (CA).

a: (g/mL); b: (ng/mL);