Table 1. Number and type of masks tested for two exposure doses.

| Type of facemask | Gamma radiation dose |  |  |
| --- | --- | --- | --- |
|  | 0 (kGy) | 15 (kGy) | 25 (kGy) |
| Surgical mask (Nos.) | 20 | 10 | 10 |
| Cloth mask (Nos.) | 20 | 10 | 10 |
| N95 mask (Nos.) | 6 | 3 | 3 |

Note: 0 kGy corresponds to control pieces (un-irradiated).

Table 2. Pressure drop across the filter for two-flow rate condition.

| S. N. | Type of mask | Flow rate (L min-1) | ΔP (pascal) for control masks | ΔP (pascal) for sterilized masks |
| --- | --- | --- | --- | --- |
| 1. | N95 | 20 | 31.32 ± 2.31 | 30.89 ± 2.64 |
|  |  | 90 | 258.9 ± 4.59 | 257.7 ± 3.59 |
| 2. | Surgical | 20 | 2.45 ± 2.03 | 2.95 ± 1.83 |
|  |  | 90 | 7.36 ± 2.08 | 7.16 ± 1.78 |
| 3. | Cloth | 20 | 11.57 ± 1.87 | 12.78 ± 2.12 |
|  |  | 90 | 73.55 ± 2.74 | 72.95 ± 1.74 |

Table 3. Filtration efficiency of three type of masks for two flow rate, control and irradiated mask for ambient aerosols.

Mask Type

Aerosols size in µm

Gamma radiation dose

0 (kGy)	15 (kGy)	25 (kGy)

20 L min-1	90 L min-1	20 L min-1	90 L min-1	20 L min-1	90 L min-1 Average FE (%) ± Uncertainty

|  | 0.3-5.0 | 20.28±1.49 | 18.84±0.33 | 20.44±1.68 | 19.21±1.51 | 19.74±1.92 | 19.63±2.67 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Cloth | 1.0-5.0 | 60.50±7.59 | 49.20±8.44 | 51.16±5.38 | 53.51±5.09 | 54.94±9.56 | 59.09±4.03 |
|  | 3.0-5.0 | 87.44±8.99 | 81.53±8.10 | 80.66±4.28 | 82.89±4.78 | 88.90±7.14 | 89.41±5.63 |
|  | 0.3-5.0 | 20.98±2.74 | 20.08±2.87 | 20.28±1.61 | 15.76±0.22 | 22.48±3.92 | 17.48±0.52 |
| Surgical | 1.0-5.0 | 59.22±9.25 | 51.03±8.88 | 50.12±2.55 | 50.89±6.49 | 56.85±5.58 | 49.89±7.63 |
|  | 3.0-5.0 | 90.36±4.69 | 80.19±8.33 | 81.65±9.23 | 77.68±12.21 | 84.60±6.28 | 73.15±3.73 |
|  | 0.3-5.0 | 99.68±0.05 | 94.98±1.98 | 69.78±1.07 | 63.12±2.22 | 70.14±1.18 | 65.44±2.48 |
| N95 | 1.0-5.0 | 99.51±0.59 | 99.99±0.01 | 99.67±0.52 | 99.63±0.59 | 99.57±0.45 | 99.05±1.53 |
|  | 3.0-5.0 | 99.99±0.01 | 99.99±0.01 | 99.99±0.01 | 99.99±0.01 | 99.99±0.01 | 99.99±0.01 |

Table 4. Comparison of N95 mask filtering efficiency for total concentration of aerosols sizes < 300 and > 300 nm for two-flow rate, control and after sterilization.

| Gamma Irradiation dose | Aerosol filtering efficiency with Uncertainty (%) |  |  |  |
| --- | --- | --- | --- | --- |
|  | Flow rate 20 L min-1 |  | Flow rate 90 L min-1 |  |
|  | 0.3 -5.0 µm | 10- 300 nm | 0.3 -5.0 µm | 10- 300 nm |
| 0 kGy | 99.98 ± 0.01 | 99.89 ± 0.03 | 99.31 ±0.05 | 98.92 ± 0.12 |
| 15 kGy | 84.28 ± 1.68 | 99.51 ± 0.21 | 91.02 ± 1.17 | 95.49 ± 0.89 |
| 25 kGy | 86.54 ± 1.63 | 99.04 ± 0.11 | 89.07 ± 1.15 | 96.16 ± 1.13 |

Table 5. Comparison of N95 respiratory mask filtering efficiency from literature.

| Authors | Expt. Parameters | Irradiation Dose |  | Filtering efficiency (%) |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | U : 0.4 m s-1, T : 22.9 ˗ 23.9ºC RH: 11.3 ˗ 21.4 % ΔP: 174 ˗ 192 pascal | Aerosols size (µm) |  | 0.3 | 0.5 | 1.0 |
| Cramer et al., 2020 |  | 0 kGy |  | 85.9 | 89.5 | 94.3 |
|  |  | 1 kGy |  | 29.2 | 41.1 | 70.3 |
|  |  | 10 kGy |  | 28.6 | 38.6 | 74.3 |
|  |  | 50 kGy |  | 24.8 | 36.7 | 69.5 |
|  |  | Aerosols size (µm) |  | 0.3 | 0.5 | 1.0 |
| Man et al., 2020 | Parameters were not mention in the paper | 0 kGy |  | 99.4 | 99.8 | 99.8 |
|  |  | 10 kGy |  | 55.4 | 79.1 | 98.1 |
|  |  | 25 kGy |  | -- | 57.5 | 98.7 |
|  | U : 0.066 m s-1, T : 25 ˗ 27 ºC RH: 65 ˗ 67 % ΔP:31.32±2.31 pascal | Aerosols size (µm) | 0.1 | 0.3 | 0.5 | 1.0 |
| Current work |  | 0 kGy | 99.8 | 99.7 | 99.8 | 99.9 |
|  |  | 15 kGy | 98.6 | 70.1 | 86.7 | 96.3 |
|  |  | 25 kGy | 96.2 | 68.5 | 87.9 | 96.7 |
| Lin et al., 2020 | T : 38 ± 2.5 ºC RH: 85 ± 5 % ΔP: 59 ˗ 75 pascal | Aerosols size (µm) | 0.007-0.882, (CMD-0.075) |  |  |  |
|  |  | 0 kGy |  | > 95 |  |  |
|  |  | 10, 25 and 30 kGy |  | 44 -77 |  |  |