Supplementary Table 3 Diagnostic performance of nurse-measured and kiosk BP values with 24-hour ABPM as reference (N = 117) *
	Daytime ABPM 
≥130/80 (reference)
	Nurse-measured
BP ≥140/90
	Kiosk BP ≥135/85

	ABPM positive for hypertension, %
	47.9
	47.9

	Sensitivity, % (95% CI)
	71.4 (57.8–82.7)
	82.3 (78.1–96.0)

	Specificity, % (95% CI)
	72.1 (59.2–82.9)
	36.1 (24.2–49.4)

	Positive predictive value, % (95% CI)
	70.2 (60.4–78.5)
	56.2 (51.0–61.3)

	Negative predictive value, % (95% CI)
	73.3 (63.8–81.0)
	78.6 (61.6–89.3)

	Positive likelihood ratio, (%)
	2.56 (1.66–3.97)
	1.40 (1.13–1.72)

	Negative likelihood ratio, (%)
	0.40 (0.25–0.62)
	0.30 (0.13–0.68)

	True positive (hypertensive), n (%)
	40 (34.2)
	50 (42.7)

	True negative (hypertensive), n (%)
	44 (37.6)
	22 (18.8)

	False positive, n (%)
	17 (14.5)
	39 (33.3)

	False negative, n (%)
	16 (13.7)
	6 (5.1)

	Accuracy, % (95% CI) 
	71.8 (62.7–79.7)
	61.6 (52.1–70.4)


*Diagnostic thresholds: 24-hour ABPM, ≥130 mmHg systolic and/or ≥80 mmHg diastolic (reference); nurse-measured, ≥140 mmHg systolic and/or ≥90 mmHg diastolic; kiosk, ≥135 mmHg systolic and/or ≥85 mmHg diastolic
Abbreviations: ABPM: ambulatory BP monitoring; BP: blood pressure; CI: confidence interval
 
