Supplementary Table S1. Results of galactomannan assays performed on two platforms using old and new reagent lots.

| Instrument | GEMINI | GEMINI | Multiskan | GEMINI | Result |
| --- | --- | --- | --- | --- | --- |
| Reagent lot | Old | Old (re) | Old | New |  |
| Case No. | Sample Index |  |  |  |  |
| S1 | 0.46 | 0.52 | 0.53 | NT | Positive |
| S2 | 6.18 | 8.63 | 10.63 | 6.23 | Positive |
| S3 | 5.49 | 6.09 | 6.04 | 4.61 | Positive |
| S4 | 6.18 | 8.63 | 11.75 | 6.23 | Positive |
| S5 | 6.18 | 8.63 | 10.56 | 6.23 | Positive |
| S6 | 5.33 | 7.17 | 7.13 | 5.46 | Positive |
| S7 | 6.18 | 8.63 | 10.14 | 6.23 | Positive |
| S8 | 6.18 | 8.63 | 11.07 | 6.23 | Positive |
| S9 | 0.49 | 0.54 | 0.59 | NT | Positive |
| S10 | 6.18 | 8.63 | 9.87 | 6.23 | Positive |
| S11 | 5.30 | 6.76 | 6.93 | 5.59 | Positive |
| S12 | 6.16 | 8.63 | 9.31 | 6.03 | Positive |
| S13 | 6.18 | 8.63 | 10.51 | 6.23 | Positive |
| S14 | 6.18 | 6.64 | 6.48 | 6.23 | Positive |
| S15 | 0.43 | 0.39 | 0.36 | 0.45 | Negative |
| S16 | 1.94 | 1.67 | 1.63 | 2.04 | Positive |

NT, not tested