Supplementary Information and Figures

Assay design, validation, and performance

Assay Design

The assay consists of 19 genes commonly mutated in CH: DNMT3A, TET2, ASXL1, TP53, PPM1D, SF3B1, SRSF2, IDH1, IDH2, JAK2, BCOR, MPL, U2AF1, U2AF2, ZRSR2, CBL, CHEK2, GNAS and GNB11–5. It includes the entire gene sequence for each gene and is ~60,000 bp in size. Unique molecular indexes (UMIs) were added to increase sensitivity and decrease error via collapsing when sequenced at high depths. The panel is sequenced at 1M reads (1000X) for standard depth and 15M reads (10,000X) for high depth for a collapsed coverage of ~3,500X using the UMIs. The assay was validated at both depths. We assessed the ability to detect mutations (SNVs and indels) in 117 replicates, across 208 variants total.

Standard Depth – Limit of Blank

A total of 47 clinical samples were tested and 57 positive calls were observed, including 1 hotspot.

Mean Distinct Coverage was 992X. Specificity (NPA) was 99.998% for all variants and 99.996% for hotspot variants.

Standard Depth – Theoretical Limit of Detection

95% confidence levels for the probit fit were between 0.9-1.2% VAF, except for non-SBS changes which were 0.6% VAF. The LoD of each SBS category is represented in Table 1.

Standard Depth – Experimental Limit of Detection

100% of variants were correctly called at the 1.85% VAF level, although these were detected at 0.6% VAF (Figure 1 A and C PPM1D Q5244*). This is slightly higher than was predicted using the probit estimate (Table 1), likely because of additional noise in this sample set, however variants were correctly called in 67% of cases down to 0.25% VAF. (Figure 1 C).

Standard Depth - Precision & Reproducibility

A total of 61 individual variants were evaluated for the precision analysis and 136 were evaluated for reproducibility. In the precision study a success rate of 100% was achieved. Overall reproducibility was 97.06% (95% CI: 94.21% - 99.91%) (Table 2). The variants not found were around the 1% VAF threshold of detection.  Tight standard deviations were observed for precision (0.21% - 2.42%). The standard deviations for reproducibility were 0.21% to 4.67% with one exception of 11.79%, due to an outlier that may have resulted from collecting samples at separate timepoints from the same individual.

Standard Depth - Accuracy and Concordance with MSK-IMPACT

62 clinical samples across 117 replicates were compared with MSK-IMPACT6 blood sequencing results. Several months may have passed between the CH and IMAPCT blood draws (max. 12 months). 208 variants were analyzed, and 197 variants were correctly detected (94.71%). Concordance between the variant VAFs was high (adj.R2 = 0.891) and is shown in Figure 1 E.

High Depth – Limit of Blank

A total of 23 clinical samples were tested and 63 positive calls were observed, including 3 hotspot calls. Mean Distinct Coverage was 3870X. Specificity (NPA) was 99.995% for all variants and 99.978% for hotspot variants.

High Depth – Theoretical Limit of Detection

The 95% confidence levels for the probit fit were between 0.3-0.6% VAF. The LOD of each single base substitution (SBS) category is represented in Table 1.

High Depth – Experimental Limit of Detection

73% of variants were correctly called at 0.3% VAF and 100% were correctly called at 0.4% VAF. (Figure 1 B and D). Of the 4 missed at 0.3%, 2 were indels and 2 were G>A substitutions. This finding is similar to the predicted LOD from the probit fit (Table 1).

High Depth – Precision and Reproducibility

A total of 14 individual variants were evaluated for the precision analysis and 16 were evaluated for reproducibility. In the precision study a success rate of 100% (14/14) was achieved. Overall reproducibility was 100% (16/16) (Table 2). Tight standard deviations were observed for precision and reproducibility, varying from 0.21% - 1.63% and 0% - 1.84%, respectively.

High Depth – Accuracy and Concordance with MSK-IMPACT

17 clinical samples were compared with MSK-IMPACT blood sequencing results. Several months may have passed between the CH and IMAPCT blood draws (max. 12 months). 39 variants were analyzed, and all variants were correctly detected (100%). Concordance between the tests was high (adj.R2 = 0.929) and is shown in Figure 1 F.

Background error assessment and confidence score

Evaluating whole-blood clinical samples from normal controls, we identified a signature of low frequency background single base substitution (SBS) errors across samples that was present after consensus collapse of redundant paired-end sequences. These dominant SBS errors were consistent with previous reports of low frequency DNA noise signatures7,8, and thus we sought to develop a methodology for in silico filtering of these errors. Aggregating low frequency SBS errors into a sample-wide background score and adjusting for total distinct coverage, we found that this score sufficiently captured the variation observed in error signatures. We next developed a logistic regression model that adjusts for background score and individual variant characteristics including mutation type (SBS specific change or indel status), variant depth and mutant allele fraction to generate a variant specific confidence score for the likelihood that a mutation is reproducible above background levels.

Performance on two different sequencing platforms

18 CH sample libraries were sequenced by both Illumina and Complete Genomics. 44 mutations were identified in the Illumina run and 39 were identified in the Complete Genomics run. 77.3% of the mutations found by Illumina sequencing were recapitulated in the Complete Genomics results (34 mutations total). The missing mutations were of low VAF (Figure 2 A), likely the relatively low mean distinct coverage (495.4X and 313.0X in Complete Genomics and Illumina, respectively) lead to unreliable variant calling at these levels. Despite this discrepancy, the correlation between methods was high (adj.R2 = 0.95, p < 2.2e-16) (Figure 2 A) leading us to conclude that Complete Genomics is equivalent to Illumina in terms of calling variants.

Using our background error score, Complete Genomics had 2X higher error score levels than the same libraries sequenced on the Illumina platform (Figure 2 B). The correlation between methods was reasonably high (adj.R2 = 0.81), indicating a portion of the error, the portion observed across both platforms, was likely the result of sample processing through extraction, library generation, PCR or other upstream processes prior to sequencing. The Phred Quality scores, an indicator of the quality of the sequence data, were high in both sequencing types. Though the Illumina Phred Quality scores were higher than Complete Genomics, although Complete Genomics appeared to be more stable than Illumina (Figure 2 C). To further investigate the differences in error between sequencing types, we assessed the error rates from by each SBS type. 64,853 genomic positions were evaluated, and the mean distinct coverage was 181X. In all SBS types, except A>G and T>C, Complete Genomics had significantly higher error rates than Illumina (Figure 2 D). Libraries were prepared for Illumina sequencing, then adapted for Complete Genomics through additional steps including adapter ligation and PCR, which is the likely cause for the differences in error rates observed here. This difference in error profiles was not affected by the difference in the depth of coverage (Supplementary Figure 1).

Supplementary Figures

Figure S1. Limits of detection (LoD) and accuracy at standard and high depth sequencing

(A) LoD for standard depth and (B) high depth sequencing. (C) Standard depth and

(D) High depth LOD limit of detection tests. Heatmaps depicting presence and absence of mutations in each run. Each box represents one sample. Expected – expected VAF.  (E) Standard depth and (F) High depth scatter plots showing high accuracy between MSK-IMPACT and CH Panel VAFs.

Figure S2: Concordance and error rates between sequencing with Illumina and Complete Genomics sequencing platforms

(A) Accuracy between Illumina and Complete Genomics. (B) Error Scores between Illumina and Complete Genomics. (C) Phred Quality scores. (D) SBS error profiles for Complete Genomics compared to Illumina.

Figure S3: Estimated error rate vs mean distinct coverage of Complete Genomics Sequencing compared to Illumina by single base substitution type shows no influence of sequencing depth on error

Figure S4: MDS has lower input cfDNA than AML

Figure S5: Higher sensitivity and confidence of low VAF calls is possible when using multiple biospecimens

(A) Simulated sensitivity when incorporating one (single – yellow), two (duo – green) or three (trio – blue) biospecimens to establish a mutation call. (B) Observed sensitivity when incorporating multiple biospecimens (whole blood, saliva and cfDNA). (C) Specificity for each model of mutation calling.

Supplementary Tables

Table S1: Limit of detection according to probit fit

Probit estimated (Est.) 95% confidence (C95) limit of detection (LoD) including a background score of 0.5, a confidence score >25 and a minimum variant depth of 5.

nonSBS – non-single base substitution, such as insertions or deletions.

Table S2: Accuracy, precision, and reproducibility data at standard and high depth

Table S3: Blast data for AML and MDS patients

BMT = bone marrow transplant

MEC = Mitoxantrone, Etoposide, and Cytarabine

Table S4: Mean distinct coverage for disease type comparisons

MDS has lower mean distinct coverage for plasma cfDNA than buffy coat