Supplementary Material

Methods and Materials

Part I: Thermal Manikin Study

To ensure the comparability and continuity of our study, the four undercarriage heating systems previously studied were used again and connected to corresponding FAW blankets for testing, which were set as Group A; Group B; Group C; Group D. The temperature in the experimental room was maintained at 21~23℃, and the relative humidity was 40~60%. In this situation, we conducted heat distribution and heating rate (thermocouple) test, air pressure difference test, heater outlet temperature stability test, FAW blanket non-woven fabric strength test (Martindale abrasion resistance test), FAW blanket holes per unit area test and hole observation test respectively.

Heat distribution and heating rate test

In the experimental room simulating the temperature and humidity of the operating room (OR), the FAW blanket was lay on the operating table and connected the corresponding heater machine. Then the manikin model was placed on the blanket (It was placed according to the product instructions). Thermocouple thermometer probes were placed on the right abdomen, inner side of left thigh, inner side of right knee joint and inner side of left lower leg of the model (as shown in the Figure-1 at the red dot). And the manikin was covered with a sheet. Turn on the thermometer and record the first temperature (room temperature state). Then when the heater was turned on, use a stopwatch to count the time simultaneously. Record the temperature every 20 seconds for the first 10 minutes, then every half a minute, and stop after 30 minutes. The aim of this test was to plot the warming curve and compare the warming effect of different FAW blankets.

Air pressure difference test

All pre-test preparation work was similar to previous heat distribution and heating rate test part. Heater was turned on. After balance for 5 minutes, the probe was inserted of the manometer into the FAW blanket. Record the pressure value at each point, find out the maximum and minimum values of each blanket, and calculate the air pressure difference. Calculation formula: pressure difference = (maximum value-minimum value) / maximum value*100%. Specific points as shown in the supplementary material figure S1 (foot side, crotch, underarm, side of the head, top of the head).

Heater outlet temperature stability test

Four products of heaters were selected, and a thermometer was placed in the heater outlet in a room simulating the temperature and humidity of the OR (Supplementary material, Figure-S2). The first temperature (room temperature state) was recorded. While turning on the heaters, a stopwatch was used to record the temperature every 10 seconds for the first 5 minutes, and every 20 seconds thereafter. Stop recording at 10 minutes. This test was aimed to plot the heating curve to compare the heating rate (equilibrium time) and stability of different heaters.

FAW blanket non-woven strength test (Martindale abrasion resistance test)

Referring to American Society for Testing and Materials (ASTM) D4966 test method, four FAW blankets were selected and the blanket non-woven fabric was cut into 38mm diameter round samples using SDL Atlas Fabric Sample Cutter. The prepared samples were loaded onto the jig. The jig with the sample and the corresponding number of weights (9kPa) were mounted on the corresponding number of grinding discs (Supplementary material, Figure S3). Enter the number of revolutions and cycles of friction and make a total of 5 cycles of 20 revolutions each. After that, the samples were removed and the wear level or ASTM grade was judged according to the comparison (Supplementary material, Figure S4). A total of 5 parallel experiments were performed for each group, and the average value was selected.

The number of holes per unit area of FAW blanket

Four FAW blankets were selected. An area of 10cm*10cm was taken at the top, middle and bottom of the FAW blanket as well as the left, middle and right (Supplementary material, Figure S5), and the number of holes was counted and recorded.

Hole observation test (cell phone microscope hole observation, with or without fiber drift)

All pre-test preparation work was similar to previous heat distribution and heating rate test part. Then turn on the machine. A cell phone microscope was glued to the cell phone camera. A special application was used to observe the holes’ structure of the FAW blanket and whether there were fibers floating around. (Supplementary material, Figure S6)