# Tables

Table 1 List of survey questions. Question Reference; Survey Question(s). A; How many gamma cameras and SPECT/CT systems do you have at your site?. B; - Do you have a routine quality control program set up at your site? - If yes, is it based on vendor recommendations, based on recommendations from a professional body or you designed your own?. C; - What tests are performed, how frequently do you do them and who performs each test? - Who has oversight of the testing program?. D; - How are the test results actioned in your practice? - Do you use the vendor-supplied software to analyse the data? - Do you check the result against the vendor's passing criteria or your own criteria? - Do you perform additional data analyses on the QC results?.

Table 2 Information on participating sites and survey respondents.

| Category | No. of Sites |
| --- | --- |
| Total | 8 |
| Type of Site: |  |
| Public Hospitals | 6 |
| Part of Network of Public Hospitals | 2 |
| State/Territory: |  |
| Australian Capital Territory | 1 |
| New South Wales | 1 |
| Queensland | 1 |
| South Australia | 1 |
| Victoria | 3 |
| Western Australia | 1 |
| Survey Respondent: |  |
| Nuclear Medicine Physicist | 7 |
| Radiology Physicist | 1 |

Table 3 Summary of daily QC tests performed as part of the QC testing schedule for gamma cameras and SPECT/CT systems across the participating sites

| Test | Site 1 | Site 2 | Site 3 | Site 4 | Site 5 | Site 6 | Site 7 | Site 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A. GAMMA CAMERA - PLANAR |  |  |  |  |  |  |  |  |
| Intrinsic uniformity verification |  |  |  |  |  |  |  | a |
| System uniformity verification |  |  |  |  |  |  |  | a |
| B. CT |  |  |  |  |  |  |  |  |
| Daily CT testb |  |  |  |  |  |  |  |  |

a Intrinsic or extrinsic/system uniformity verification performed as per vendors' specification

b Vendor-specific CT test

Table 4 Summary of QC tests and profession(s) responsible for performing the tests in the QA programs for gamma cameras and SPECT/CT systems across the participating sites.

| Test | Site 1 | Site 2 | Site 3 | Site 4 | Site 5 | Site 6 | Site 7 | Site 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A. GAMMA CAMERA – PLANAR |  |  |  |  |  |  |  |  |
| Intrinsic and/or system uniformity verification (no tuning) | - | - | Yearly (MP) | Weekly (NMT) | - | - | - | - |
| High-count intrinsic and/or system uniformity verification (no tuning) | Yearly (MP) | - | - | - | - | - | - | Monthly (MP) |
| Intrinsic uniformity verification (post-tuning) | Weekly (NMT) | Weekly (NMT) | Weekly (NMT) | Weekly (NMT) | - | - | Weekly (MP) | - |
| High-count intrinsic uniformity verification (post-tuning) | Monthly (NMT) | Monthly (NMT) | Monthly (MP) | When requireda | Monthly (NMT) | Monthly (NMT/MP) | Monthly (MP) | Monthly (MP) |
| Intrinsic uniformity using off-peak PHA windows | Yearly (NMT) | - | Yearly (MP) | - | - | - | - | Ad-hoc (MP) |
| Intrinsic energy resolution | - | - | Yearly (MP) | - | - | - | - | - |
| System spatial resolution | Yearly (MP) | - | Monthly/ Yearly b (MP) | - | - | Yearly (NMT/MP) | Ad-hoc (MP) | - |
| System planar sensitivity | Yearly (MP) | - | Yearly/ Six-Monthlyc (MP) | Yearly (MP) | Ad-hoc (MP) | Monthly (NMT/MP) | Ad-hoc/ After service (MP) | Ad-hoc (MP) |
| Multiple window spatial resolution | - | - | Yearly (MP) | - | - | - | - | - |
| Collimator hole alignment | - | - | Yearly (MP) | - | - | - | - | - |
| B. GAMMA CAMERA - WHOLE-BODY |  |  |  |  |  |  |  |  |
| Whole-body spatial resolution | - | - | Yearly (MP) | - | - | - | - | - |
| Whole-body count rate variability & exposure time correction | - | - | Yearly (MP) | - | - | - | - | - |
| Whole-body exposure time correction | - | - | Yearly (MP) | - | - | - | - | - |
| C. SPECT |  |  |  |  |  |  |  |  |
| Centre of rotation | Quarterly (NMT) | Monthly (NMT) | Monthly (MP) | Weekly (NMT) | Monthly (NMT) | Yearly (NMT/MP) | Ad-hoc (MP) | Monthly (MP) |
| SPECT image quality | Six-Monthly (MP) | - | Yearly (MP) | Ad-hoc (MP) | Ad-hoc (MP) | - | - | Ad-hoc (MP) |
| SPECT spatial resolution | Yearly (MP) | - | Yearly (MP) | - | - | - | - | - |
| SPECT uniformity & sensitivity | Yearly (MP) | - | Yearly (MP) | - | - | - | - | - |
| D. SPECT/CT |  |  |  |  |  |  |  |  |
| SPECT/CT alignment calibration/verification | At serviced & Yearlye (MP) | At serviced & After servicef (MP) | At serviced & Ad-hoc (MP) | At serviced & Ad-hoc (MP) | At serviced | At serviced | At serviced | At serviced & Ad-hoc (MP) |
| E. CT |  |  |  |  |  |  |  |  |
| CT compliance testing | - | Yearly (MP) | - | - | 2-Yearly (MP) | - | Two-Yearly/ After servicef (MP) | Yearly/ After servicef (MP) |

MP = Medical Physicist(s); Tech = NMT = Nuclear Medicine Technologist(s); PHA = Pulse Height Analyser

a Performed by service engineers when required (based on daily and weekly uniformity results).

b Monthly qualitative and annual system spatial resolution tests.

c Six-monthly planar sensitivity tests performed if site is enrolled in clinical trials requiring quantitation.

d SPECT/CT alignment calibration and alignment check performed by service engineers during preventive maintenance service.

e Yearly SPECT/CT alignment check not performed by physicist(s) if test summary on review of service report from vendor engineer.

f After major service or repair.

Table 5 List of tests for gamma cameras and SPECT/CT systems in the NEMA NU-1 2018 testing standard [19].

| System | Test |
| --- | --- |
| Gamma Camera | Intrinsic Spatial Resolution |
|  | Intrinsic Spatial Linearity |
|  | Intrinsic Energy Resolution |
|  | Intrinsic Flood Field Uniformity |
|  | Multiple Window Spatial Registration |
|  | Intrinsic Count Rate Performance in Air |
|  | Intrinsic Spatial Resolution at 75 000 counts per second (cps) |
|  | Intrinsic Flood Field Uniformity at 75 000 cps |
|  | System Spatial Resolution without Scatter |
|  | System Spatial Resolution with Scatter |
|  | System Planar Sensitivity and Collimator Penetration and Scatter |
|  | Detector Shielding |
|  | System Count Rate Performance with Scatter |
| Whole-Body | Whole-Body System Spatial Resolution without Scatter |
| SPECT | System Alignment |
|  | SPECT Reconstructed Spatial Resolution without Scatter |
|  | SPECT Reconstructed Spatial Resolution with Scatter |
|  | System Volume Sensitivity |
|  | Detector-Detector Sensitivity Variation |
|  | Tomographic Contrast and Absolute Quantification Accuracy |
| SPECT/CT | SPECT/CT Co-Registration Accuracy |

Table 6 Version(s) of the NEMA NU-1 standard used in publications from sampled professional bodies.

| Professional Body | Publication | Publication Year | NU-1 Version(s) |
| --- | --- | --- | --- |
| AAPM [22] | Report of AAPM Task Group 177 | 2019 | NU 1-2012 [27] |
| ANZSNM [23] | Minimum QC Schedule for Gamma Cameras | 2013 | NU 1-2007 [28] |
| EANM [24] | Routine quality control recommendations for nuclear medicine instrumentation | 2010 | NU 1-2001 [29] NU 1-2007 [28] |
| IAEA [25] | Human Health Series 6 – Quality Assurance for SPECT Systems | 2009 | NU 1-1994 [30], NU 1-2001 [29] |

Table 7 Comparison of recommended QC tests and minimum testing frequencies for gamma cameras and SPECT/CT systems in publications from professional bodies.

| Test Name | Test Description | NEMA Test | AAPM [22] | ANZSNM1 [23] | EANM [24] | IAEA1 [25] |
| --- | --- | --- | --- | --- | --- | --- |
| A. GAMMA CAMERA - PLANAR |  |  |  |  |  |  |
| Intrinsic energy resolution | Characterisation of ability to distinguish between different photon energies, expressed as FWHM of photopeak. |  | Yearly | Yearly | - | - |
| Energy window peaking for Tc-99m | Checking and centring of energy window on Tc-99m photopeak. |  | Daily | Daily | Daily | Daily/Weekly |
| Energy window setting for other radionuclides | Checking and centring of energy windows for other isotopes used clinically. |  | When used | When used | When used | Six-monthly |
| Intrinsic flood field uniformity for Tc-99m | Qualitative and quantitative uniformity assessment without degradation effects from collimators using Tc-99m point source. |  | Daily2 | Daily2 | Daily2 | Weekly |
| Intrinsic flood field uniformity using asymmetric (off-peak) PHA windows for Tc-99m | Qualitative and quantitative uniformity assessment using Tc-99m source with off-peak energy window to check for crystal hydration and unbalance in PMT tuning. |  | Yearly | Yearly | - | Six-Monthly |
| Intrinsic high-count flood field uniformity for Tc-99m | Qualitative and quantitative uniformity assessment using Tc-99m point source with reduced statistical variations effects. |  | - | Weekly | - | - |
| Intrinsic flood uniformity for other radionuclides | Qualitative and quantitative uniformity assessment for other isotopes used clinically. |  | Yearly | - | Quarterly | Six-Monthly |
| Intrinsic spatial resolution & linearity | Visual and/or quantitative assessment of spatial resolution and linearity without degradation effects of collimators. |  | Weekly3 | Monthly4 | Six-Monthly5 | Six-Monthly3 |
| Intrinsic count rate performance | Assessment of count losses due to dead time effects at high counting rates. |  | Yearly | - | - | Six-Monthly |
| System flood field uniformity | Qualitative and quantitative uniformity assessment using Co-57/Tc-99m flood source under clinical conditions. |  | Daily2 | Daily2 | Daily2 | Daily |
| System flood field uniformity for all collimators | Qualitative and quantitative uniformity assessment for other isotopes used clinically. |  | Yearly | System dependent | - | Six-Monthly |
| System spatial resolution and linearity | Spatial resolution and linearity assessment with the degradation effects of collimators. |  | Weekly3 | Six-monthly4 | Six-Monthly5 | - |
| System planar sensitivity | Characterisation of ability to detect incident photons. Typically expressed in counts per megabecquerel per second. |  | Yearly | Six-Monthly | Daily6 | Six-Monthly |
| Multiple window spatial registration | Assessment of ability to resolve different photon energies when more than one PHA is used simultaneously. |  | Yearly (Optional) | Six-Monthly | Six-Monthly/Yearly | Six-Monthly |
| B. GAMMA CAMERA – WHOLE-BODY |  |  |  |  |  |  |
| Whole-body spatial resolution | Characterisation of spatial resolution in directions parallel and perpendicular to motion (with and/or without scatter). |  | Yearly (Optional) | Yearly | Yearly | After Major Repairs |
| Whole-body uniformity | Check of consistency in uniformity and sensitivity irrespective of scan position. |  | Yearly (Optional) | Yearly | Yearly | - |
| C. SPECT |  |  |  |  |  |  |
| SPECT uniformity | Uniformity assessment as detector head rotates over 360 degrees. |  | Quarterly8/Yearly | Yearly | Six-Monthly8 | - |
| SPECT resolution without scatter | Characterisation of spatial resolution in air to check for degradation during image acquisition and reconstruction. |  | Yearly | Six-Monthly9 | Six-Monthly | Six-Monthly |
| SPECT resolution with scatter | Characterisation of spatial resolution under clinical conditions to check for degradation during image acquisition and reconstruction |  | Yearly | Yearly9 | - | Six-Monthly |
| SPECT image quality | Tomographic contrast and quantitative accuracy assessment. |  | Quarterly | - | Six-Monthly10 | Six-Monthly10 |
| Centre of rotation offset and alignment | Check of centre of rotation offsets in X and Y axes, and alignment of detector head tilt in the Y-axis. |  | Monthly11 | Weekly/Monthly12 | Weekly/Monthly13 | Weekly/Monthly14 |
| Variations of uniformity and sensitivity with angle | Characterisation of variations in detector non-uniformity and sensitivity at different angular positions |  | - | Six-Monthly | Six-Monthly10 | - |
| D. SPECT/CT |  |  |  |  |  |  |
| SPECT/CT co-registration | Characterisation of accuracy in registration between SPECT and CT images. |  | As per Vendor | Yearly | Monthly | - |
| SPECT/CT image quality | Uniformity, noise analysis and contrast detectability assessment for CT attenuation corrected SPECT images. |  | Quarterly | Six-Monthly | - | - |
| E. CT |  |  |  |  |  |  |
| CT number accuracy | Check of CT number accuracy for known materials, including water. |  | Daily | As per Vendor/Legislation | Monthly | - |
| CT dose assessment | Dose output assessment using CTDIair and/or CTDIvol measurements. |  | Yearly | As per Vendor/Legislation | As per Legislation | - |
| CT image quality assessment | CT number, noise, uniformity, slice thickness and artefact assessment in acquired CT images. |  | Daily | As per Vendor/Legislation | Daily15 | - |

Co-57 = Cobalt-57; CT = Computed Tomography; CTDIair = CT Dose Index in Air; CTDIvol = Volume CT Dose Index; FHWM = Full Width at Half Maximum; PHA = Pulse-Height Analyser; PMT = Photomultiplier Tube; SPECT = Single Photon Emission Computed Tomography; SPECT/CT Single Photon Emission Computed Tomography/Computed Tomography; Tc-99m = Technetium-99m

1 Minimum recommended testing frequencies listed

2 As an alternative to daily system/intrinsic flood field uniformity

3 As an alternative to system/intrinsic spatial resolution

4 Intrinsic or system spatial resolution not specified.

5 Intrinsic or system spatial resolution, depending on convenience

6 As part of daily intrinsic/extrinsic flood field uniformity tests

7 As part of centre of rotation (COR) verification tests

8 As part of SPECT image quality testing

9 Tomographic resolution with or without scatter not specified

10 As part of overall SPECT system performance

11 Monthly or frequency as recommended by vendor

12 Weekly COR checks if system is unstable

13 Frequency adjusted depending on system stability

14 For a stable system, perform monthly COR verifications

15 As part of daily QC procedures specified by vendors