Supporting Information for

Quantitative Phase Imaging with a Compact Meta-microscope

Junyi Wang1, Rongtao Yu1, Xin Ye1, Jiacheng Sun1, Jian Li1, Chunyu Huang1, Xingjian Xiao1, Jitao Ji1, Wenjing Shen1, Zuoxiu Tie1, Chen Chen1,*, Shining Zhu1, and Tao Li1,*

1National Laboratory of Solid State Microstructures, Key Laboratory of Intelligent Optical Sensing and Manipulation, Jiangsu Key Laboratory of Artificial Functional Materials, College of Engineering and Applied Sciences, Nanjing University, Nanjing 210093, China

*Correspondence should be addressed to Chen Chen (chenchen2021@nju.edu.cn) or to Tao Li (taoli@nju.edu.cn)

Supplementary Note 1. Multi-plane axial intensity derivative approximation method

Before the retrieval of phase at the focal plane z = 0, we first need to approximate the axial intensity derivative of the focal plane, . We can simply use the finite difference between two intensity images to approximate the derivative as

(S1)

where  is the defocus distance,  is the defocus intensity, and  is the in-focus intensity. When  is small enough, the finite difference can work as a viable approximation mathematically. However, small  means small contrast between the defocused image and in-focus image. Under experiment conditions, this may lead to poor SNR and result in the reconstruction of noise artifacts. Large  in turn decreases the accuracy of the finite difference approximation, which attenuates the high-frequency components and blurs the phase reconstruction. When multiple images are available, we can carry out a simple polynomial curve fitting1, as in least squares fitting. Then the estimated analytical polynomial function can be expressed as

(S2)

where n is the order of the polynomial function. Theoretically or under a low-noise level, the nth -order derivative can be approximated by n+1 or more images. However, in practice, the noise may compromise the accuracy of direct intensity fitting, leading to overfitting and creating low frequency artifacts in the retrieved phase profile. An intensity spectrum fitting method2 was proposed to tackle this problem, termed Gaussian process regression TIE (GP-TIE). Under the coherent, paraxial, and weak object approximation, the defocused intensity spectrum can be written as

(S3)

where  is the DC term, and  and  are the Fourier transform of amplitude and phase distributions. In practice, the intensity spectrum will deviate from this ideal sinusoidal oscillation pattern due to the noise and partially coherent illumination conditions. However, this prior knowledge can serve as a restriction to avoid the overfitting of noise. Based on the Gaussian process regression method to fit the axial intensity oscillation to a continuous sinusoidal function, the axial intensity derivative in the frequency domain can be expressed as

(S4)

(S5)

(S6)

where ,  are hyper-parameters and  is the covariance matrix of the output given by the input  and . From the multi-plane measurements, we can derive the intensity derivative with high approximation accuracy based on the GP-TIE method.

Supplementary Note 2. Theoretical comparison between the hyperbolic phase and aplanatic phase under microscopic imaging configuration

According to the general Snell’s law, the phase gradient at the interface of the metalens working at the normalized incidence can be written as

(S7)

Supplementary Figure S1. The parameters of the simulated metalens are diameter D = 0.2 mm, and NA = 0.58, the ideal focal distance is 0.281 mm. The working wavelength is set as 450 nm. a, b Ray tracing simulation results using hyperbolic phase profile under the microscopic imaging configuration of magnification of 4 and 1. c, d Ray tracing simulation results using aplanatic phase profile under the microscopic imaging configuration of magnification of 4 and 1.

where  is the incident wavelength, and r is the radial coordinate, f is the focal length. This is how the normally used hyperbolic phase of the metalens is derived. If the phase modulation follows the PB phase method, then we can assume the phase gradient at the interface is independent of the incident wavelength. From the geometric relation between f and r, we can finally derive the chromatic dispersion equation as

(S8)

where  is the center wavelength and f0 is the corresponding focal length. From Eq. S8, f is both radially and wavelength dependent. This means when the incident wavelength deviates from the center wavelength, the hyperbolic metalens will display spherical aberration, positive spherical aberration when  and negative one when , specifically. Apart from this problem, the hyperbolic phase is derived under normalized incidence, which is not suitable to describe the microscopic imaging conditions where samples are usually very close to the imaging optics making the wavefront more like spherical wavefront. If the hyperbolic metalens is applied directly in microscopic imaging, there will be spherical aberration even at the center wavelength, especially in high numerical aperture (NA) situations. The aplanatic metalens3 is proposed to greatly alleviate the aforementioned problems. We implemented the ray tracing method (Ansys Zemax OpticStudio) to compare the focusing performance of hyperbolic and aplanatic phases under the imaging magnification of 4 and 1. Considering the limited calculation resources, we scaled down the diameter of the metalens to 0.2 mm and kept the NA fixed at 0.58, which resulted in a focal length of 0.28 mm. The working wavelength is set as 450 nm, and the working distance was set to match the magnification configuration of 4 as the situation in the main text. As is shown in Fig. S1c, the metalens with the aplanatic phase designed for the imaging magnification of 4 can focus the incident light perfectly without spherical aberration under the corresponding configuration. When the working distance is changed to a magnification configuration of 1, the focal spot is not as perfect but still falls inside the Airy disk, and the image distance is close to the ideal 0.56 mm, as is shown in Fig. S1d. By contrast, the focal spots of the metalens designed by the hyperbolic phase are all larger than the Airy disk, and the focal distances are not as accurate. This result validates the slight alteration in imaging magnification will not bring significant aberration-induced error to the phase retrieval, and the aplanatic phase is more favorable in microscopic applications.

Reference

1.	Waller, L., Tian, L. & Barbastathis, G. Transport of Intensity imaging with higher order derivatives. Opt Express 18, 12552 (2010).

2.	Jingshan, Z., Claus, R. A., Dauwels, J., Tian, L. & Waller, L. Transport of Intensity phase imaging by intensity spectrum fitting of exponentially spaced defocus planes. Opt Express 22, 10661 (2014).

3.	Chen, C. et al. Spectral tomographic imaging with aplanatic metalens. Light Sci Appl 8, 99 (2019).

Supplementary Figure S2. a Optical setup for the chromatic dispersion characterization. Objective: Mitutoyo 20X, NA = 0.42. b Optical setup for imaging experiments. Objective: Mitutoyo 10X, NA = 0.28. CPF: circularly polarized film. L1: Jcoptix, OLD2440-T2M, f = 7.5 cm. L2: Jcoptix, OLD3266-T2M, f = 20 cm. L3: Jcoptix, OLD2438-T2M, f = 6 cm. FS: field stop. AS: aperture stop. TL: tube lens, Jcoptix, OLD2466-T2M, f = 20 cm. Filters used in the experiments are hard-coated bandpass filters from Thorlabs, Inc.

Supplementary Figure S3. a, b Defocused images of the 4T1 and HeLa cells captured by the meta-microscope at targeted wavelengths. Scale bar: 20 μm.

Supplementary Figure S4. Parameter sweep results of the SiNx nanofins (H = 1000 nm, P = 300 nm) at the center wavelength 450 nm. a, c Simulated transmittance under x and y polarized normal incidence. b, d Simulated phase modulation under x and y polarized normal incidence. e Simulated polarization conversion ratio (PCR) of each nanofin from LCP to RCP. The nanofin (Dx = 90 nm, Dy = 215 nm) with PCR = 99.89% at the center wavelength 450 nm and average PCR = 95% from 420 nm to 480 nm at an interval of 10 nm is chosen (marked by red star).