Supplementary Table 1. Report on image processing and radiomic features extraction

| Acquisition and reconstruction | Acquisition parameters; Siemens Biograph mCT-S (64). Whole-body PET; Whole-body CT; Thin-section chest CT. 18F-FDG activity (MBq)*; 350-550 (3.70-7.77MBq/kg); –; –. Min/bed position; 2.5; –; –. Crystal; LSO; –; –. Reconstruction; TrueX + TOF; –; –. Matrix (pixels); 200 × 200; 512 × 512; 512 × 512. Resolution (mm); 4.0; 0.7; 0.7. Slice thickness (mm); 3.0; 3.0; 1.0. Voxel size (mm3); 4.07 4.07 3.0; 0.98 0.98 3.0; 0.98 0.98 1.0. Slices; –; 64; 64. Voltage (kV); –; 120; 120. Tube current (mA); –; 210 (adjusted by auto mA); 110 (adjusted by auto mA). Reconstruction; Gaussian filtering with a full-width at half maximum of 5.0mm; 2 iterations and 21 subsets; B30f medium smooth; B60f sharp. *Administered activity was calculated according to the European Association of Nuclear Medicine (EANM) guidelines, version 2.0, from February 2015 [22]. |
| --- | --- |
| Approach | The images were analyzed as a volume (3D). |
| Software | LIFEx v7.2.0 (IMIV, CEA, Inserm, CNRS, Univ. Paris-Sud, Université Paris Saclay, CEA-SHFJ, 91400, Orsay, France) |
| Data availability | All the original patient DICOM files are stored in the institutional PACS. |

| Data conversion |  |  |  |
| --- | --- | --- | --- |
| Procedure | None |  |  |
| Image post-acquisition processing |  |  |  |
| Procedure | None |  |  |
| Segmentation |  |  |  |
| ROI | The volume of interest (VOI) included the tumour lesions. Radiomic features were extracted from Whole-body PET images and Whole-body CT images, through the same VOI. |  |  |
| Procedure | The VOIs were defined semi-automatically on whole-body PET images by using the LIFEx package (version 7.2.0, http://www.lifexsoft.org), with a threshold of 40% of the maximum standardized uptake value (SUVmax), and copied to the Whole-body CT images. |  |  |
| Interpolation |  |  |  |
| Voxel dimensions | Not applicable |  |  |
| Image interpolation method | Not applicable |  |  |
| Image intensity rounding | Not applicable |  |  |
| ROI interpolation method | Not applicable |  |  |
| ROI partial volume | Not applicable |  |  |
| Re-segmentation |  |  |  |
| ROI mask criteria | None |  |  |
| Discretization |  |  |  |
| Discretisation method | Whole-body PET: 64 bins from 0 to 25 |  | Whole-body CT: 400 bins from -1000 to 3000 |
| Feature calculation |  |  |  |
| Feature set | Conventional Indices |  |  |
|  | PET | CT |  |
|  | SUVmean | HUmean |  |
|  | SUVmin | Humin |  |
|  | SUVstd | HUstd |  |
|  | SUVmax | HUmax |  |
|  | SUVQ1 | HUQ1 |  |
|  | SUVQ2 | HUQ2 |  |
|  | SUVQ3 | HUQ3 |  |
|  | TLG (mL) |  |  |
|  | First Order Features |  |  |
|  | SHAPE |  |  |
|  | Volume (mL) |  |  |
|  | Sphericity |  |  |
|  | Compacity |  |  |
|  | Histogram (HISTO) |  |  |
|  | Skewness |  |  |
|  | Kurtosis |  |  |
|  | Entropy_log10 |  |  |
|  | Entropy_log2 |  |  |
|  | Uniformity |  |  |
|  | Second Order Features |  |  |
|  | Grey Level Co-occurrence Matrix (GLCM) |  |  |
|  | Entropy_log2 |  |  |
|  | Entropy_log10 |  |  |
|  | Angular Second Moment (ASM) |  |  |
|  | Contrast |  |  |
|  | Dissimilarity |  |  |
|  | Inverse Difference Moment (IDM) |  |  |
|  | Normalised Inverse Difference Moment (NL_IDM) |  |  |
|  | Correlation |  |  |
|  | Grey-Level Run Length Matrix (GLRLM) |  |  |
|  | Short-Runs Emphasis (SRE) |  |  |
|  | Long-Runs Emphasis (LRE) |  |  |
|  | Low Grey-Level Run Emphasis (LGRE) |  |  |
|  | High Grey-Level Run Emphasis (HGRE) |  |  |
|  | Short-Run Low Grey-Level Emphasis (SRLGE) |  |  |
|  | Short-Run High Grey-Level Emphasis (SRHGE) |  |  |
|  | Long-Run Low Grey-Level Emphasis (LRLGE) |  |  |
|  | Long-Run High Grey-Level Emphasis (LRHGE) |  |  |
|  | Grey-Level Non-Uniformity (GLNU) |  |  |
|  | Run Length Non-Uniformity (RLNU) |  |  |
|  | Run Percentage (RP) |  |  |
|  | Neighbourhood Gray-Tone Difference Matrix (NGTDM) |  |  |
|  | Coarseness |  |  |
|  | Contrast |  |  |
|  | Busyness |  |  |
|  | Complexity |  |  |
|  | Strength |  |  |
|  | Gray Level Size Zone Matrix（GLSZM） |  |  |
|  | Small-Zone Emphasis (SZE) |  |  |
|  | Large-Zone Emphasis (LZE) |  |  |
|  | Low Gray-Level Zone Emphasis (LGZE) |  |  |
|  | High Gray-Level Zone Emphasis (HGZE) |  |  |
|  | Small-Zone Low Grey-Level Emphasis (SZLGE) |  |  |
|  | Small-Zone High Grey-Level Emphasis (SZHGE) |  |  |
|  | Large-Zone Low Grey-Level Emphasis (LZLGE) |  |  |
|  | Large-Zone High Grey-Level Emphasis (LZHGE) |  |  |
|  | Grey Level Non-Uniformity (GLNU) |  |  |
|  | Zone Size Non-Uniformity (ZSNU) |  |  |
|  | Zone Percentage (ZP) |  |  |
| Feature parameters | The parameters were calculated according to the formulas described in the software manual (LIFEx v7.2.0) (IMIV, CEA, Inserm, CNRS, Univ. Paris-Sud, Université Paris Saclay, CEA-SHFJ, 91400, Orsay, France) |  |  |