Supplementary Information

Urinary cadmium concentration is associated with the severity and clinical outcomes of COVID-19: a bicenter observational cohort study

Authors: Li-Chung Chiu, Chung-Shu Lee, Ping-Chih Hsu, et al.

Measurements of blood cadmium concentration

Blood cadmium measurements were conducted using inductively coupled plasma mass spectrometry (ICP-MS). Blood specimens were collected in 6 mL plastic blood collection tubes containing K2EDTA as an anticoagulant (BD, Franklin Lakes, NJ, USA), and stored at 4°C. Cadmium concentration was quantified using ICP-MS on a PerkinElmer NexION 350X instrument (Waltham, MA, USA), and analyzed using no-gas mode. Blood specimens (500 μL) were diluted 10 times with 1.5% (w/v) nitric acid (JT Baker, Phillipsburg, NJ, USA) solution containing yttrium as an internal standard. The cadmium and yttrium standards were purchased from AccuStandard (New Haven, CT, USA). A standard calibration curve with a range of 0 to 40 μg/L was created. The calibration curve had a correlation coefficient of more than 0.995. Level 1 control in-house prepared control and level 2 control Seronorm trace elements whole blood control (Sero, Billingstad, Norway) were used and analyzed at the start and end of each analytical run and after every 10 samples. The lower limit of quantitation (LOQ) for blood cadmium with ICP-MS was 0.5 μg/L. Therefore, values below 0.5 μg/L were assigned to this LOQ (i.e., 0.5 μg/L) for analysis.

Measurements of urine cadmium concentration

Urinary cadmium measurements were also conducted using ICP-MS. Urine specimens were collected in 10 mL metal-free plastic collection tubes and stored at 4°C. Urinary cadmium concentration was quantified using ICP-MS on an Agilent 7800 ICP-MS instrument (Santa Clara, CA, USA) and analyzed using no-gas mode. Urine specimens (500 μL) were diluted 10 times with 1.5% (w/v) nitric acid (JT Baker, Phillipsburg, NJ, USA) solution containing yttrium as an internal standard. The standard range was 2.67 to 355.84 nmol/L. The calibration curve had a correlation coefficient of more than 0.995. BIO-RAD Lyphochek® Urine Metal Control Levels 1, 2, and 3 (Hercules, CA, USA) were used and analyzed at the beginning and end of each analytical run; they were then analyzed again after every 10 samples. The lower LOQ for urinary cadmium was 2.67 nmol/L, and values below the LOQ were assigned to this LOQ for analysis. Each urinary cadmium result was corrected using the urinary creatinine level of the same urine sample.