## Sheet1
Source | Location | Study Type | Age, y, Mean (SD) | BMI, Mean (SD) | Sample Size, No. | Males, No. (%) | Wearable devices | Wearable devices |  | Cuffed devices | WBP, Mean±SD, mm Hg |  | CBP, Mean±SD, mm Hg
[8 empty] | measurement sites | measurement technologies |  | SBP1 | DBP1 | SBP2 | DBP2
Ali A. Aboughaly,15 2017 | German University Cairo, Egypt | prospective clinical trial | 44.1(20.2) | NA | 30 | 19 (63.3%) | a Non-Invasive Cuff-less Blood Pressure Monitor (NICBPM) | wrist | PPG and ECG | the automated oscillometric system (model BP101A, cuff-based apparatus) | 118.6±11.19 | NA | 118±12 | 79±14
C.C.Y. Poon,32 2005 | Hong Kong, China | prospective clinical trial | 57(29) | NA | 85 | 36(42%) | a stainless steel electrodes and a pair of light emitting diode and photodetector (IR91-21C and IR91-21B, Everlight, Taiwan) | finger | PPG and ECG | an automatic BP meter (Colin BP-8800, Japan and Omron HEM-907, Japan) | 122±21 | 64±9 | 122±21 | 63±9
Erez Kachel,20 2021(1) | Tiberias, Israel | prospective, comparative clinical trial | 60(15) | 26.9(2.6) | 10 | 6(60%) | a chest-patch | chest | PPG | an AL (AB-0023 Art-LineTM Kits Single Channel, Biometrix, Gronsveld, The Netherlands, and IntelliVue MX500 Patient Monitor, Philips Medical Systems) | 120±15 | 61±9 | 120±14 | 61±8
Erez Kachel,20 2021(2) | Tiberias, Israel | prospective, comparative clinical trial | 60(15) | 26.9(2.6) | 10 | 6(60%) | a wrist watch | wrist | PPG | an AL (AB-0023 Art-LineTM Kits Single Channel, Biometrix, Gronsveld, The Netherlands, and IntelliVue MX500 Patient Monitor, Philips Medical Systems) | 120±15 | 61±9 | 120±14 | 61±8
F. Scalise,36 2019 | Monza, ITALY | prospective clinical trial | NA | NA | 100 | NA | a new cuffless BP device(CL-BP, Rooti RX, Rooti LabsÔ) | NA | ECG | a standard device (C-BP, SpacelabsÔ 90207) | 131.1±15.9 | 80.2±9.7 | 125.4±10.9 | 75±8.3
Fen Miao,26 2020 | Shenzhen, Guangdong, China | prospective clinical trial | NA | NA | 30 | NA | one-channel ECG signals | NA | ECG | A monitoring system (Mindray N12, Shenzhen, China) | 136.85±18.65 | 72.67±5.93 | 137.07±19.54 | 73.42±8.17
Hakuo Takahashi,33 2021 | Mukou City, Kyoto, Japan | prospective clinical trial | 50(11.7) | NA | 86 | 38(44%) | Omron HEM-7600T blood pressure measuring device | upper arm | oscillometry | the standard mercury sphygmomanometer | 129.4±22.24 | 84±15.92 | 130±23.2 | 83±17.2
Jing Liu,25 2020 | Hong Kong, China | prospective clinical trial | 70.2(5.4) | 23.2(3.2) | 22 | 17(77%) | a wearable MWPPG prototype device | finger | MWPPG | a mercury sphygmomanometer with double-tube stethoscope. | 129.24±11.49 | 65.0±7.41 | 127.8±13.4 | 66.0±10.0
Jong Hyeon Ahn,16 2021 | Seoul, South Korea | prospective clinical trial | 66.9(9.1) | 23.6(3.3) | 56 | 33 (58.9%) | a smartwatch (SM-R850, Samsung Electronics, Suwon, Gyeonggi, South Korea) | wrist | PPG | an aneroid sphygmomanometer (reference device, ri-san;Rudolf Riester, Jungingen, Germany) | 118.9±22.76 | 77.5±16.57 | 118.6±23.2 | 76.3±17.2
Joon Ho Moon,27 2020 | Seoul, South Korea | prospective clinical trial | 57.1(17.9) | 25.6(3.9) | 35 | 17(49%) | the InBodyWATCH | both wrists | PPG and ECG | a manual sphygmomanometer (BPBIO220, InBody Co., Ltd.) | 129.4±22.8 | 69.5±11.5 | 127.2±20.9 | 69.7±12.3
Kazuomi Kario,21 2019 | Tochigi, Japan | prospective clinical trial | 66.1(10.8) | 23.4(4.8) | 50 | 30(60%) | a wrist-worn watch-type oscillometric BP monitoring (WBPM) device | wrist | oscillometry | an ambulatory BP monitoring (ABPM) device, A&D TM-2441 | 132.8±17.4 | 78.9±10.3 | 132.0±17.1 | 82.1±9.7
LiangYe Li,24 2020 | Wuhan, China | prospective clinical trial | NA | NA | 17 | 3(18%) | a wearable blood pressure monitoring wristband based on optical fiber | wrist | PTT | the commercial blood pressure (BP) monitor OMRON HEM-8102A (an intelligent electronic sphygmomanometer) | 114±6.66 | 66.35±5.35 | 113.76±7.35 | 66.24±6.03
Mitsuo Kuwabara,34 2019(1) | Shimotsuke, Tochigi, Japan | prospective clinical trial | 58.8(11.6) | NA | 85 | 47(55%) | the Omron HEM-6410T-ZM | wrist | oscillometry | a standard mercury sphygmomanomete | 125.4±18.72 | 76.2±13.15 | 126.3±20.2 | 77.3±14.5
Mitsuo Kuwabara,34 2019(2) | Shimotsuke, Tochigi, Japan | prospective clinical trial | 54.9(12.0) | NA | 85 | 42(49%) | the Omron HEM-6410T-ZL | wrist | oscillometry | a standard mercury sphygmomanomete | 129.7±18.73 | 81±10.84 | 127.3±20.1 | 80.3±12.9
Mitsuo Kuwabara,35 2020 | Shimotsuke, Tochigi, Japan | prospective clinical trial | 54.6(13.3) | NA | 85 | 44(52%) | the Omron HEM-9601T | wrist | oscillometry | a calibrated mercury sphygmomanometer | 126.4±20.89 | 80.8±13.52 | 125.2±22.0 | 79.7±14.6
Sang-Suk Lee,23 2011 | Wonju, Korea | prospective clinical trial | NA | NA | 13 | NA | a cuffless wrist-wearable pulsimeter with a Hall device | wrist | Hall device | electronic or mercury liquid BP meters. | 137.82±2.04 | 68.04±1.61 | 137.31±4.21 | 71.31±4.94
Sheikh Mohammed Shariful Islam,19 2020 | Burwood, Australia | prospective clinical trial | 20.3(5.4) | 23.6(3.3) | 20 | 10(50%) | a wearable BP monitoring device (Tmart Technologies) | NA | NA | an ambulatory BP device (Model TM-2430; A & D Medical Corp) | 125±5 | 77±9 | 126±10 | 75±6
Siraporn Sakphrom,28 2021 | Nakhon Si Thammarat, Thailand | prospective clinical trial | NA | NA | 60 | NA | a smart devices (PRO) | wrist | NA | standard medical equipment of OMRON HEM-7130 | 90.85±10.77 | 61.48±7.82 | 90.63±11.35 | 61.37±7.88
Sugondo Hadiyoso,18 2021 | Bandung, Indonesia | prospective clinical trial | NA | NA | 10 | NA | the commercial wrist digital BP meter from Sunrom | wrist | oscillometry | Omron 7120 | 119.5±17.88 | 80.7±10.57 | 119.9±15.69 | 84.9±11.87
Syed Ghufran Khalid,22 2018 | Chelmsford, UK | prospective clinical trial | NA | NA | 32 | NA | Raw PPG signal waveforms | NA | PPG | noninvasive BP (NIBP) data | 109.1±12.8 | 59.8±7.2 | 108.9±16.8 | 59.2±9.4
Tessa S Schoot,29 2016 | Nijmegen, Netherlands | prospective clinical trial | 54.1(14.5) | NA | 37 | 18（49%） | the Checkme Health Monitor (Viatom Technology) | finger | NA | the validated Vital Signs Monitor 300 series (Welch Allyn, Skaneateles Falls, NY, USA) | 138.4±25.2 | NA | 136.6±21.8 | NA
Xiangyu Fan,17 2018(1) | Hong Kong, China | prospective clinical trial | NA | NA | 12 | NA | a wearable liquid-capsule sensor platform embedded | wrist | LCPS and ECG | cuff-based device | 113.6±3.0 | 71.0±1.5 | 113.4±2.4 | 73.8±2.2
Xiangyu Fan,17 2018(2) | Hong Kong, China | prospective clinical trial | NA | NA | 12 | NA | a wearable liquid-capsule sensor platform embedded | wrist | PPG and ECG | cuff-based device | 113.4±2.5 | 70.9±1.3 | 113.4±2.4 | 73.8±2.2
Yiming Zhang,30 2020 | Hangzhou, China | prospective clinical trial | 27.96(7.12) | NA | 22 | 15(68%) | a cuff-less monitoring device | finger | PPG and ECG | the Omron-8713 electronic sphygmomanometer | 112.492±8.75 | 74.425±7.26 | 112.49±12.23 | 74.42±9.87
Abbreviations: BMI, body mass index (calculated as weight in kilograms divided by height in meters squared); PPG, photoplethysmography; ECG, electrocardiograph; MWPPG, multi-wavelength photoplethysmography; PTT, pulse transit time; LCPS, liquid-capsule pressure sensor; WBP, wearable devices blood pressure; CBP, cuffed devices blood pressure; SBP, systolic blood pressure; DBP, diastolic blood pressure; NA, not applicable; SD, standard deviation.