Supplementary Information

Supplementary Fig. 1 Temperature fields generated with currents of 128 MHz and 100 kHz are equivalent. (a) The curves of temperature change induced by sinusoidal currents of 128 MHz and 100 kHz at different locations in the phantom. Currents were injected through the electrode, the resulting temperature changes was measured by three fluoroptic probes which were fixed at the electrode surface, about 5 mm, and about 10 mm from the electrode centre. For each frequency, experiments were performed at 3 levels of heating, with 3 repetitions for each level. (b) Histogram of the differences in the temperature fields produced by two frequencies of current. The mean temperature during the last 30s of each curve was used. The differences between two heating current were significantly with ± 0.1 °C. Sample size n = 27, analysed with one-tailed t test with significance level of 0.05.

Supplementary Fig. 2 Customized design of the MRI compatible protection apparatus and procedures in the animal study. (a) Detailed structure of the apparatus, including the base, chamber, and cover. The white arrows and the dashed line indicate the typical locations for fixation with ceramic screws. The screws between the base and the chamber can be replaced with metallic screws to increase strength, in which case the metallic screws need to be removed before MRI. (b) Surgical scene of implantation. (c) Overall procedure of the animal study.

Supplementary Fig. 3 Individualized blood perfusion in the animal study. The temperature change measured by fluoroptic probe at the electrode surface during the cooling phase was fitted to the simulated temperature change to determine the perfusion parameter for each experiment. For perfusion parameter in simulations, the range of 2.7025 mL/100g/min to 81.0750 mL/100g/min with a step length of 2.7025 mL/100g/min was traversed. Higher perfusion parameter was excluded due to insignificant improvements and the potential overfitting. The probe measurement for the first 700 s of cooling phase were used.

Supplementary Table 1 Material parameters for electrothermal simulation

| Parameter | Value |
| --- | --- |
| Phantom |  |
| Medium density (kg/m3) | 1000 |
| Medium heat capacity (J/kg/K) | 4200 |
| Medium heat conductivity (W/m/K) | 0.6 |
| Medium electrical conductivity (S/m) | 0.36 |
| Bioheat model |  |
| Tissue density (kg/m3) | 1045.5 |
| Tissue heat capacity (J/kg/K) | 3630 |
| Tissue heat conductivity (W/m/K) | 0.51 |
| Tissue electrical conductivity (S/m) | 0.37 |
| Blood temperature (°C) | 37 |
| Blood density (kg/m3) | 1049.8 |
| Blood heat capacity (J/kg/K) | 3617 |
| Perfusion (mL/100 g/min) — animal | 52.6 |
| Perfusion (mL/100 g/min) — patient | 55.9 |

Supplementary Table 2 Uncertainties of simulation for phantom

| Uncertainty source | Value 1 | Value 2 | Sensitivity | SD | Uncertainty |
| --- | --- | --- | --- | --- | --- |
| Phantom |  |  |  |  |  |
| Medium density (kg/m3) | 1000 | 100.93 | 0.06%/% | 9.3 (0.93%) | 0.05% |
| Medium heat capacity (J/kg/K) | 4200 | 4413 | 0.04%/% | 213.36 (5.08%) | 0.2% |
| Medium heat conductivity (W/m/K) | 0.6 | 0.63 | 0.90%/% | 0.03 (4.78%) | 4.30% |
| Boundary temperature (°C) | 37 | 35 | 4.32%/% | 0.5 (1.35%) | 5.84% |
| Heating duration (s) | 998 | 1008 |  |  | 0.02% |
| Relative tolerance | 0.01 | 0.001 |  |  | 0.03% |
| Min mesh size (mm) | 0.005 | 0.1 |  |  | 0.01% |
| Time step (s) | 0.005~2 | 0.1~5 |  |  | 0.02% |
| Combined |  |  |  |  | 7.26% |

Supplementary Table 3 Uncertainties of simulations for animal and patient subject

| Uncertainty source | Value 1 | Value 2 | Sensitivity | SD | Uncertainty |
| --- | --- | --- | --- | --- | --- |
| Bioheat model |  |  |  |  |  |
| Medium density (kg/m3) | 1041 | 1050 | 0.03%/% | 6.36 (0.61%) | 0.02% |
| Medium heat capacity (J/kg/K) | 3578 | 3682 | 0.03%/% | 73.54 (2.03%) | 0.06% |
| Medium heat conductivity (W/m/K) | 0.49 | 0.54 | 0.71%/% | 0.022 (4.3%) | 3.10% |
| Blood density (kg/m3) | 1025 | 1060 | 0.18%/% | 16.56 (1.58%) | 0.28% |
| Blood heat capacity (J/kg/K) | 3300 | 3900 | 0.17%/% | 301.44 (8.33%) | 1.42% |
| Perfusion — animal (ml/100g/min) | 38.29 | 66.87 | 0.18%/% | 14.29 (27.17%) | 4.89% |
| Perfusion — patient (ml/100g/min) | 45.98 | 65.74 | 0.18%/% | 9.88(17.69%) | 3.18 % |
| Blood temperature (°C) | 37 | 36.7 | 19.07%/% | 0.3 (0.81%) | 15.46% |
| Boundary temperature (°C) | 37 | 35 | 0.03%/% | 0.5 (1.35%) | 0.04% |
| Heating duration — animal (s) | 365 | 370 |  |  | 0.06% |
| Heating duration — patient (s) | 180 | 185 |  |  | 0.15% |
| Convergence: relative tolerance | 0.01 | 0.001 |  |  | 0.01% |
| Min mesh size (mm) | 0.005 | 0.1 |  |  | 0.01% |
| Time step (s) | 0.005~2 | 0.1~5 |  |  | 0.01% |
| Combined — animal |  |  |  |  | 16.57% |
| Combined — patient |  |  |  |  | 16.15% |