## Overview
Overview of sample conditions, symptom status and extracted and used data for meta-analysis
|  | Used Data | [11 empty] | Extracted Data | [6 empty] | Type of RT-PCR used (LOD as per manufacturer's instructions [if publicly available]) | RT-PCR sample type | Persons tested | Predominant SARS-CoV-2 variant
|  | Author | Study location | Sample type | IFU conform | Sample condition | Sample size | Sensitivity | lower 95% CI | upper 95% CI | Specificity | lower 95% CI | upper 95% CI | Sample size | Sensitivity | lower 95% CI | upper 95% CI | Specificity | lower 95% CI | upper 95% CI
| Abbott, BinaxNOW (LFA) | ADVIA Centaur SARS-CoV 2 Ag, Siemens Healthineers (lab-based, chemiluminescent immuoassay; LOD:624 TCID50/ML )
f133.1 |  | Peck Palmer, s35 | USA | NP | yes | frozen | 347 | 85,3* | 80,5* | 89,3* | 97.3 | 90.7 | 99.7 | 347 | 85.6 | 81.5 | 89.1 | 97.3 | 90.7 | 99.7 | XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha, Delta
| Abbott, Panbio (nasal sampling, LFA) | AFIAS COVID-19 Ag, Boditech Med Inc. (POC, fluorescence immunoassay, reguires reader; LOD: N/A)
f42.1 |  | Baccani, f165.3 | Italy | NP | no | unclear | 81 | 37.5 | 18,8* | 59,4* | 100 | 93,7* | 100 | 81 | 37.5 | 20.4 | 57.4 | 100 | 93.7 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | unclear | unclear
|  | Lanari, K9.2 | Italy | NP | no | fresh | 765 | 86.4 | 75 | 93.9 | 98.3 | 97.1 | 99.1 | 765 | 86.4 | 75 | 93.9 | 98.3 | 97.1 | 99.1 | ELITe InGenius (LOD: N/A); XpertXpress (LOD: 0.02 PFU/ml); Simplexa (LOD: 500 copies/ml) | AN | symptomatic, high risk contact, screening independent of symptoms | wildype, Alpha
Atellica Solution (lab-based, chemiluminescen immunoassay; LOD: N/A)
|  | Hörber, S20.1 | Germany | AN/MT | yes | fresh | 447 | 88.5 | 83.7 | 92,3* | 99.5 | 97.4 | 99.9 | 447 | 88.5 | 83.7 | 91.9 | 99.5 | 97.4 | 99.9 | TaqPath (LOD: 1.98 * 10^3 NDU/ml); | NP | symptomatic, high risk contact, screening independent of symptoms | Omicron
| Becton, Dickinson and Company, BD Antigen Test (LFA) | BD Veritor, Becton, Dickinson and Company (LFA; requires reader; LOD as per manufacturer's instructions: 1.4 * 10^2 TCID50/ml)
f45.1 |  | Pekosz, a28.1 | USA | NP | yes | frozen | 251 | 96.4 | 81,7* | 99,9* | 98.7 | 96.1 | 99.7 | 251 | 96.4 | 82.3 | 99.4 | 98.7 | 96.1 | 99.7 | Lyra (LOD: 0.8 copies/µl) | NP | symptomatic | wildtype
|  | Van de Moeren, a39.1 | Netherlands | NP | yes | frozen | 351 | 94.1 | 71,3* | 99,9* | 100 | 98.9 | 100 | 351 | 94.1 | 71.1 | 100 | 100 | 98.9 | 100 | Cobas (LOD: 0.009 TCID50/ml); RealTime (LOD: 200 copies/ml) | NP | symptomatic | wildtype
|  | Young, a43.1 | USA | NP | yes | frozen | 251 | 76,3* | 59,8* | 88,6* | 99,5* | 97,4* | 99,9* | 251 | 76.3 | 60.8 | 87 | 99.5 | 97.4 | 99.9 | Lyra (LOD: 0.8 copies/µl) | NP | symptomatic | wildtype
|  | Caruana, f34.1 | Switzerland | NP | yes | stored | 532 | 41.2 | 32,1* | 50,8* | 99,8* | 98,7* | 100* | 532 | 41.2 | # | # | 99.7 | # | # | Cobas (LOD: 0.009 TCID50/ml); GeneXpert (LOD: 0.02 PFU/ml); VIASURE (LOD: >=10 RNA copies/reaction) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Marti, f46.2 | USA | AN/MT | yes | unclear | 32 | 93.75 | 79,2* | 99,2* | # | # | # | 32 | 93.75 | # | # | 100 | # | # | XpertXpress (LOD: 0.02 PFU/ml) | AN/MT | symptomatic | wildtype
|  | Kilic, f71.1 | USA | AN/MT | yes | fresh | 1384 | 66.4 | 57 | 74.9 | 98.8 | 98.1 | 99.3 | 1384 | 66.4 | 57 | 74.9 | 98.8 | 98.1 | 99.3 | Simplexa (LOD: 500 copies/ml) | NP | symptomatic | wildtype
|  | Yin, f82.1 | Belgium | NP | yes | fresh | 177 | 87.7 | 80 | 92.7 | # | # | # | 177 | 87.7 | 80 | 92.7 | # | # | # | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic | wildtype
|  | Koeleman, f103.1 | Netherlands | NP | yes | stored, frozen | 40 | 77.5 | 61,5* | 89,2* | # | # | # | 40 | 77.5 | 61.1 | 88.6 | # | # | # | ELITe InGenius (LOD: N/A); inhouse (LOD: N/A); XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic | wildtype
|  | Christensen, f109.1 | USA | AN/MT | no | fresh | 278 | 86.7 | 75,4* | 94,1* | 99.5 | 97.4 | 99.9 | 278 | 86.7 | 75.8 | 93.1 | 99.5 | 97.4 | 99.9 | Lyra (LOD: 0.8 copies/µl) | NP | symptomatic | wildtype
|  | Frediani, f139.1 | USA | NP | no | frozen | 48 | 50* | # | # | # | # | # | 48 | 50 | # | # | # | # | # | Simplexa (LOD: 500 copies/ml) | NP | unclear | wildtype
|  | Karon, f148.1 | USA | NP | yes | frozen | 347 | 66.5 | 59,4* | 73* | 100 | 97.5 | 100 | 347 | 66.5 | 59.6 | 72.7 | 100 | 97.5 | 100 | TibMolBiol (LOD: 1-10 copies/reaction) | NP | symptomatic, screening independent of symptoms | unclear
|  | Wee, k17.1 | Singapore | unclear | no | fresh | 6665 | 93.6 | 78.6 | 99.2 | 99.8 | 99.7 | 99.9 | 6665 | 93.6 | 78.6 | 99.2 | 99.8 | 99.7 | 99.9 | unclear | unclear | symptomatic, high risk contact, screening independent of symptoms | Delta
|  | Greub, s16.3 | Switzerland | NP | yes | stored at 4°C | # | 71* | 61,1* | 79,6* | 99,5* | 97,2* | 100* | # | 70.9 | 58.1 | 81.8 | 99.7 | # | # | Cobas (LOD: 0.009 TCID50/ml); VIASURE (LOD: >=10 RNA copies/reaction); GeneXpert (LOD: 0.02 PFU/ml); Allplex (LOD: 4,167 copies/ml); inhouse (LOD: N/A) | NP | unclear | wildtype
|  | Navero-Castillejos, s31.1 | Spain | NP | yes | frozen | 130 | 77,5* | 66,8* | 86,1* | 100 | # | # | 130 | 77.22 | # | # | 100 | # | # | XpertXpress (LOD: 0.02 PFU/ml); TaqPath (LOD: 1.98 * 10^3 NDU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Neukam, s32.2 | Argentina | NP | yes | fresh | 584 | 66,7* | 48,2* | 82* | 99.8 | 99 | 100 | 584 | 74.4 | 59.8 | 85.1 | 99.8 | 99 | 100 | RealStar (LOD: 10 NDU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | Gamma
|  | Robinson, s37.1 | USA | AN/MT | yes | frozen | 99 | # | # | # | # | # | # | 99 | # | # | # | # | # | # | RealStar (LOD: 10 NDU/ml); Cobas (LOD: 0.009 TCID50/ml); PerkinElmer (LOD: 30.467 copies/ml); Aptima (LOD: 0.01 TCID50/ml); NeuMoDx (LOD: 150 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Wells, s49.1 | USA | unclear | no | unclear | 409 | # | # | # | # | # | # | 409 | # | # | # | # | # | # | unclear | unclear | screening independent of symptoms | wildtype
| Dräger, Antigen Test SARS-CoV-2 (LFA) | Bioeasy 2019-nCov Antigen Rapid Test Kit,Shenzhen Bioeasy Biotechnology (LFA; requires reader; LOD not publicly available from manufacturer)
f79.1 |  | Krüger, a17.1 | Germany | NP | yes | fresh | 727 | 66.7 | 38,4* | 88,2* | 93.1 | 91* | 94,9* | 727 | 66.7 | 41.7 | 84.9 | 93.1 | 91 | 94.8 | Cobas (LOD: 0.009 TCID50/ml); TibMolBiol (LOD: 1-10 copies/reaction); Allplex (LOD: 4,167 copies/ml); Genesig (LOD: 0.33 copies/µl); RealTime (LOD: 200 copies/ml) | NP | symptomatic, high risk contact | wildtype
|  | Porte, a31,1 | Chile | NP | yes | stored | 127 | 93.9 | 86,3* | 98* | 100 | 92,1* | 100* | 127 | 93.9 | 86.5 | 97.4 | 100 | # | # | RealTime (LOD: 200 copies/ml); Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic | wildtype
|  | Weitzel, a41.1 | Chile | NP | yes | frozen | 111 | 85 | 75,3* | 92* | 100 | 88,8* | 100* | 111 | 85 | 75.6 | 91.2 | 100 | 89 | 100 | Simplexa (LOD: 500 copies/ml) | NP | symptomatic | wildtype
|  | Parada-Ricart, a48.1 | Spain | NP | yes | fresh | 172 | 73,1* | 52,2* | 88,4* | 85,6* | 78,9* | 90,9* | 172 | # | # | # | # | # | # | Lyra (LOD: 0.8 copies/µl) | NP | unclear | wildtype
|  | Krüger, s25.1 | Germany | NP | yes | fresh | 727 | 66.7 | 38,4* | 88,2* | 93.1 | 91 | 94,9* | 727 | 66.7 | 41.7 | 84m8 | 93.1 | 91 | 94.8 | TibMolBiol (LOD: 1-10 copies/reaction); Allplex (LOD: 4,167 copies/ml); Cobas (LOD: 0.009 TCID50/ml); RealTime (LOD: 200 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
| Becton, Dickinson and Company, BD Antigen Test (LFA) | CoviNag-ELISA, XEMA medica (lab-based, Ezyme-linked immunosorbent assay; LOD: 40 pg/ml)
|  | Mayanskiy, f188.1 | Russia | NP | no | frozen | 277 | 90.1 | 84.8 | 94 | 75.8 | 65.9 | 84 | 277 | 90.1 | 85.1 | 93.8 | 75.8 | 66.5 | 83.5 | Multiplex (LOD: 10 copies NA/tube) | NP | unclear | wildtype
Elecsys® Anti-SARS-CoV-2-Antigen-Assays, Roche Diagnostics GmbH (lab-based, Electrochemiluminescence immunoassay; LOD: 22,5 TCID50/ml)
|  | Nörz, f78.1 | Germany | NP | no | stored | 3139 | 60.2 | 55.2 | 65.1 | 99.9 | 99.6 | 100 | 3139 | 60.2 | 55.2 | 65.1 | 99.9 | 99.6 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | unclear | wildtype
|  | Ostermann, f157.4 | Germany | unclear | yes | stored | 408 | 31.42 | 22,7* | 41,2* | 97.69 | 95.31 | 99,1* | 408 | 31.42 | 23.34 | 40.83 | 97.69 | 95.31 | 98.88 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | unclear | Alpha
|  | Iqbal, f171.1 | Pakistan | NP | no | stored | 170 | 72.4 | 63,8* | 80* | 97,7* | 87,7* | 99,9* | 170 | 72.4 | 62.4 | 78.8 | 97.7 | 87.7 | 99.9 | TibMolBiol (LOD: 1-10 copies/reaction) | AN/MT | symptomatic | Alpha
|  | Audige, f184.1 | Switzerland | NP | yes | frozen | 407 | 87.2 | 81.5 | 91.6 | 100 | 98.3 | 100 | 407 | 87.2 | 81.5 | 91.6 | 100 | 98.3 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Audige, f184.2 | Switzerland | saliva | yes | frozen | 307 | 40.2 | 29.9 | 51.3 | 99.5 | 97.4 | 100 | 307 | 40.2 | 29.9 | 51.3 | 99.5 | 97.4 | 100 | Cobas (LOD: 0.009 TCID50/ml) | saliva | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Ben Abdelhanin, k4.1 | Belgium | NP | yes | stored | 225 | 65.85 | 56.76 | 74.16 | 100 | 96.49 | 100 | 225 | 65.85 | 56.76 | 74.16 | 100 | 96.49 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic, screening independent of symptoms | Alpha
|  | Montalvo Villalba, k13.1 | Cuba | NP | yes | fresh | 132 | 82,5* | 70.9 | 90.9 | 97.06 | 89.9 | 99.5 | 132 | 82.81 | 71.79 | 90.12 | 97.06 | 89.9 | 99.46 | STAT-NAT (LOD: 10 copies/reaction) | NP | high risk contact | wildtype
|  | Montalvo Villalba, k13.2 | Cuba | NP | yes | fresh | 59 | 79.41 | 62,1* | 91,3* | 92 | 74* | 99* | 59 | 79.41 | 63.2 | 89.65 | 92 | 75.03 | 98.58 | STAT-NAT (LOD: 10 copies/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Montalvo Villalba, k13.4 | Cuba | NP | yes | fresh | 119 | 78.43 | 64,7* | 88,7* | 86.76 | 76,4* | 93,8* | 119 | 78.43 | 65.37 | 87.51 | 86.76 | 76.72 | 92.88 | STAT-NAT (LOD: 10 copies/reaction) | NP | screening independent of symptoms | wildtype
|  | Kolesova, s24.1 | Italy | NP | no | fresh | 110 | 84,2* | 74* | 91,6* | 100 | # | # | 110 | 85 | # | # | 100 | # | # | MutaPLEX® (LOD: 10 copies/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | unclear
|  | Lau, s26.1 | Singapore | NP | yes | frozen | 350 | 42.6 | 30 | 55.9 | 99.7 | 98.1 | 100 | 350 | 42.6 | 30 | 55.9 | 99.7 | 98.1 | 100 | Cobas (LOD: 0.009 TCID50/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | Delta
|  | Salvagno, s40.1 | Italy | NP | no | fresh | 160 | 42,7* | 33.3 | 52.5 | 100 | 93 | 100 | 160 | 43 | 33.3 | 52.5 | 100 | 93 | 100 | RealStar (LOD: 10 NDU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | Delta
| Innova Medical Group, INNOVA SARS-CoV-2 Antigen Rapid Qualitative Test (LFA) | Euroimmun SARS-CoV-2 Ag ELISA, EUROIMMUN Medizinische Labordiagnostika AG (lab-based, Enzyme-linked immunosorbent assay; LOD: 3.12x10^22 TCID50/ml)
f131.1 |  | Ostermann, f157.3 | Germany | unclear | yes | stored | 410 | 17.76 | 11 | 26.3 | 99.67 | 98.15 | 99.98 | 410 | 17.76 | 11.57 | 26.08 | 99.67 | 98.15 | 99.98 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | unclear | Alpha
|  | Miłosz, s29.1 | Poland | NP | yes | fresh | 137 | 100 | # | # | 98.84 | # | # | 137 | 100 | # | # | 98.84 | # | # | EURORealTime (LOD: 150 cp/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
| SD Biosensor/Roche, Standard Q (nasal sampling, LFA) | Exdia COVID-19 Ag, Precision Biosensor (POC; requires reader; LOD: 1.51 x 10^6 TCID50/mL)
|  | Caruana, f34.2 | Switzerland | NP | yes | stored | 532 | 48.3 | 38,8* | 57,8* | 99.5 | 98,3* | 99,9* | 532 | 48.3 | # | # | 99.5 | # | # | Cobas (LOD: 0.009 TCID50/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Greub, s16.2 | Switzerland | NP | yes | stored | # | 81* | 71,9* | 88,2* | 99.5 | # | # | # | 80.6 | 68.6 | 89.6 | 99.5 | # | # | Cobas (LOD: 0.009 TCID50/ml); VIASURE (LOD: >=10 RNA copies/reaction); GeneXpert (LOD: 0.02 PFU/ml); Allplex (LOD: 4,167 copies/ml); inhouse (LOD: N/A) | NP | unclear | wildtype
| SD Biosensor/Roche, Standard Q (nasal sampling, LFA) | FIAflex™ Fluorescent Immunoassay Analyzer, ACON Biotech (POC, fluorescence immunoassay; LOD not publicly available)
|  | Karon, f148.3 | USA | NP | yes | frozen | 347 | 88.3 | 83 | 92.5 | 100 | 97.5 | 100 | 347 | 88.3 | 83.1 | 92.1 | 100 | 97.5 | 100 | TibMolBiol (LOD: 1-10 copies/reaction) | NP | symptomatic, screening independent of symptoms | unclear
FREND COVID-19 Ag, NanoEntek (POC; requires reader; LOD as per manufacturer's instructions: 1.47 * 10^2 TCID50/ml)
|  | Bruzzone, f86.3 | Italy | unclear | yes | stored | 60 | 93.3 | 83,8* | 98,2* | # | # | # | 60 | 93.3 | 84.1 | 97.4 | # | # | # | Allplex (LOD: 4,167 copies/ml) | unclear | unclear | wildtype
|  | Orsi, f191.1 | Italy | NP | no | fresh | 110 | 93.3 | 83.8 | 98.2 | 100 | 92.9 | 100 | 110 | 93.3 | 83.8 | 98.2 | 100 | 92.9 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic | wildtype
Gazelle, Hemex Health (POC, fluorescence immunoassay; LOD not publicly available)
|  | Mehta, s28.1 | India, USA | AN/MT | yes | fresh | 1486 | 96.3 | # | # | 99.7 | # | # | 1486 | 96.3 | # | # | 99.7 | # | # | TaqPath (LOD: 1.98 * 10^3 NDU/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | Delta
HISCL SARS-COV-2 AG ASSAY KIT, Sysmex Europe (lab-based, chemiluminescence enzyme immunoassay; LOD not publicly available)
|  | Linssen, s27.2 | Netherlands | NP | yes | frozen | 541 | 75 | 47.6 | 92.7 | 99.3 | 98.1 | 99.8 | 541 | 75 | 47.6 | 92.7 | 99.3 | 98.1 | 99.8 | TibMolBiol (LOD: 1-10 copies/reaction); Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | Alpha, Delta
|  | Linssen, s27.3 | Netherlands | NP | yes | frozen | 969 | 90.9 | 80 | 97 | 83.3 | 80.7 | 85.6 | 969 | 90.9 | 80 | 97 | 83.3 | 80.7 | 85.6 | TibMolBiol (LOD: 1-10 copies/reaction); Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | Alpha, Delta
|  | Saito, s39.1 | Japan | NP | yes | frozen | 115 | # | # | # | # | # | # | 115 | # | # | # | # | # | # | QIAamp (LOD: N/A) | NP | unclear | unclear
iChroma COVID-19 Ag Test, Boditech Medical (POC; fluorescence immunoassay; LOD not publicly available from manufacturer)
|  | FINDx, f39.1 | Switzerland | NP | yes | fresh | 232 | 73.2 | 57,1* | 85,8* | 100 | 98 | 100 | 232 | 73.2 | 58.1 | 84.3 | 100 | 98 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
LIAISON® SARS-COV-2 AG, DiaSorin (lab-based, Chemiluminescent Immunoassay; LOD: 22.0 TCID50/ml)
|  | Lefever, f70.1 | Belgium | NP | yes | frozen | 414 | 67,6* | 60,8* | 74* | 100 | 98,3* | 100 | 414 | 67.7 | # | # | 100 | # | # | VIASURE (LOD: >=10 RNA copies/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Van der Moeren, f144.1 | Netherlands | NP | yes | fresh | 248 | 73 | 61,4* | 82,6* | 100 | 97.9 | 100 | 248 | 73 | 61.3 | 82.7 | 100 | 97.9 | 100 | Alinity (LOD: 100 copies/ml) | NP | symptomatic | Delta
|  | Ostermann, f157.2 | Germany | unclear | yes | frozen | 410 | 19.63 | 12,6* | 28,4* | 99.67 | 98.15 | 99.98 | 410 | 19.63 | 13.21 | 28.15 | 99.67 | 98.15 | 99.98 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | unclear | Alpha
|  | Salvagno, f159.1 | Italy | NP | no | fresh | 421 | 30,9* | 25,7* | 36,5* | 100 | 97 | 100 | 421 | 31 | 25 | 36 | 100 | 97 | 100 | RealStar (LOD: 10 NDU/ml) | NP | unclear | Alpha
|  | Hartard, f166.1 | France | NP | no | stored | 378 | 84.8 | 71,1* | 93,7* | 99.4 | 97,8* | 99,9* | 378 | 84.8 | 74.4 | 95.2 | 99.4 | 98.6 | 100 | unclear | NP | symptomatic, screening independent of symptoms | alpha, Beta
|  | Alghounaim, s3.1 | Kuwait | AN/MT | yes | stored, frozen | 897 | 43.4 | 30.6 | 56.8 | 99.9 | 99.3 | 100 | 897 | 43.4 | 30.6 | 56.8 | 99.9 | 99.3 | 100 | TaqPath (LOD: 1.98 * 10^3 NDU/ml); | AN/MT | unclear | unclear
|  | Fernández-Rivas, s23.2 | Spain | AN/MT | yes | fresh | 732 | 81.48 | 74.6 | 87.1 | 100 | 99.4 | 100 | 732 | 81.48 | 74.6 | 87.1 | 100 | 99.4 | 100 | Allplex (LOD: 4,167 copies/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Petonnet, s36.1 | France | NP | yes | fresh | 242 | 63,2* | 52,2* | 73,3* | 100 | 97.6 | 100 | 242 | 63.2 | 53.1 | 73.4 | 100 | 97.6 | 100 | Cobas (LOD: 0.009 TCID50/ml), Simplexa (LOD: 500 copies/ml), BioFire (LOD: 2.2E-02 TCID50/ml), NeuMoDx (LOD: 150 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Uster, s48.1 | Switzerland | NP | yes | frozen | 208 | 75.8 | 68.4 | 82.2 | 100 | 92.5 | 100 | 208 | 75.8 | 68.4 | 82.2 | 100 | 92.5 | 100 | Cobas (LOD: 0.009 TCID50/ml); Liat System (LOD: 0.012 TCID50/ml) | NP | unclear | wildtype, Alpha
|  | Uster, s48.2 | Switzerland | NP | yes | frozen | 137 | 76.5 | 67.8 | 83.8 | 100 | 81.5 | 100 | 137 | 76.5 | 67.8 | 83.8 | 100 | 81.5 | 100 | Cobas (LOD: 0.009 TCID50/ml); Liat System (LOD: 0.012 TCID50/ml) | NP | unclear | Omicron
Lumipulse® G SARS-CoV-2 Ag, Fujirebio (chemiluminescent enzyme immunoassay; LOD as per manufacturer's instructions: 0.47 log10 TCID50/ml)
|  | Hirotsu, f47.1 | Japan | NP | yes | stored | 1029 | 92.5 | 79,6* | 98,4* | 100* | 99,6* | 100* | 1029 | 92.5 | # | # | 100 | # | # | StephOnePlus Real Time PCR (LOD: unclear) | NP | symptomatic, high risk contact | unclear
|  | Gili, f57.2 | Italy | NP | yes | fresh | 1738 | 100 | 96* | 100* | 92.1 | 90,7* | 93,4* | 1738 | 100 | # | # | 92.1 | # | # | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Gili, f57.1 | Italy | NP | yes | stored | 226 | 90.5 | 82,8* | 95,6* | 91.6 | 85,5* | 95,7* | 226 | 90.5 | # | # | 91.6 | # | # | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic, high risk contact | wildtype
|  | Asai, f74.1 | Japan | saliva | yes | fresh | 305 | 77.8 | 65,5* | 87,3* | 98.3 | 95,8* | 99,5* | 305 | 77.8 | # | # | 98.3 | # | # | BD MAX system (LOD: 640 copies/ml) | saliva | symptomatic, high risk contact | wildtype
|  | Matsuzaki, f124.2 | Japan | NP | yes | fresh | 128 | 87.8 | 75,2* | 95,4* | 98.7 | 93,1* | 100* | 128 | 87.8 | # | # | 98.7 | # | # | QuantiTect (LOD: N/A) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Ostermann, f157.1 | Germany | unclear | yes | stored | 410 | 52.34 | 42,5* | 62,1* | 97.03 | 94,4* | 98,6* | 410 | 52.34 | 42.96 | 61.55 | 97.03 | 94.45 | 98.43 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | unclear | Alpha
|  | Baccani, f165.1 | Italy | NP | no | unclear | 201 | 87.9 | 71,8* | 96,6* | 95.8 | 91,6* | 98,3* | 201 | 87.9 | 73.7 | 95.8 | 95.8 | 92 | 98.1 | Allplex (LOD: 4,167 copies/ml) | NP | unclear | unclear
|  | Tanimoto, f198.1 | Japan | NP | yes | stored | 68 | 80 | 64,4* | 90,9* | 96.4 | 81,7* | 99,9* | 68 | 80 | 65.2 | 89 | 96.4 | 82.3 | 99.8 | StephOnePlus Real Time PCR (LOD: unclear) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Tanimoto, f198.2 | Japan | saliva | yes | stored | 116 | 61.8 | 47,7* | 74,6* | 100 | 94.1 | 100 | 116 | 61.8 | 48.6 | 73.5 | 100 | 94.1 | 100 | StephOnePlus Real Time PCR (LOD: unclear) | saliva | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Adreani, k3.1 | France | NP | yes | stored | 239 | 87.6 | 81.7 | 93.5 | 72.7 | 63.4 | 82 | 239 | 87.6 | 81.7 | 93.5 | 72.7 | 63.4 | 82 | LightCycler480 (LOD: N/A); Alinity m (LOD: 100 copies/ml) | NP | unclear | unclear
|  | Loconsole, k11.1 | Italy | NP | no | unclear | 911 | 94.9 | 91,1* | 97,4* | 97.4 | 95,9* | 98,5* | 911 | 94.9 | 91.9 | 97 | 97.4 | 96.5 | 98.1 | TaqPath (LOD: 1.98 * 10^3 NDU/ml); | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Sberna, k15.1 | Italy | NP | yes | fresh | 513 | 84 | 79.1 | 88.1 | 89.1 | 84.4 | 92.7 | 513 | 84 | 79.1 | 88.1 | 89.1 | 84.4 | 92.7 | Allplex (LOD: 4,167 copies/ml); Easy (LOD: N/A) | NP | unclear | unclear
|  | Hirotsu, s18.1 | Japan | NP | yes | unclear | 154 | 92.3 | 84,8* | 96,9* | 100 | 94,3* | 100 | 154 | 92.3 | 89.2 | 92.3 | 100 | 95.5 | 100 | XpertXpress (LOD: 0.02 PFU/ml); StephOnePlus Real Time PCR (LOD: unclear) | NP | unclear | unclear
|  | Nomoto, s33.1 | Japan | AN/MT | yes | frozen | 128 | 90,6* | 75* | 98 | 100 | 96,2* | 100 | 128 | 91 | 75 | 98 | 100 | 96 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic | wildtype, Alpha, Delta
|  | Nomoto, s33.2 | Japan | NP | yes | frozen | 128 | 96,9* | 83,8* | 99,9* | 100 | 96,2* | 100 | 128 | 97 | 84 | 100 | 100 | 96 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic | wildtype, Alpha, Delta
|  | Salvetti, s41.1 | Italy | NP | no | fresh | 2659 | 91,4* | 84,7* | 95,8* | 93.1 | 92 | 94 | 2659 | 91.4 | 84.3 | 95.6 | 93.1 | 92 | 94 | XpertXpress (LOD: 0.02 PFU/ml); ELITe MGB (LOD: 714 gEq/ml) | NP | screening independent of symptoms | Alpha
| LumiraDx, COVID-19 SARS-CoV-2 Antigen Test (automated immunoassay; requires reader) | LumiraDx COVID-19 SARS-CoV-2 Antigen Test, LumiraDx UK Ltd (POC, Microfluidic fluorescence immunoassay; LOD as per manufacturer's instructions: 32 TCID50/ml)
f58.1 |  | Kohmer, f32.1 | Germany | NP | yes | fresh | 100 | 50 | 38.1 | 61.9 | 100 | 86.8 | 100 | 100 | 50 | 38.1 | 61.9 | 100 | 86.8 | 100 | Cobas (LOD: 0.009 TCID50/ml) | NP | unclear | wildtype
|  | Drain, f43.1 | UK/USA | AN/MT | yes | fresh | 257 | 97.6 | 91,6* | 99,7* | 96.6 | 92,6* | 98,7* | 257 | 97.6 | 91.6 | 99.3 | 96.6 | 92.7 | 98.4 | Cobas (LOD: 0.009 TCID50/ml) | AN/MT | symptomatic, screening independent of symptoms | wildtype
|  | Drain, f43.2 | UK/USA | NP | yes | fresh | 255 | 97.5 | 86,8* | 99,9* | 97.7 | 94.7 | 99,2* | 255 | 97.5 | 87.1 | 99.6 | 97.7 | 94.7 | 99 | Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, screening independent of symptoms | wildtype
f58.1 |  | Krüger, f58.1 | Germany | AN/MT | yes | fresh | 761 | 82.2 | 75* | 88* | 99.3 | 98.3 | 99.7 | 761 | 82.2 | 75.2 | 87.5 | 99.3 | 98.3 | 99.7 | Allplex (LOD: 4,167 copies/ml); Cobas (LOD: 0.009 TCID50/ml); TibMolBiol (LOD: 1-10 copies/reaction) | NP | symptomatic, high risk contact | wildtype
|  | Micocci, f77.1 | UK | NP | yes | fresh | 239 | 75* | 34,9* | 96,8* | 96,1* | 92,7* | 98,2* | 239 | # | # | # | # | # | # | unclear | NP | symptomatic, screening independent of symptoms | unclear
|  | Bruzzone, f86.2 | Italy | unclear | yes | stored | 23 | 69.6 | 47,1* | 86,8* | # | # | # | 23 | 69.6 | 49.3 | 84.4 | # | # | # | Allplex (LOD: 4,167 copies/ml) | unclear | unclear | wildtype
|  | Bianco, f102.1 | Italy | AN/MT | yes | fresh | 907 | 90.3 | 86.3 | 93.4 | 92.1 | 89.7 | 94.1 | 907 | 90.3 | 86.3 | 93.4 | 92.1 | 89.7 | 94.1 | XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Caramello, f108.1 | Italy | NP | no | fresh | 149 | 86.6 | 78,2* | 92,7* | 88.5 | 76,6* | 95,6* | 149 | # | # | # | # | # | # | unclear | NP | symptomatic, high risk contact | wildtype
|  | Blairon, f113.1 | Belgium | NP | no | stored | 198 | 61.1 | 52,8* | 68,9* | 100 | 92,7* | 100 | 198 | 61.1 | 53.3 | 68.9 | 100 | 100 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | unclear | Alpha
|  | Dierks, f119.1 | Germany | NP | yes | fresh | 444 | 45.45 | 20.22 | 73.26 | 99.54 | 98.17 | 99.88 | 444 | 45.45 | 20.22 | 73.26 | 99.54 | 98.17 | 99.88 | Cobas (LOD: 0.009 TCID50/ml); Genesig (LOD: 0.33 copies/µl); XpertXpress (LOD: 0.02 PFU/ml) | NP | unclear | unclear
|  | Fernandez, f128.1 | Spain | AN/MT | no | unclear | 46 | 87.5 | 67,6* | 97,3* | 100 | 84,6* | 100* | 46 | 87.5 | # | # | 100 | # | # | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic, high risk contact | unclear
|  | Karon, f148.2 | USA | NP | yes | frozen | 347 | 83.2 | 77,3* | 88,2* | 100 | 97.5 | 100 | 347 | 83.2 | 77.4 | 87.8 | 100 | 97.5 | 100 | TibMolBiol (LOD: 1-10 copies/reaction) | NP | symptomatic, screening independent of symptoms | unclear
|  | Leli, f151.1 | Italy | AN/MT | no | fresh | 792 | 68.7 | 60.9 | 75.5 | 95.2 | 93,2* | 96.7 | 792 | 68.7 | 60.9 | 75.5 | 95.2 | 93.1 | 96.7 | Alinity (LOD: 100 copies/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Cento, f196.1 | Italy | NP | no | fresh | 960 | 85,6* | 81,4* | 89,1* | 97,2* | 95,6* | 98,4* | 960 | 85 | 82 | 89 | 97 | 96 | 98 | Simplexa (LOD: 500 copies/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Burdino, k6.1 | Italy | AN/MT | no | unclear | 1232 | 90.1 | 86.2 | 93.1 | 99.4 | 98.6 | 99.8 | 1232 | 90.1 | 86.2 | 93.1 | 99.4 | 98.6 | 99.8 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); XpertXpress (LOD: 0.02 PFU/ml); Simplexa (LOD: 500 copies/ml); TaqPath (LOD: 1.98 * 10^3 NDU/ml) | NP | unclear | Alpha
|  | Denine, k7.1 | Italy | AN/MT | no | unclear | 191 | 94.1 | # | # | 91.9 | # | # | 191 | 94.1 | # | # | 91.9 | # | # | Simplexa (LOD: 500 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Drain, k8.1 | USA | AN/MT | no | frozen | 222 | 82.1 | 63.1 | 93.9 | 100 | 98,1* | 100* | 222 | 82.1 | 63.1 | 93.9 | 100 | 98.1 | 100 | TaqPath (LOD: 1.98 * 10^3 NDU/ml); Cobas (LOD: 0.009 TCID50/ml) | AN/MT | high risk contact | wildtype
|  | FINDdx, k20.1 | Germany | AN/MT | yes | fresh | 761 | 82.2 | 75* | 88* | 99.3 | 98.3 | 99.7 | 761 | 82.2 | 75.2 | 87.5 | 99.3 | 98.3 | 99.7 | Cobas (LOD: 0.009 TCID50/ml); Allplex (LOD: 4,167 copies/ml); TibMolBiol (LOD: 1-10 copies/reaction); RealStar (LOD: 10 NDU/ml); RealTime (LOD: 200 copies/ml); LightMix (LOD: N/A) | NP | symptomatic, high risk contact | wildtype
|  | FINDdx, k20.2 | Brazil | AN/MT | yes | fresh | 251 | 86.2 | 74,6* | 93,9* | 99 | 96.3 | 99,9* | 251 | 86.2 | 75.1 | 92.8 | 99 | 96.3 | 99.7 | lab-developed | NP | symptomatic, high risk contact | Gamma
|  | Catellan, s6.1 | Italy | AN/MT | yes | fresh | 282 | 76.3 | 70,3* | 81,7* | 94.4 | 84,6* | 98,8* | 282 | 76.3 | 70.8 | 81.8 | 94.4 | 88.3 | 100 | One Step Real Time (LOD: N/A) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Elli, s9.1 | Italy | AN/MT | yes | fresh | 234 | 34,2* | 19.6 | 51,4* | 92,3* | 87,7* | 95,7* | 234 | 34 | 19.6 | 51.3 | 92 | 87.6 | 95.6 | GeneFinder(LOD: 0,5 cq/ml); XpertXpress (LOD: 0.02 PFU/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Greub, s16.1 | Switzerland | NP | yes | stored | # | 99* | 94,6* | 100* | 99 | # | # | # | 98.6 | 92.6 | 100 | 99 | # | # | BD MAX system (LOD: 640 copies/ml); GeneXpert (LOD: 0.02 PFU/ml); Cobas (LOD: 0.009 TCID50/ml) | NP | unclear | wildtype
|  | Kirby, s21.1 | USA | NP | yes | stored | 489 | # | # | # | # | # | # | 489 | # | # | # | # | # | # | Abbott (LOD: 100 copies/ml); Alinity (LOD: 100 copies/ml) | NP | symptomatic, screening independent of symptoms | wildtype, Alpha, Delta
|  | Sitoe, s44.1 | Mozambique | AN/MT | yes | stored | 696 | 78 | 73.5 | 82 | 95.8 | 92.9 | 97.7 | 696 | 78 | 73.5 | 82 | 95.8 | 92.9 | 97.7 | Abbott (LOD: 100 copies/ml) | OP | symptomatic, high risk contact, screening independent of symptoms | Alpha, Delta
| Quidel, Sofia SARS Antigen FIA (LFA; requires reader) | mariPOC SARS-CoV-2, ArcDia International Ltd (lab-based, Fluorescence immunoassay; LOD: 140 TCID50/ml corresponding to 84,500 copies/ml)
|  | Koskinen, f197.1 | Finland | NP | yes | stored | 13 | 100 | 75.3 | 100 | # | # | # | 13 | 100 | 75.3 | 100 | # | # | # | Allplex (LOD: 4,167 copies/ml); Amplidiag (LOD: N/A) | NP | symptomatic | unclear
|  | Koskinen, f197.2 | Finland | NP | yes | frozen | 45 | 84.4 | 70.5 | 93.5 | # | # | # | 45 | 84.4 | 70.5 | 93.5 | # | # | # | Allplex (LOD: 4,167 copies/ml); Amplidiag (LOD: N/A) | NP | symptomatic, high risk contact, screening independent of symptoms | unclear
|  | Gunell, s17.1 | Finland | NP | yes | stored | 939 | 78.3 | # | # | 100 | # | # | 939 | 78.3 | # | # | 100 | # | # | in-house PCR | NP | symptomatic, high risk contact, screening independent of symptoms | Alpha
|  | Gunell, s17.2 | Finland | NP | yes | stored | 23 | 73.9 | # | # | # | # | # | 23 | 73.9 | # | # | # | # | # | in-house PCR | NP | symptomatic, high risk contact, screening independent of symptoms | Alpha
| Quidel, Sofia SARS Antigen FIA (LFA; requires reader) | PCL COVID-19 Ag Rapid FIA test, PCL (POC; Fluorescnece immunoassay; LOD: 0,1ng/ml (r-N protein)/ 1000 PFU (Active Virus))
|  | Al-Alawi, k2.1 | Oman | NP | yes | unclear | 87 | 69.8 | 55.66 | 81.66 | 94.1 | 80.32 | 99.28 | 87 | 69.8 | 55.66 | 81.66 | 94.1 | 80.32 | 99.28 | LifeRiver (LOD: 1×10^3 copies/mL); Cobas (LOD: 0.009 TCID50/ml); Sansuro Biotech (LOD: 200 copies/ml) | NP | symptomatic | wildtype
|  | Horban, s19.1 | Poland | NP | yes | unclear | 266 | 40 | # | # | 97,8* | 94,9* | 99,3* | 266 | 40 | # | # | 98 | # | # | Genesig (LOD: 0.33 copies/µl) | NP | unclear | unclear
| Quidel, Sofia SARS Antigen FIA (LFA; requires reader) | RapidTesta, Sekisui Medical Co., Ltd. (POC, Immunochromatography assay; LOD: NP/nasal 1.1x10^2 TCID50/ml)
|  | Suzuki, f194.1 | Japan | NP | no | fresh | 1127 | 78.4 | 67.3 | 87.1 | 97.6 | 96.5 | 98.5 | 1127 | 78.4 | 67.3 | 87.1 | 97.6 | 96.5 | 98.5 | in-house | NP | symptomatic, high risk contact, screening independent of symptoms | Alpha
| Quidel, Sofia SARS Antigen FIA (LFA; requires reader) | S-PLEX, MesoScale (lab-based, Electrochemiluminescence immunoassays; LOD: 62 fg/mL)
|  | Bachmann, f154.1 | USA | AN/MT | yes | frozen | 170 | 91,3* | 83,6* | 96,2* | 94,9* | 87,4* | 98,6* | 170 | 86 | 77 | 92 | 96 | 89 | 99 | BioRad CFX96 (LOD: N/A) | AN/MT | symptomatic | wildtype
|  | Wang, k1.1 | USA | NP | yes | fresh | 200 | 88 | 80 | 93.6 | 93 | 86.1 | 97.1 | 200 | 88 | 80 | 93.6 | 93 | 86.1 | 97.1 | Panther Fusion (LOD: 1x10^-2 TCID50/ml) | NP | unclear | wildtype
|  | Wang, k1.2 | USA | NP | yes | frozen | 250 | 72.9 | 60.9 | 82.8 | 96.1 | 92.2 | 98.4 | 250 | 72.9 | 60.9 | 82.8 | 96.1 | 92.2 | 98.4 | Panther Fusion (LOD: 1x10^-2 TCID50/ml) | NP | unclear | wildtype
|  | Pollock, k19.1 | USA | NP | yes | frozen | 141 | 79 | 70 | 87 | 98 | 87 | 100 | 141 | 79 | 70 | 87 | 98 | 87 | 100 | Panther Fusion (LOD: 1x10^-2 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | unclear
|  | Pollock, k19.2 | USA | NP | yes | frozen | 85 | 91 | 77 | 98 | 100 | 93 | 100 | 85 | 91 | 77 | 98 | 100 | 93 | 100 | Panther Fusion (LOD: 1x10^-2 TCID50/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | unclear
| Quidel, Sofia SARS Antigen FIA (LFA; requires reader) | Sofia SARS Antigen FIA, Quidel (POC, Fluorescence immunoassay; LOD as per manufacturer's instructions: 1.13 * 10^2 TCID50/ml)
f114.1 |  | Beck, a04.1 | USA | NP | yes | fresh | 346 | 77 | 64,5* | 86,8* | 99.6 | 98,1* | 100* | 346 | 77 | # | # | 99.6 | # | # | Hologic (LOD: 600 NDU/ml) | AN/MT | symptomatic | unclear
f114.2 |  | Porte, a32.1 | Chile | NP | yes | frozen | 64 | 93.8 | 79,2* | 99,2* | 96.9 | 83,8* | 99,9* | 64 | 93.8 | 79.9 | 98.3 | 96.9 | 84.3 | 99.4 | Genesig (LOD: 0.33 copies/µl) | NP | symptomatic | unclear
|  | Herrera, a46.1 | USA | unclear | no | unclear | 1172 | 76.8 | 72.6 | 80.5 | 99.2 | 98.2 | 99.7 | 1172 | 76.7 | 72.6 | 80.5 | 99.2 | 98.2 | 99.7 | unclear | unclear | symptomatic | wildtype
|  | Marti, f46.1 | USA | AN/MT | yes | unclear | 427 | 72 | 56,3* | 84,7* | 99,7* | 98,6* | 100* | 427 | 72 | # | # | 99 | # | # | XpertXpress (LOD: 0.02 PFU/ml) | AN/MT | symptomatic | wildtype
|  | Jaaskelainen, f50.1 | Finland | NP | yes | frozen | 188 | 80.4 | 73,1* | 86,5* | 100 | 91,2* | 100* | 188 | 80.4 | 73.3 | 86 | 100 | # | # | in-house | NP | symptomatic, screening independent of symptoms | wildtype
|  | Brihn, f98.1 | USA | AN/MT | no | fresh | 2039 | 65.8 | 57.6 | 73.3 | 99.4 | 98.9 | 99.7 | 2039 | 65.8 | 57.6 | 73.3 | 99.4 | 98.9 | 99.7 | Fulgent (LOD: 5 copies/µl) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Bornemann, f112.1 | Germany | NP | yes | fresh | 1391 | 57.1 | 46,3* | 67,5* | 99.3 | 98,7* | 99,7* | 1391 | 57.1 | # | # | 99.3 | # | # | VIASURE (LOD: >=10 RNA copies/reaction); Rida (LOD: 4,3 copies/ml); GeneXpert (LOD: 0.02 PFU/ml) | NP | unclear | wildtype
|  | Harris, f114.1 | USA | AN/MT | yes | fresh | 885 | 82.3 | 77,5* | 86,4* | 98,8* | 97,5* | 99,5* | 885 | 82.3 | # | # | 98.2 | # | # | 2019-nCov CDC EUA Kit (LOD: 10^0.5 copies/µl) | NP | symptomatic, high risk contact | wildtype
|  | Harris, f114.2 | USA | AN/MT | yes | fresh | 1538 | 0 | # | # | 100 | # | # | 1538 | 0 | # | # | 100 | # | # | 2019-nCov CDC EUA Kit (LOD: 10^0.5 copies/µl) | NP | symptomatic, high risk contact | wildtype
|  | Smith, f122.1 | USA | NP | yes | fresh | 2887 | 76.6 | 70,7* | 81,9* | 99.7 | 99,5* | 99,9* | 2887 | 76.6 | 71 | 82 | 99.7 | 99 | 100 | RealTime (LOD: 200 copies/ml); Cobas (LOD: 0.009 TCID50/ml); XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Frediani, f139.1 | USA | NP | no | frozen | 48 | 56,25* | # | # | # | # | # | 48 | 56 | # | # | # | # | # | Simplexa (LOD: 500 copies/ml) | NP | unclear | wildtype
|  | Bachmann, f154.2 | USA | AN/MT | yes | fresh | 170 | 73,6* | 64,4* | 81,6* | 98,3* | 91,1* | 100 | 170 | 74 | 64 | 82 | 98 | 91 | 100 | 2019-nCov CDC EUA Kit (LOD: 10^0.5 copies/µl) | NP | symptomatic | wildtype
|  | Al-Alawi, k2.2 | Oman | NP | yes | unclear | 78 | 64.3 | 50.36 | 76.64 | 100 | 84.56 | 100 | 78 | 64.3 | 50.36 | 76.64 | 100 | 84.56 | 100 | LifeRiver (LOD: 1×10^3 copies/mL); Cobas (LOD: 0.009 TCID50/ml); Sansuro Biotech (LOD: 200 copies/ml) | NP | symptomatic | wildtype
|  | Mitchell, k12.1 | USA | AN/MT | yes | fresh | 144 | 33.3 | # | # | 100 | # | # | 144 | 33.3 | # | # | 100 | # | # | XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Mitchell, k12.2 | USA | AN/MT | yes | fresh | 148 | 87.8 | # | # | 100 | # | # | 148 | 87.8 | # | # | 100 | # | # | XpertXpress (LOD: 0.02 PFU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Agard, s1.1 | USA | AN/MT | yes | stored | 109 | 31 | # | # | # | # | # | 109 | 31 | # | # | 100 | # | # | XpertXpress (LOD: 0.02 PFU/ml); Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic, screening independent of symptoms | wildtype, Alpha
|  | Bornemann, s4.1 | Germany | NP | yes | fresh | 7863 | 62,9* | 57,8* | 67,8* | 99,4* | 99,2* | 99,5* | 7863 | 62.9 | 58 | 68 | 99.4 | 99 | 100 | VIASURE (LOD: >=10 RNA copies/reaction); Rida (LOD: 4,3 copies/ml); GeneXpert (LOD: 0.02 PFU/ml); Allplex (LOD: 4,167 copies/ml); LightMix (LOD: N/A) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Černila, s7.1 | Slowenia | NP | no | fresh | 132 | 90 | 55.5 | 99.75 | 97.52 | 92.93 | 99.49 | 132 | 90 | 55.5 | 99.75 | 97.52 | 92.93 | 99.49 | unclear | NP | symptomatic | unclear
|  | Černila, s7.2 | Slowenia | NP | no | fresh | 672 | 94,8* | 92,6* | 96,5* | 95* | 89,3* | 98,1* | 672 | 81.43 | 73.98 | 87.5 | 98.87 | 97.55 | 99.58 | unclear | NP | symptomatic, screening independent of symptoms | unclear
|  | Freeman, s11.1 | USA | AN/MT | yes | fresh | 249 | 81,8* | 67,3* | 91,8* | 99 | 96,5* | 99,9* | 249 | 82 | 68 | 91 | 99 | 97 | 100 | XpertXpress (LOD: 0.02 PFU/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Freeman, s11.2 | USA | AN/MT | yes | fresh | 138 | 64 | 39 | 84 | 99 | 96 | 100 | 138 | 64 | 39 | 84 | 99 | 96 | 100 | XpertXpress (LOD: 0.02 PFU/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Gourgeon, s15.1 | France | NP | yes | frozen | 282 | 59.1 | 50.2 | 67.6 | 99 | 94.6 | 100 | 282 | 59.1 | 50.2 | 67.6 | 99 | 94.6 | 100 | in-house; Alinity (LOD: 100 copies/ml); RealStar (LOD: 10 NDU/ml); Hologic (LOD: 600 NDU/ml); TaqPath (LOD: 10 GCE/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha, Delta
| SD Biosensor/Roche, Standard Q (nasal sampling, LFA) | Standard F, SD Biosensor/Roche (POC, Fluorescence immunoassay; LOD not publicly available from manufacturer)
f35.4 |  | Drevinek, a10.1 | Czech Republic | NP | yes | fresh | 591 | 62.3 | 55,6* | 68,7* | 99.5 | 98 | 99.9 | 591 | 62.3 | 55.8 | 68.4 | 99.5 | 98 | 99.9 | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic, high risk contact | wildtype
f15.2 |  | Liotti, a22.1 | Italy | NP | yes | stored | 359 | 47.1 | 37.1 | 57.1 | 98.4 | 96 | 99.6 | 359 | 47.1 | 37.1 | 57.1 | 98.4 | 96 | 99.6 | Allplex (LOD: 4,167 copies/ml); RealStar (LOD: 10 NDU/ml); Cobas (LOD: 0.009 TCID50/ml); Simplexa (LOD: 500 copies/ml) | NP | unclear | unclear
f45.2 |  | Porte, a32.2 | Chile | NP | yes | frozen | 64 | 90.6 | 75* | 98* | 96.9 | 83,8* | 99,9* | 64 | 90.6 | 75.8 | 96.8 | 96.9 | 84.3 | 99.4 | GeneXpert (LOD: 0.02 PFU/ml); Cobas (LOD: 0.009 TCID50/ml) | NP | symptomatic | unclear
|  | FINDdx, a63.2 | Brazil | NP | no | unclear | 453 | 77.5 | 69* | 84,6* | 97.9 | 95.7 | 99,2* | 453 | 77.5 | 69.2 | 84.1 | 97.9 | 95.7 | 99 | Simplexa (LOD: 500 copies/ml) | NP | symptomatic, high risk contact | wildtype
|  | FINDdx, a63.1 | Germany | NP | no | fresh | 676 | 69.2 | 52,4* | 83* | 96.9 | 95.2 | 98 | 676 | 69.2 | 53.6 | 81.4 | 96.9 | 95.2 | 98 | TaqPath (LOD: 1.98 * 10^3 NDU/ml) | NP | symptomatic, high risk contact | wildtype
a21.1 |  | Osterman, f20.1 | Germany | NP | yes | fresh, stores, frozen | 549 | 60.85 | 53,5* | 67,8* | 97.78 | 95.68 | 99* | 549 | 60.85 | 53.7 | 67.5 | 97.78 | 95.68 | 98.9 | XpertXpress (LOD: 0.02 PFU/ml); TaqPath (LOD: 1.98 * 10^3 NDU/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Bruzzone, f86.1 | Italy | unclear | yes | stored | 60 | 86.7 | 75,4* | 94,1* | # | # | # | 60 | 86.7 | 75.8 | 93.1 | # | # | # | Allplex (LOD: 4,167 copies/ml) | unclear | unclear | wildtype
|  | Menchinelli, f145.1 | Italy | NP | yes | fresh | 2898 | 68,5* | 61,6* | 74,9* | 97* | 96,3* | 97,6* | 2898 | # | # | # | # | # | # | Allplex (LOD: 4,167 copies/ml); Simplexa (LOD: 500 copies/ml); Cobas (LOD: 0.009 TCID50/ml); Hologic (LOD: 600 NDU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Baccani, f165.2 | Italy | NP | no | unclear | 93 | 35.7 | 18,6* | 55,9* | 100 | 94,5* | 100 | 93 | 35.7 | 20.1 | 54.2 | 100 | 96.2 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | unclear | unclear
|  | Krahn, f167.1 | Germany | OP | yes | unclear | 3110 | 59.4 | 48,9* | 69,3* | 99 | 98,5* | 99,3* | 3110 | 59.4 | # | # | 99 | # | # | Cobas (LOD: 0.009 TCID50/ml); Hologic (LOD: 600 NDU/ml); GeneXpert (LOD: 0.02 PFU/ml); RealStar (LOD: 10 NDU/ml) | OP | symptomatic, screening independent of symptoms | wildtype
|  | Orsi, f191.2 | Italy | NP | no | fresh | 110 | 86.7 | 75.4 | 94.1 | 100 | 92.9 | 100 | 110 | 86.7 | 75.4 | 94.1 | 100 | 92.9 | 100 | Allplex (LOD: 4,167 copies/ml) | NP | symptomatic | wildtype
|  | Weitzel, f193.1 | Chile | NP | yes | unclear | 773 | # | # | # | 91.3 | 89.1 | 93.2 | 773 | # | # | # | 91.3 | 89.1 | 93.2 | XpertXpress (LOD: 0.02 PFU/ml); Simplexa (LOD: 500 copies/ml) | NP | screening independent of symptoms | wildtype
|  | Lanari, k9.1 | Italy | NP | no | fresh | 381 | 53.8 | 33,4* | 73,4* | 99.7 | 98.4 | 100 | 381 | 53.8 | 35.4 | 71.4 | 99.7 | 98.4 | 100 | XpertXpress (LOD: 0.02 PFU/ml); Simplexa (LOD: 500 copies/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha
|  | Leiner,k10.1 | Germany | AN/MT | no | unclear | 4076 | 62.3 | # | # | 97.6 | # | # | 4076 | 62.3 | # | # | 97.6 | # | # | Allplex (LOD: 4,167 copies/ml) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | wildtype
|  | Gourgeon, s15.2 | France | NP | yes | frozen | 299 | 47,3* | 39,1* | 55,6* | 100 | 97.6 | 100 | 299 | 47 | 38.8 | 55.3 | 100 | 97.6 | 100 | in-house; Alinity (LOD: 100 copies/ml); RealStar (LOD: 10 NDU/ml); Hologic (LOD: 600 NDU/ml); TaqPath (LOD: 10 GCE/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha, Delta
|  | Kliegr, s22.1 | Czech Republic | unclear | no | unclear | 3907 | 79.5 | 75.7 | 98.7 | 98.3 | 99.1 | # | 3907 | 79.5 | 75.7 | 98.7 | 98.3 | 99.1 | # | unclear | unclear | symptomatic, high risk contact, screening independent of symptoms | Delta
|  | Yu, s51.1 | Korea | NP | no | fresh | 554 | 67.1 | 60,5* | 73,3* | # | # | # | 554 | 67.1 | 60.7 | 73 | # | # | # | PowerCek (LOD: 5,400 NDU/mL) | NP | symptomatic, high risk contact, screening independent of symptoms | Delta
|  | García-Fernández, s13.1 | Spain | NP | yes | fresh | 663 | 84 | 76,4* | 89,9* | 99.6 | 98,7* | 100* | 663 | 84 | 76.1 | 89.7 | 99.6 | 98.5 | 99.9 | VIASURE (LOD: >=10 RNA copies/reaction) | NP | symptomatic | wildtype
VITROS Immunodiagnostic Products SARS-CoV-2 Antigen test, Ortho Clinical Diagnostics (lab-based, Chemiluminescent Immunoassay; LOD: 5.0x102 -3.0x10^3 TCID50/ml)
|  | Favresse, f31.1 | Belgium | NP | yes | fresh | 188 | 83.3 | 74.4 | 90.2 | 100 | 96.1 | 100 | 188 | 83.3 | 74.4 | 90.2 | 100 | 96.1 | 100 | Cepheid Xpert Xpress | NP | symptomatic, screening independent of symptoms | wildtype
|  | Matsuzaki, f124.1 | Japan | NP | yes | fresh | 128 | 75.5 | 61,1* | 86,7* | 100 | 95,4* | 100* | 128 | 75.5 | # | # | 100 | # | # | RealStar (LOD: 10 NDU/ml) | NP | symptomatic, screening independent of symptoms | wildtype
|  | Levett, f173.1 | Canada | unclear | no | stored | 528 | 84,3* | 73,6* | 91,9* | 100 | 99,2* | 100* | 528 | 84.2 | # | # | 100 | # | # | unclear | unclear | unclear | unclear
|  | Levett, f173.2 | Canada | unclear | no | stored | 24 | 83,3* | 51,6* | 97,9* | 100 | # | # | 24 | 83 | # | # | 100 | # | # | unclear | unclear | unclear | unclear
|  | Paul, k14.1 | India | NP | yes | fresh | 148 | 73.47 | 63.59 | 81.88 | 100 | 92.89 | 100 | 148 | 73.47 | 63.59 | 81.88 | 100 | 92.89 | 100 | RealStar (LOD: 10 NDU/ml) | NP | symptomatic, high risk contact | unclear
|  | Fourti, s10.1 | France | NP | yes | fresh | 1655 | # | # | # | 100 | 99.8 | 100 | 1655 | # | # | # | 100 | 99.8 | 100 | Alinity (LOD: 100 copies/ml); Hologic (LOD: 600 NDU/ml) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype
WANTAI SARS-CoV-2 Antigen Rapid Test, WANTAI BioPharm (POC, fluorescence immunoassay; LOD: 5pg/mL, 60TCIDso/mL
|  | Gourgeon, s15.3 | France | NP | yes | frozen | 267 | 79.7 | 71.3 | 86.5 | 98 | 93 | 99.8 | 267 | 79.7 | 71.3 | 86.5 | 98 | 93 | 99.8 | in-house; Alinity (LOD: 100 copies/ml); RealStar (LOD: 10 NDU/ml); Hologic (LOD: 600 NDU/ml); TaqPath (LOD: 10 GCE/reaction) | NP | symptomatic, high risk contact, screening independent of symptoms | wildtype, Alpha, Delta
|  | Routsias, s38.1 | Greece | AN/MT | yes | fresh | 450 | 95.71 | # | # | 100 | # | # | 450 | 95.71 | # | # | 100 | # | # | LiliF (LOD: N/A) | AN/MT | symptomatic, high risk contact, screening independent of symptoms | unclear
0 |  | * these values were either not inlcuded in the respective publication but could be calculated from other values presented, or were recalculated to align statistical methods
Abbreviations: POC = Point-of-care; LOD = Limit of detection; IFU = Intructions for use; NP = nasopharyngeal; AN /MT = anterior nasal / mid-turbinate; OP = oropharnygeal; CI = confidence interval
Underlined author: head-to-head studies
Underlined sample size: not peer-reviewed
Underlined stuy location: pooled samples
ADVIA Centaur SARS-CoV 2 Ag, Siemens Healthineers (https://www.fda.gov/media/156971/download)
AFIAS COVID-19 Ag, Boditech Med Inc.(http://web.labindustrias.com/wp-content/uploads/2022/02/Boditech-COVID-19-Ag-Control-Rev.04.pdf)
Atellica Solutions, Siemens Healthineers (https://dmec.moh.gov.vn/documents/10182/29929418/upload_00064323_1651242713783.pdf)
Becton, Dickinson and Company, BD Veritor (https://www.fda.gov/media/139755/download)
CoviNag ELISA, XEMA medica (https://xema-medica.com/medical-diagnostics/catalog)
Elecsys® Anti-SARS-CoV-2-Antigen-Assays, Roche Diagnostics GmbH (https://www.fda.gov/media/137605/download)
Euroimmun SARS-CoV-2 Ag ELISA, EUROIMMUN Medizinische Labordiagnostika AG (https://www.euroimmun.com/products/infection-diagnostics/id/sars-cov-2-infection-covid-19/)
Exdia COVID-19 Ag, Precision Biosensor (https://zaidamed-shop.de/downloads/GA_e_COVID-19%20Ag_Test.pdf)
FIAflex™ Fluorescent Immunoassay Analyzer, ACON Biotech (https://www.google.com/search?client=safari&rls=en&q=FIAflex%E2%84%A2+Fluorescent+Immunoassay+Analyzer&ie=UTF-8&oe=UTF-8)
FREND COVID-19 Ag, NanoEntek (http://nanoentek.com/theme/nanont2_en/shop/02/product01_view.php?it_id=1608092888#detail_down)
Gazelle, Hemex Health (https://hemexhealth.com)
HISCL SARS-COV-2 AG ASSAY KIT, Sysmex Europe (https://www.sysmex-europe.com/products/products-detail/hiscl-sars-cov-2-ag-assay-kit.html)
iChroma COVID-19 Ag Test, Boditech Medical (http://web.labindustrias.com/wp-content/uploads/2022/02/ichroma-COVID-19-Ag-Rev.09.pdf)
LIAISON® SARS-COV-2 AG, DiaSorin (https://www.fda.gov/media/147310/download)
Lumipulse® G SARS-CoV-2 Ag, Fujirebio (https://www.fujirebio.com/en/products-solutions/lumipulse-g-sars-cov2-ag)
LumiraDx COVID-19 SARS-CoV-2 Antigen Test, LumiraDx UK Ltd (https://www.fda.gov/media/141302/download)
mariPOC SARS-CoV-2, ArcDia International Ltd (https://www.arcdia.com/maripoc/tests/covid-19-tests/)
PCL COVID-19 Ag Rapid FIA test, PCL (http://www.pclchip.com/en/sub02-3.php?page=01&sub_page=02)
RapidTesta, Sekisui Medical Co., Ltd. (https://www.sekisuimedical.jp/news/release/210727.html)
S-PLEX, MesoScale (https://www.mesoscale.com/en/products/s-plex-sars-cov-2-n-kit-sector-1-pl-k150adhs/)
Sofia SARS Antigen FIA, Quidel (https://www.quidel.com/sites/default/files/product/documents/EF1438906EN00.pdf)
STANDARD F, SD Biosensor (https://www.sdbiosensor.com/product/product_view?product_no=177)
VITROS Immunodiagnostic Products SARS-CoV-2 Antigen test, Ortho Clinical Diagnostics (https://www.fda.gov/media/145071/download)
WANTAI SARS-CoV-2 Antigen Rapid Test, WANTAI BioPharm (https://www.fda.gov/media/140030/download)

## Raw data
| studyID_sub1 | Study ID | Sub_1 | Sub_2 | Sub_3 | Author.x | Test assessed | Technology | Connectivity | Turnaround Time | Expected Location | Throughput[tests/h] | Head to Head? | Peer reviewed? | Type of study | Retro/Prosp? | Study location.x | Type of location | Study Population | Screening criteria | Age (Median) | Age (mean) | Age (lower Range) | Age (upper Range) | Age (IQR 1) | Age (IQR 3) | Age (STDev) | Sex (Number of Female) | Reference Standard | Type of RT-PCR | Type of RT-PCR_notes | Type of sample (AG) | sample_adjusted | Type of sample (PCR) | sample_adjusted (PCR) | Testing procedure | Population size | of asymptomatic | of symptomatic | of hospitalized | of children (<18 years) | of adults (> or = 18 years) | CT value (median) | CT value (mean) | CT value (lower range) | CT value (upper range) | CT value (IQR 1) | CT value (IQR 3) | CT Value (STDev) | Viral load (median) | Viral load (mean)[log10 copies/uL] | Viral load (lower Range)[log10RNA copies/ml] | Viral load (upper Range) [log10RNA copies/ml | Viral load (IQR 1) | Viral load (IQR 3) | Viral load (STDev) | Duration of symptoms (median) | Duration of symptoms (mean) | Duration of symptoms (lower range) | Duration of symptoms (upper range) | Duration of symptoms (IQR 1) | Duration of symptoms (IQR 3) | Duration of symptoms (STDev) | Antigen levels (quantity vs COI) | Antigen levels (Median)[pg/ml] | Antigen levels mean [log10pg/mL] | Antigen levels (lower Range) [pg/ml] | Antigen levels (upper Range)[pg/ml] | Anitgen levels (IQR 1) [pg/ml] | Antigen levels (IQR 3) [pg/ml] | Antigen levels (STDev) | True positives | False positives | True negatives | False negatives | Sensitivity | Lower 95 CI sensitivity | Upper 95 CI sensitivity | Specificity | Lower 95 CI specificity | Upper 95 CI specificity | LOD by Ct value | LOD by viral load | viral load UNIT | Setting | Start Date | End Date | Variant
1 | a04_1_0_0 | a04 | 1 | 0 | 0 | Beck, E. T., et al. | Sofia | FIA | yes | 15 | POC | # | no | yes | cohort | prospective | USA | clinical | adult, children | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Hologic | TMA, Hologic Aptima SARS-CoV-2 Assay (Hologic, Marlborough, MA) | NP | NP | nasal | AN/MT | IFU | 346 | 0 | 346 | # | 122 | 224 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 47 | 1 | 284 | 14 | 77 | 64.5 | 86.8 | 99.6 | 98.1 | 100 | # | # | # | HIC | unclear | unclear | unclear
2 | a10_1_0_0 | a10 | 1 | 0 | 0 | Drevinek, P., et al. | STANDARD F | FIA | yes | 30 | POC | # | no | no | cohort | prospective | Czech Republic | testing site | unclear | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex® 2019 n-CoV Assay (Seegene) | NP | NP | NP/OP | NP | IFU | 591 | 301 | 290 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 139 | 2 | 366 | 84 | 62.3 | 55.6 | 68.7 | 99.5 | 98 | 99.9 | # | # | # | HIC | 2020-10-01 | 2020-10-31 | wildtype
3 | a22_1_0_0 | a22 | 1 | 0 | 0 | Liotti, F. M., et al. | STANDARD F | FIA | yes | 30 | POC | # | no | yes | unclear | retrospective | Italy | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | RealStar; Allplex; Simplexa; Cobas | RealStar® SARS-CoV-2 RT-PCR Kit (altona Diagnostics), Allplex® 2019 n-CoV Assay (Seegene), Simplexa™ COVID-19 Direct kit (DiaSorin), Cobas® SARS-CoV-2 Test (Roche) | NP | NP | NP | NP | IFU | 359 | # | # | # | # | # | 31.3 | # | 15.3 | 39.7 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 49 | 4 | 251 | 55 | 47.1 | 37.1 | 57.1 | 98.4 | 96 | 99.6 | # | 5*10^2 or 2*10^6 | TCID50/ml or cp/ml | HIC | unclear | unclear | unclear
4 | a28_1_0_0 | a28 | 1 | 0 | 0 | Pekosz, A., et al. | BD Veritor | DIA | yes | # | POC | # | no | yes | case-control | retrospective | USA | clinical | adult | symptomatic | # | # | # | # | # | # | # | # | virus culture | Lyra | Lyra® SARS-CoV-2 Assay (Quidel) | NP | NP | Nasal / NP or OP | NP | IFU | 251 | 0 | 251 | # | 0 | 251 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0 | 7 | # | # | # | # | # | # | # | # | # | # | # | 27 | 3 | 220 | 1 | 96.4 | 81.7 | 99.9 | 98.7 | 96.1 | 99.7 | # | # | # | HIC | 2020-06-05 | 2020-06-11 | wildtype
5 | a31_1_0_0 | a31 | 1 | 0 | 0 | Porte, L., et al. | Bioeasy | FIA | yes | # | POC | # | no | yes | case-control | retrospective | Chile | clinical | adult, children | symptomatic | # | 38 | 1 | 91 | 29.5 | 44 | # | 59 | RT-PCR | RealTime; Cobas | RealTime SARS-CoV-2 assay M2000 (Abbott), Cobas® SARS-CoV-2 Test (Roche) | NP/OP | NP | NP/OP | NP | IFU | 127 | 0 | 127 | # | 16 | 111 | 17.7 | 20 | # | # | 14.1 | 25.1 | # | # | # | # | # | # | # | # | 2 | # | 0 | 12 | 1 | 4 | # | # | # | # | # | # | # | # | # | 77 | 0 | 45 | 5 | 93.9 | 86.3 | 98 | 100 | 92.1 | 100 | # | # | # | HIC | first weeks | first weeks | wildtype
6 | a32_1_0_0 | a32 | 1 | 0 | 0 | Porte, L., et al. | Sofia | FIA | yes | 15 | POC | # | yes | no | case-control | retrospective | Chile | clinical | adult | symptomatic | 39 | # | # | # | 36.7 | 57 | # | 31 | RT-PCR | Genesig | Genesig | NP/OP | NP | NP/OP | NP | IFU | 64 | 12 | 52 | # | 0 | 64 | 17.95 | # | # | # | 16.4 | 22.4 | # | # | # | # | # | # | # | # | # | # | 0 | 5 | # | # | # | # | # | # | # | # | # | # | # | 30 | 1 | 31 | 2 | 93.8 | 79.2 | 99.2 | 96.9 | 83.8 | 99.9 | # | # | # | HIC | unclear | unclear | unclear
7 | a32_2_0_0 | a32 | 2 | 0 | 0 | Porte, L., et al. | STANDARD F | FIA | yes | 30 | POC | # | yes | no | case-control | retrospective | Chile | clinical | adult | symptomatic | 39 | # | # | # | 36.7 | 57 | # | 31 | RT-PCR | Genesig | Genesig | NP/OP | NP | NP/OP | NP | IFU | 64 | 12 | 52 | # | 0 | 64 | 17.95 | # | # | # | 16.4 | 22.4 | # | # | # | # | # | # | # | # | # | # | 0 | 5 | # | # | # | # | # | # | # | # | # | # | # | 29 | 1 | 31 | 3 | 90.6 | 75 | 98 | 96.9 | 83.8 | 99.9 | # | # | # | HIC | unclear | unclear | unclear
8 | a39_1_0_0 | a39 | 1 | 0 | 0 | Van der Moeren, N., et al. | BD Veritor | DIA | yes | 15 | POC | # | no | no | cohort | prospective | Netherlands | testing site | adult | symptomatic | # | # | # | # | # | # | # | NA | RT-PCR | Cobas, RealTime | Cobas, RealTime | NP/OP | NP | NP/OP | NP | IFU | 351 | 352 | 0 | 0 | 0 | 352 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 16 | 2 | 332 | 1 | 94.1 | 71.3 | 99.9 | 99.4 | 97.9 | 99.9 | # | # | # | HIC | 2020-09-26 | 2020-10-07 | wildtype
9 | a41_1_0_0 | a41 | 1 | 0 | 0 | Weitzel, T., et al. | Bioeasy | FIA | yes | 10 to 20 | POC | # | yes | yes | case-control | retrospective | Chile | clinical | unclear | symptomatic | 40 | # | # | # | # | # | # | 44 | RT-PCR | Simplexa | Simplexa COVID-19 Direct EUA RT-PCR | NP/OP | NP | NOP | NP | IFU | 111 | 0 | 111 | # | # | # | # | 22.5 | 10.7 | 37.7 | # | # | # | # | # | # | # | # | # | # | 2 | # | # | # | 1 | 5 | # | # | # | # | # | # | # | # | # | 68 | 0 | 31 | 12 | 85 | 75.3 | 92 | 100 | 88.8 | 100 | # | # | # | HIC | 2020-03-16 | 2020-04-26 | wildtype
10 | a43_1_0_0 | a43 | 1 | 0 | 0 | Young, S., et al. | BD Veritor | DIA | yes | 15 | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic | 43 | 45 | 18 | 90 | # | # | 16.6 | 163 | RT-PCR | Cobas | Cobas m2000 | NP | NP | NP/OP | NP | IFU | 251 | 0 | 251 | # | 0 | 251 | # | 22.74 | # | # | # | # | 5.48 | # | # | # | # | # | # | # | 3 | 3.2 | # | # | # | # | # | # | # | # | # | # | # | # | # | 29 | 1 | 212 | 9 | 76.3 | 59.8 | 88.6 | 99.5 | 97.4 | 99.9 | # | # | # | HIC | 2020-06-05 | 2020-06-11 | wildtype
11 | a46_1_0_0 | a46 | 1 | 0 | 0 | Herrera, V., et al. | Sofia | FIA | yes | # | POC | # | no | no | cohort | prospective | USA | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | N/A | unclear | non-IFU | 1172 | 0 | 1172 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 352 | 6 | 707 | 107 | 76.8 | 72.6 | 80.5 | 99.2 | 98.2 | 99.7 | # | # | # | HIC | unclear | unclear | wildtype
12 | a58_1_0_0 | a58 | 1 | 0 | 0 | Parada-Ricart, E., et al. | Bioeasy | FIA | yes | # | POC | # | no | yes | cohort | unclear | Spain | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Lyra | Lyra SARS-CoV-2 | NP | NP | NP | NP | IFU | 172 | 44 | 128 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 19 | 21 | 125 | 7 | 73.1 | 52.2 | 88.4 | 85.6 | 78.9 | 90.9 | # | # | # | HIC | 2020-04-01 | 2020-04-30 | wildtype
13 | a63_1_0_0 | a63 | 1 | 0 | 0 | FINDdx | STANDARD F | FIA | yes | # | POC | # | no | no | cohort | prospective | Brazil | both | adult, children | symptomatic, high risk contact | # | 39 | 0 | 95 | # | # | # | 268 | RT-PCR | Simplexa | DiaSorin | NP | NP | NP | NP | non-IFU | 453 | 29 | 421 | # | # | # | 23.4 | # | # | # | 19.9 | 27.7 | # | # | # | # | # | # | # | # | 4 | # | # | # | 3 | 6 | # | # | # | # | # | # | # | # | # | 93 | 7 | 326 | 27 | 77.5 | 69 | 84.6 | 97.9 | 95.7 | 99.2 | # | Omicron:8,8x10^4 | genome copies/ml | LMIC | 2020-07-11 | 2020-09-09 | wildtype
14 | a63_2_0_0 | a63 | 2 | 0 | 0 | FINDdx | STANDARD F | FIA | yes | # | POC | # | no | no | cohort | prospective | Germany | testing site | adult | symptomatic, high risk contact | # | 37.8 | 18 | 85 | # | # | # | 360 | RT-PCR | TibMolBiol | TaqMan | NP/OP | NP | NOP | NP | non-IFU | 676 | 159 | 517 | # | 0 | 676 | 24.7 | # | # | # | 21.73 | 27.81 | # | # | # | # | # | # | # | # | 3 | # | # | # | 2 | 5 | # | # | # | # | # | # | # | # | # | 27 | 20 | 617 | 12 | 69.2 | 52.4 | 83 | 96.9 | 95.2 | 98 | # | Omicron: 5,0x10^2 | pfu/ml | HIC | 2020-06-15 | 2020-09-23 | wildtype
15 | f102_1_0_0 | f102 | 1 | 0 | 0 | Bianco, G., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Italy | clinical | adult, children | symptomatic, screening independent of symptoms | # | 47.9 | 0 | 94 | # | # | # | 505 | RT-PCR | Xpert Xpress | Xpert Xpress | nasal | AN/MT | NP | NP | IFU | 907 | 676 | 231 | # | 165 | 742 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4 | # | # | # | 2 | 7 | # | # | # | # | # | # | # | # | # | 269 | 48 | 561 | 29 | 90.3 | 86.3 | 93.4 | 92.1 | 89.7 | 94.1 | # | # | # | HIC | 2020-10-01 | 2020-12-01 | wildtype
16 | f103_1_0_0 | f103 | 1 | 0 | 0 | Koeleman, J. G. M., et al. | BD Veritor | DIA | yes | 15 | POC | # | yes | yes | case-control | prospective | Netherlands | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | ELITe InGenius; XpertXpress; in-house | ELITe InGenius, GeneXpert Xpress, in-house | NP | NP | NP | NP | IFU | 40 | 0 | 40 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 31 | NA | NA | 9 | 77.5 | 61.5 | 89.2 | NA | NA | NA | # | # | # | HIC | 2020-10-24 | 2020-11-15 | wildtype
17 | f108_1_0_0 | f108 | 1 | 0 | 0 | Caramello, V., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Italy | clinical | unclear | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | NP | NP | NP | NP | non-IFU | 149 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 84 | 6 | 46 | 13 | 86.6 | 78.2 | 92.7 | 88.5 | 76.6 | 95.6 | # | # | # | HIC | 26.10.2020 | 2020-11-10 | wildtype
18 | f109_1_0_0 | f109 | 1 | 0 | 0 | Christensen, K., et al. | BD Veritor | DIA | yes | # | POC | # | no | no | cohort | retrospective | USA | clinical | adult | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Lyra | Lyra SARS-CoV-2 | nasal | AN/MT | nasal | NP | non-IFU | 278 | 0 | 278 | # | 0 | 278 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7 | # | # | # | # | # | # | # | # | # | # | # | 52 | 1 | 217 | 8 | 86.7 | 75.4 | 94.1 | 99.5 | 97.4 | 99.9 | # | # | # | HIC | 2020-10-16 | 2020-10-30 | wildtype
19 | f112_1_0_0 | f112 | 1 | 0 | 0 | Bornemann, L., et al. | Sofia | FIA | yes | # | POC | # | no | yes | cohort | prospective | Germany | clinical | adult, children | unclear | 68 | # | 4 | 102 | # | # | # | 767 | RT-PCR | Viasure; Rida; XpertXpress | Viasure; Rida; XpertXpress | NP | NP | NP | NP | IFU | 1391 | # | # | # | # | # | 26.2 | # | 13 | 41 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 52 | 9 | 1291 | 39 | 57.1 | 46.3 | 67.5 | 99.3 | 98.7 | 99.7 | # | # | # | HIC | 2020-11-01 | 2020-11-30 | wildtype
20 | f113_1_0_0 | f113 | 1 | 0 | 0 | Blairon, L., et al. | LumiraDx | FIA | yes | 15 | POC | # | yes | yes | case-control | prospective, retrospective | Belgium | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | NP | NP | NP | NP | non-IFU | 198 | # | # | # | # | # | 25.83 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 91 | 0 | 49 | 58 | 61.1 | 52.8 | 68.9 | 100 | 92.7 | 100 | # | # | # | HIC | 2021-02-02 | 2021-03-03 | Alpha
21 | f114_1_0_0 | f114 | 1 | 0 | 0 | Harris, D. T., et al. | Sofia | FIA | yes | 15-30 | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | 2019-nCov CDC EUA Kit | CDC PCR | nasal | AN/MT | NP | NP | IFU | 885 | 0 | 885 | # | 0 | 885 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 5 | # | # | # | # | # | # | # | # | # | # | # | 251 | 7 | 573 | 54 | 82.3 | 77.5 | 86.4 | 98.8 | 97.5 | 99.5 | # | # | # | HIC | 2020-06-01 | 2020-08-01 | wildtype
22 | f114_2_0_0 | f114 | 2 | 0 | 0 | Harris, D. T., et al. | Sofia | FIA | yes | 15-30 | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | 2019-nCov CDC EUA Kit | CDC PCR | nasal | AN/MT | NP | NP | IFU | 1538 | 1538 | 0 | # | 0 | 1538 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0 | 0 | 1525 | 13 | 0 | NA | NA | 100 | NA | NA | # | # | # | HIC | 2020-06-01 | 2020-08-01 | wildtype
23 | f119_1_0_0 | f119 | 1 | 0 | 0 | Dierks, S., et al. | LumiraDx | FIA | yes | # | POC | # | no | yes | cohort | prospective | Germany | testing site | adult | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, Genesig, Xpert Xpress | Cobas, Genesig, Xpert Xpress | NP | NP | NP | NP | IFU | 444 | # | # | # | 0 | 444 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7 | 0 | 433 | 4 | 63.6 | 30.8 | 89.1 | 100 | 99.2 | 100 | # | # | # | HIC | unclear | unclear | unclear
24 | f122_1_0_0 | f122 | 1 | 0 | 0 | Smith, R., et al. | Sofia | FIA | yes | 15 | POC | # | no | yes | cross-sectional | prospective | USA | clinical | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | RealTime; Cobas; XpertXpress | Abbott, Cobas, XpertXpress | NP | NP | NP | NP | IFU | 2887 | 1206 | 1675 | # | # | # | # | 22.7 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 180 | 7 | 2645 | 55 | 76.6 | 70.7 | 81.9 | 99.7 | 99.5 | 99.9 | # | # | # | HIC | 21.09.2020 | 2020-12-03 | wildtype
25 | f124_1_0_0 | f124 | 1 | 0 | 0 | Matsuzaki, N., et al. | VITROS | CLIA | yes | 48 | lab based | 130 | yes | yes | cohort | prospective | Japan | clinical | unclear | symptomatic, screening independent of symptoms | 58 | # | # | # | 47 | 74 | # | 23 | RT-PCR | Qiagen | Qiagen | NP | NP | NP | NP | IFU | 128 | # | 49 | # | # | # | 30 | # | # | # | 24 | 35 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 37 | 0 | 79 | 12 | 75.5 | 61.1 | 86.7 | 100 | 95.4 | 100 | # | # | # | HIC | 2020-12-22 | 2021-01-07 | wildtype
26 | f124_2_0_0 | f124 | 2 | 0 | 0 | Matsuzaki, N., et al. | Lumipulse G | CLEIA | yes | # | lab based | # | yes | yes | cohort | prospective | Japan | clinical | unclear | symptomatic, screening independent of symptoms | 58 | # | # | # | 47 | 74 | # | 23 | RT-PCR | QuantiTect | QuantiTect | NP | NP | NP | NP | IFU | 128 | # | 49 | # | # | # | 30 | # | # | # | 24 | 35 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 43 | 1 | 78 | 6 | 87.8 | 75.2 | 95.4 | 98.7 | 93.1 | 100 | # | # | # | HIC | 2020-12-22 | 2021-01-07 | wildtype
27 | f128_1_0_0 | f128 | 1 | 0 | 0 | Fernandez, M. D., et al | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | unclear | Spain | clinical | adult | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | nasal | AN/MT | NP | NP | non-IFU | 46 | 27 | 19 | # | 0 | 46 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 21 | 0 | 22 | 3 | 87.5 | 67.6 | 97.3 | 100 | 84.6 | 100 | # | # | # | HIC | unclear | unclear | unclear
28 | f139_1_0_0 | f139 | 1 | 0 | 0 | Frediani, J., et al. | BD Veritor | DIA | yes | # | POC | # | yes | yes | case-control | retrospective | USA | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Simplexa | DiaSorin | NP | NP | NP | NP | non-IFU | 48 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 24 | NA | NA | 24 | 50 | NA | NA | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-08-01 | wildtype
29 | f139_2_0_0 | f139 | 2 | 0 | 0 | Frediani, J., et al. | Sofia | FIA | yes | # | POC | # | yes | yes | case-control | retrospective | USA | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Simplexa | DiaSorin | NP | NP | NP | NP | non-IFU | 48 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 27 | NA | NA | 21 | 56.25 | NA | NA | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-08-01 | wildtype
30 | f144_1_0_0 | f144 | 1 | 0 | 0 | Van der Moeren, N., et al. | LIAISON | CLIA | yes | # | lab based | # | no | yes | case-control | retrospective | Netherlands | testing site | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Alinity | Alinity M | NOP | NP | NOP | NP | IFU | 248 | 0 | 248 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 54 | 0 | 174 | 20 | 73 | 61.4 | 82.6 | 100 | 97.9 | 100 | # | # | # | HIC | 2021-11-01 | 2021-11-01 | Delta
31 | f145_1_0_0 | f145 | 1 | 0 | 0 | Menchinelli, G., et al. | STANDARD F | FIA | yes | 15-30 | POC | # | no | yes | cohort | prospective | Italy | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Simplexa; Cobas; Hologic | Seegene Allplex, Diasorin Simplexa, Cobas, Hologic Aptima | NP | NP | NP | NP | IFU | 2898 | 1815 | 1083 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 137 | 81 | 2617 | 63 | 68.5 | 61.6 | 74.9 | 97 | 96.3 | 97.6 | # | # | # | HIC | 2020-12-01 | 2021-01-31 | wildtype
32 | f148_1_0_0 | f148 | 1 | 0 | 0 | Karon, B. S., et al. | BD Veritor | DIA | yes | 15 | POC | # | yes | yes | case-control | retrospective | USA | unclear | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TibMolBiol | TaqMan | NP | NP | NP | NP | IFU | 347 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 131 | 0 | 150 | 66 | 66.5 | 59.4 | 73 | 100 | 97.5 | 100 | # | # | # | HIC | unclear | unclear | unclear
33 | f148_2_0_0 | f148 | 2 | 0 | 0 | Karon, B. S., et al. | LumiraDx | FIA | yes | 12 | POC | # | yes | yes | case-control | retrospective | USA | unclear | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TibMolBiol | TaqMan | NP | NP | NP | NP | IFU | 347 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 164 | 0 | 150 | 33 | 83.2 | 77.3 | 88.2 | 100 | 97.5 | 100 | # | # | # | HIC | unclear | unclear | unclear
34 | f148_3_0_0 | f148 | 3 | 0 | 0 | Karon, B. S., et al. | FIAflex | FIA | yes | # | POC | # | yes | yes | case-control | retrospective | USA | unclear | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TibMolBiol | TaqMan | NP | NP | NP | NP | IFU | 347 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 174 | 0 | 150 | 23 | 88.3 | 83 | 92.5 | 100 | 97.5 | 100 | # | # | # | HIC | unclear | unclear | unclear
35 | f151_1_0_0 | f151 | 1 | 0 | 0 | Leli, C., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | retrospective | Italy | clinical | adult | symptomatic, screening independent of symptoms | 71 | # | # | # | 53 | 82.7 | # | 389 | RT-PCR | Alinity | Alinity | nasal | AN/MT | NP | NP | non-IFU | 792 | 377 | 415 | # | 0 | 792 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 114 | 30 | 596 | 52 | 68.7 | 60.9 | 75.5 | 95.2 | 93.2 | 96.7 | # | # | # | HIC | 2020-10-01 | 2021-01-01 | wildtype
36 | f154_1_0_0 | f154 | 1 | 0 | 0 | Bachmann, C., M., et al | MSD S-PLEX | ECL | yes | # | lab based | # | yes | yes | cohort | prospective | USA | testing site | adult | symptomatic | # | # | 18 | # | # | # | # | # | RT-PCR | BioRad CFX96 | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel, BioRad CFX96 | nasal | AN/MT | nasal | AN/MT | IFU | 170 | 0 | 170 | # | 0 | 170 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7 | # | # | # | # | # | # | # | # | # | # | # | 84 | 4 | 74 | 8 | 91.3 | 83.6 | 96.2 | 94.9 | 87.4 | 98.6 | # | # | # | HIC | 2020-12-01 | 2020-12-31 | wildtype
37 | f154_2_0_0 | f154 | 2 | 0 | 0 | Bachmann, C., M., et al | Sofia | FIA | yes | # | POC | # | yes | yes | cohort | prospective | USA | testing site | adult | symptomatic | # | # | 18 | # | # | # | # | # | RT-PCR | 2019-nCov CDC EUA Kit | CDC 2019-nCoV Real-Time RT-PCR Diagnostic Panel | nasal | AN/MT | NP | NP | IFU | 170 | 0 | 170 | # | 0 | 170 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 7 | # | # | # | # | # | # | # | # | # | # | # | 81 | 1 | 59 | 29 | 73.6 | 64.4 | 81.6 | 98.3 | 91.1 | 100 | # | # | # | HIC | 2020-12-01 | 2020-12-31 | wildtype
38 | f157_1_0_0 | f157 | 1 | 0 | 0 | Osterman, A., et al | Lumipulse G | CLEIA | yes | # | lab based | # | yes | yes | cohort | retrospective | Germany | clinical | adult | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Cobas; GeneXpert | Seegen Allplex, Cobas, Xpert Xpress | unclear | unclear | NP/OP/URT | NP | IFU | 410 | # | # | 410 | 0 | 410 | # | # | # | # | # | # | # | 6045 | # | 83 | 1548572803 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 56 | 9 | 294 | 51 | 52.34 | 42.5 | 62.1 | 97.03 | 94.4 | 98.6 | # | # | # | HIC | 2021-02-01 | 2021-03-01 | Alpha
39 | f157_2_0_0 | f157 | 2 | 0 | 0 | Osterman, A., et al | LIAISON | CLIA | yes | # | lab based | # | yes | yes | cohort | retrospective | Germany | clinical | adult | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Cobas; GeneXpert | Seegen Allplex, Cobas, Xpert Xpress | unclear | unclear | NP/OP/URT | NP | IFU | 410 | # | # | 410 | 0 | 410 | # | # | # | # | # | # | # | 6045 | # | 83 | 1548572803 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 21 | 1 | 302 | 86 | 19.63 | 12.6 | 28.4 | 99.67 | 98.15 | 99.98 | # | # | # | HIC | 2021-02-01 | 2021-03-01 | Alpha
40 | f157_3_0_0 | f157 | 3 | 0 | 0 | Osterman, A., et al | Euroimmun | ELISA | yes | # | lab based | # | yes | yes | cohort | retrospective | Germany | clinical | adult | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Cobas; GeneXpert | Seegen Allplex, Cobas, Xpert Xpress | unclear | unclear | NP/OP/URT | NP | IFU | 410 | # | # | 410 | 0 | 410 | # | # | # | # | # | # | # | 6045 | # | 83 | 1548572803 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 19 | 1 | 302 | 88 | 17.76 | 11 | 26.3 | 99.67 | 98.15 | 99.98 | # | # | # | HIC | 2021-02-01 | 2021-03-01 | Alpha
41 | f157_4_0_0 | f157 | 4 | 0 | 0 | Osterman, A., et al | Elecsys | ECLIA | yes | # | lab based | # | yes | yes | cohort | retrospective | Germany | clinical | adult | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Cobas; GeneXpert | Seegen Allplex, Cobas, Xpert Xpress | unclear | unclear | NP/OP/URT | NP | IFU | 408 | # | # | 408 | 0 | 410 | # | # | # | # | # | # | # | 6045 | # | 83 | 1548572803 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 33 | 7 | 296 | 72 | 31.42 | 22.7 | 41.2 | 97.69 | 95.31 | 99.1 | # | # | # | HIC | 2021-02-01 | 2021-03-01 | Alpha
42 | f159_1_0_0 | f159 | 1 | 0 | 0 | Salvagno, G., L., et al. | LIAISON | CLIA | yes | 42 | lab based | 136 | no | yes | cohort | prospective | Italy | clinical | unclear | unclear | 48 | # | # | # | 31 | 59 | # | 227 | RT-PCR | RealStar | Realstar | NP | NP | NP | NP | non-IFU | 421 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 93 | 0 | 120 | 208 | 30.9 | 25.7 | 36.5 | 100 | 97 | 100 | # | # | # | HIC | 2021-03-29 | 2021-04-18 | Alpha
43 | f165_1_0_0 | f165 | 1 | 0 | 0 | Baccani, I., et al. | Lumipulse G | CLEIA | yes | 30 | lab based | # | no | yes | unclear | retrospective | Italy | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | NP | NP | NP | NP | non-IFU | 201 | # | # | # | # | # | 30.825 | 30.5061 | 15.99 | 39.37 | 27.61 | 35.6125 | 14.4321 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | pg/ml | 0.37 | 250.9211 | 0.01 | 5000 | 0.19 | 2.62 | 1022.738 | 29 | 7 | 161 | 4 | 87.9 | 71.8 | 96.6 | 95.8 | 91.6 | 98.3 | # | # | # | HIC | unclear | unclear | unclear
44 | f165_2_0_0 | f165 | 2 | 0 | 0 | Baccani, I., et al. | STANDARD F | FIA | yes | 20 | POC | # | no | yes | unclear | retrospective | Italy | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | NP | NP | NP | NP | non-IFU | 93 | # | # | # | # | # | 30.825 | 30.5061 | 15.99 | 39.37 | 27.61 | 35.6125 | 14.4321 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | 0.21 | 4.7696 | 0 | 87.12 | 0.12 | 0.34 | 18.3601 | 10 | 0 | 65 | 18 | 35.7 | 18.6 | 55.9 | 100 | 94.5 | 100 | # | # | # | HIC | unclear | unclear | unclear
45 | f165_3_0_0 | f165 | 3 | 0 | 0 | Baccani, I., et al. | AFIAS | FIA | yes | 12 | POC | # | no | yes | unclear | retrospective | Italy | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | NP | NP | NP | NP | non-IFU | 81 | # | # | # | # | # | 30.825 | 30.5061 | 15.99 | 39.37 | 27.61 | 35.6125 | 14.4321 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | 0.14 | 3.396 | 0 | 47.76 | 0.1 | 0.19 | 10.0879 | 9 | 0 | 57 | 15 | 37.5 | 18.8 | 59.4 | 100 | 93.7 | 100 | # | # | # | HIC | unclear | unclear | unclear
46 | f166_1_0_0 | f166 | 1 | 0 | 0 | Hartard, C., et al. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | cohort | unclear | France | testing site | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | MGISP-960 | NP | NP | NP | NP | non-IFU | 378 | 99 | 221 | # | # | # | 18.805 | 19.394 | 11.947 | 32.971 | 16.2415 | 20.6457 | 4.7558 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 8864.65 | 14832.9731 | 26.426 | 83476 | 874.8025 | 32693.75 | 25651.1018 | 39 | 2 | 330 | 7 | 84.8 | 71.1 | 93.7 | 99.4 | 97.8 | 99.9 | # | # | # | HIC | unclear | unclear | Alpha, Beta
47 | f167_1_0_0 | f167 | 1 | 0 | 0 | Kahn, M., et al. | STANDARD F | FIA | yes | 10-30 | POC | # | no | yes | cohort | retrospective | Germany | clinical | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Cobas; Hologic; GeneXpert; RealStar | Cobas, GeneXpert, Realstar | OP | OP | OP | OP | IFU | 3110 | 2340 | 249 | # | # | # | # | 26.5 | 13 | 42.7 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 57 | 31 | 2983 | 39 | 59.4 | 48.9 | 69.3 | 99 | 98.5 | 99.3 | # | # | # | HIC | 2020-06-23 | 2020-11-26 | wildtype
48 | f171_1_0_0 | f171 | 1 | 0 | 0 | Iqbal, B., et al. | Elecsys | ECLIA | yes | 15-18 | lab based | # | no | yes | cross-sectional | prospective | Pakistan | clinical | unclear | symptomatic | # | 45 | 14 | 89 | # | # | # | 87 | RT-PCR | TibMolBiol | TibMol | NP | NP | nasal/OP | AN/MT | non-IFU | 170 | 0 | 170 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 92 | 1 | 42 | 35 | 72.4 | 63.8 | 80 | 97.7 | 87.7 | 99.9 | # | # | # | LMIC | 2021-04-01 | 2021-04-30 | Alpha
49 | f173_1_0_0 | f173 | 1 | 0 | 0 | Levett, P., N., et al. | VITROS | CLIA | yes | # | lab based | # | no | yes | case-control | retrospective | Canada | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | unclear | unclear | non-IFU | 24 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 10 | 0 | 12 | 2 | 83.3 | 51.6 | 97.9 | 100 | NA | NA | # | 0.5 | TCID 50 /ml | HIC | unclear | unclear | unclear
50 | f173_2_0_0 | f173 | 2 | 0 | 0 | Levett, P., N., et al. | VITROS | CLIA | yes | # | lab based | # | no | yes | cohort | prospective | Canada | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | unclear | unclear | non-IFU | 528 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 59 | 0 | 458 | 11 | 84.3 | 73.6 | 91.9 | 100 | 99.2 | 100 | # | 0.5 | TCID 50 /ml | HIC | unclear | unclear | unclear
51 | f184_1_0_0 | f184 | 1 | 0 | 0 | Audige, A., et al. | Elecsys | ECLIA | yes | # | lab based | # | no | yes | case-control | retrospective | Switzerland | unclear | adult, children | symptomatic, high risk contact, screening independent of symptoms | 35 | 39 | 13 | 98 | 26 | 50 | 17 | 189 | RT-PCR | Cobas | Cobas | NP | NP | NP | NP | IFU | 407 | 61 | 338 | # | # | # | 22.79 | 23.88 | 13.98 | 39.05 | 19.53 | 27.2 | 5.59 | # | # | # | # | # | # | # | 2 | 2.88 | 1 | 28 | 1 | 3 | 2.67 | # | # | # | # | # | # | # | # | 163 | 0 | 220 | 24 | 87.2 | 81.5 | 91.6 | 100 | 98.3 | 100 | # | # | # | HIC | 2020-10-20 | 2021-02-28 | wildtype, Alpha
52 | f184_2_0_0 | f184 | 2 | 0 | 0 | Audige, A., et al. | Elecsys | ECLIA | yes | # | lab based | # | no | yes | case-control | retrospective | Switzerland | unclear | adult, children | symptomatic, high risk contact, screening independent of symptoms | 32 | 35 | 16 | 77 | 25 | 41 | 14 | 134 | RT-PCR | Cobas | Cobas | saliva | saliva | saliva | saliva | IFU | 307 | 56 | 244 | # | # | # | 30.28 | 29.16 | 15.32 | 38.16 | 26.11 | 32.97 | 5.24 | # | # | # | # | # | # | # | 2 | 3.02 | 1 | 15 | 1 | 4 | 2.44 | # | # | # | # | # | # | # | # | 35 | 1 | 219 | 52 | 40.2 | 29.9 | 51.3 | 99.5 | 97.4 | 100 | # | # | # | HIC | 2020-10-20 | 2021-02-28 | wildtype, Alpha
53 | f191_1_0_0 | f191 | 1 | 0 | 0 | Orsi, A., et al. | FREND | FIA | yes | 3 | POC | # | yes | yes | cohort | prospective | Italy | clinical | unclear | symptomatic | # | # | # | # | # | # | # | 110 | RT-PCR | Allplex | Seegene | NP | NP | NP | NP | non-IFU | 110 | 0 | 110 | # | # | # | 26 | 25.78 | 18 | 34 | 23 | 28 | 3.89 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 56 | 0 | 50 | 4 | 93.3 | 83.8 | 98.2 | 100 | 92.9 | 100 | # | # | # | HIC | 2020-11-01 | 2020-11-30 | wildtype
54 | f191_2_0_0 | f191 | 2 | 0 | 0 | Orsi, A., et al. | STANDARD F | FIA | yes | 30 | POC | # | yes | yes | cohort | prospective | Italy | clinical | unclear | symptomatic | # | # | # | # | # | # | # | 110 | RT-PCR | Allplex | Seegene | NP | NP | NP | NP | non-IFU | 110 | 0 | 110 | # | # | # | 26 | 25.78 | 18 | 34 | 23 | 28 | 3.89 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 52 | 0 | 50 | 8 | 86.7 | 75.4 | 94.1 | 100 | 92.9 | 100 | # | # | # | HIC | 2020-11-01 | 2020-11-30 | wildtype
55 | f193_1_0_0 | f193 | 1 | 0 | 0 | Weitzel, T., et al. | STANDARD F | FIA | yes | # | POC | # | no | yes | cohort | prospective | Chile | clinical | unclear | screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | XpertXpress; Smplexa | XpertXpress; Smplexa | NP | NP | NP | NP | IFU | 773 | 773 | 0 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0 | 67 | 706 | 0 | NA | NA | NA | 91.3 | 89.1 | 93.2 | # | # | # | HIC | 2020-09-01 | 2021-01-01 | wildtype
56 | f194_1_0_0 | f194 | 1 | 0 | 0 | Suzuki, H., et al. | RapidTesta | DIA | yes | 10 | POC | # | yes | no | cohort | prospective | Japan | testing site | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | inhouse | inhouse | NP | NP | NP | NP | non-IFU | 1127 | 325 | 805 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 58 | 25 | 1028 | 16 | 78.4 | 67.3 | 87.1 | 97.6 | 96.5 | 98.5 | # | # | # | HIC | 2021-04-20 | 2021-05-31 | Alpha
57 | f196_1_0_0 | f196 | 1 | 0 | 0 | Cento, V., et al. | LumiraDx | FIA | yes | # | POC | # | no | yes | cohort | prospective | Italy | clinical | adult | symptomatic, screening independent of symptoms | 66 | # | # | # | 45 | 79 | # | 405 | RT-PCR | Simplexa | Simplexa | NP | NP | NP | NP | non-IFU | 960 | # | # | # | 0 | 960 | 30.8 | # | # | # | 26.1 | 32.8 | # | # | # | # | # | # | # | # | 6 | # | # | # | 4 | 7 | # | # | # | # | # | # | # | # | # | 297 | 17 | 596 | 50 | 85.6 | 81.4 | 89.1 | 97.2 | 95.6 | 98.4 | # | # | # | HIC | 2020-10-21 | 2020-12-09 | wildtype
58 | f197_1_0_0 | f197 | 1 | 0 | 0 | Koskinen, J., et al. | mariPOC | FIA | yes | 20 | POC | 12.5 | no | no | case-control | prospective | Finland | clinical | adult, children | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Amplidiag | Allplex, Amplidiag | NP | NP | NP | NP | IFU | 13 | # | 13 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 13 | NA | NA | 0 | 100 | 75.3 | 100 | NA | NA | NA | # | 2.7 | TCID50/test | HIC | unclear | unclear | unclear
59 | f197_2_0_0 | f197 | 2 | 0 | 0 | Koskinen, J., et al. | mariPOC | FIA | yes | 20-55 | POC | 12.5 | no | no | case-control | prospective | Finland | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Allplex; Amplidiag | Allplex, Amplidiag | NP | NP | NP | NP | IFU | 45 | # | # | # | # | # | 25.3 | 25.7509 | 34 | # | 23.02 | 29.03 | 4.2923 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 38 | NA | NA | 7 | 84.4 | 70.5 | 93.5 | NA | NA | NA | # | 2.7 | TCID50/test | HIC | unclear | unclear | unclear
60 | f198_1_0_0 | f198 | 1 | 0 | 0 | Tanimoto., et al. | Lumipulse G | CLEIA | yes | 35-60 | lab based | 120 | no | yes | case-control | prospective | Japan | unclear | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | StephOnePlus | StephOnePlusReal | NPS | NP | NPS | NP | IFU | 68 | # | # | # | # | # | 25.45 | 26.73 | 14.8 | 41.7 | 19.8 | 31.6 | 8.062 | 13115 | 1054422 | 0.2669 | 15478184 | 234.6 | 561491 | 2846383 | # | # | # | # | # | # | # | log10pg/mL | # | 11887184.8 | 0.01 | 138298 | # | # | # | 32 | 1 | 27 | 8 | 80 | 64.4 | 90.9 | 96.4 | 81.7 | 99.9 | # | 42.8 | RNAcopies | HIC | 2020-07-01 | 2020-11-30 | wildtype
61 | f198_2_0_0 | f198 | 2 | 0 | 0 | Tanimoto., et al. | Lumipulse G | CLEIA | yes | 35-60 | lab based | 120 | no | yes | case-control | prospective | Japan | unclear | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | StephOnePlus | StephOnePlusReal | saliva | saliva | saliva | saliva | IFU | 116 | # | # | # | # | # | 30.8 | 30.43 | 18.5 | 42.7 | 25.46 | 36.5 | 6.501 | 373.4 | 1054422 | 0.2669 | 15478184 | 234.6 | 561491 | 2846383 | # | # | # | # | # | # | # | log10pg/mL | # | 11887184.8 | 0.01 | 138298 | # | # | # | 34 | 0 | 61 | 21 | 61.8 | 47.7 | 74.6 | 100 | 94.1 | 100 | # | 42.8 | RNAcopies | HIC | 2020-07-01 | 2020-07-01 | wildtype
62 | f20_1_0_0 | f20 | 1 | 0 | 0 | Osterman, A., et al. | STANDARD F | FIA | yes | 15-30 | POC | # | yes | yes | cohort | retrospective | Germany | both | adult, children | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | XpertXpress, TaqPath | Xpress, TaqPath | NP/OP | NP | NP/OP | NP | IFU | 549 | # | # | # | 60 | 300 | 27 | # | 14 | 40 | # | # | # | 4545000 | # | # | # | 1228750 | 21773400 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 115 | 8 | 352 | 74 | 60.85 | 53.5 | 67.8 | 97.78 | 95.68 | 99 | # | # | # | HIC | 2020-03-04 | 2020-10-19 | wildtype
63 | f31_1_0_0 | f31 | 1 | 0 | 0 | Favresse, J., et al. | VITROS | CLIA | yes | 48 | lab based | 150 | yes | yes | cohort | prospective | Belgium | clinical | adult, children | symptomatic, screening independent of symptoms | 55 | # | 1 | 97 | # | # | # | 104 | RT-PCR | RealStar | RealStar | NP | NP | NP | NP | IFU | 188 | 118 | 70 | # | # | # | 22.3 | # | 12.6 | 38.2 | # | # | # | # | # | # | # | # | # | # | 3 | # | # | # | 2 | 4 | # | # | # | # | # | # | # | # | # | 80 | 0 | 92 | 16 | 83.3 | 74.4 | 90.2 | 100 | 96.1 | 100 | # | # | # | HIC | 2020-11-17 | 2020-11-25 | wildtype
64 | f32_1_0_0 | f32 | 1 | 0 | 0 | Kohmer, N., et al. | LumiraDx | FIA | yes | 12 | POC | # | yes | yes | cohort | prospective | Germany | testing site | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas | Cobas 6800 | NP | NP | NP | NP | IFU | 100 | # | # | # | # | # | 29.38 | 29.17 | 22.13 | 36.46 | 25.32 | 33.09 | 4.11 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 37 | 0 | 26 | 37 | 50 | 38.1 | 61.9 | 100 | 86.8 | 100 | # | # | # | HIC | 2020-11-01 | 2020-11-14 | wildtype
65 | f34_1_0_0 | f34 | 1 | 0 | 0 | Caruana, G., et al. | BD Veritor | DIA | yes | 20 | POC | # | yes | yes | cohort | prospective | Switzerland | clinical | unclear | symptomatic, screening independent of symptoms | 73 | # | # | # | 54.8 | 84 | # | 236 | RT-PCR | unclear | unclear | NP | NP | NP | NP | IFU | 532 | 239 | 293 | # | # | # | 27.5 | # | # | # | 20.6 | 33.3 | # | 539000 | # | # | # | 15000 | 32500000 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 47 | 1 | 417 | 67 | 41.2 | 32.1 | 50.8 | 99.8 | 98.7 | 100 | # | # | # | HIC | 2020-11-07 | 2020-12-16 | wildtype
66 | f34_2_0_0 | f34 | 2 | 0 | 0 | Caruana, G., et al. | Exdia | FIA | yes | 20 | POC | # | yes | yes | cohort | prospective | Switzerland | clinical | unclear | symptomatic, screening independent of symptoms | 73 | # | # | # | 54.8 | 84 | # | 236 | RT-PCR | GeneXpert; Cobas; Viasure | GeneXpert, Cobas 6800;Viasue | NP | NP | NP | NP | IFU | 532 | 239 | 293 | # | # | # | 27.5 | # | # | # | 20.6 | 33.3 | # | 539000 | # | # | # | 15000 | 32500000 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 55 | 2 | 416 | 59 | 48.3 | 38.8 | 57.8 | 99.5 | 98.3 | 99.9 | # | # | # | HIC | 2020-11-07 | 2020-12-16 | wildtype
67 | f39_1_0_0 | f39 | 1 | 0 | 0 | FINDdx | ichroma | FIA | yes | # | POC | # | no | no | cohort | prospective | Switzerland | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | # | 36.3 | 16 | 76 | # | # | # | 129 | RT-PCR | Cobas | Cobas | NP | NP | NP | NP | IFU | 232 | 200 | 32 | # | 0 | 232 | 22.7 | # | # | # | 18.9 | 28.9 | # | # | # | # | # | # | # | # | 1.5 | # | # | # | 1 | 3 | # | # | # | # | # | # | # | # | # | 30 | 0 | 191 | 11 | 73.2 | 57.1 | 85.8 | 100 | 98 | 100 | # | 1,66 x 10^3 pfu/ml | # | HIC | 2021-01-21 | 2021-01-29 | wildtype, Alpha
68 | f43_1_0_0 | f43 | 1 | 0 | 0 | Drain, P.K., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | UK/USA | both | adult, children | symptomatic, screening independent of symptoms | # | 34 | 0 | 90 | # | # | 14.5 | 142 | RT-PCR | Cobas | Cobas 6800 | AN | AN/MT | AN | AN/MT | IFU | 257 | 98 | 159 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 4 | # | # | # | # | 2.9 | # | # | # | # | # | # | # | # | 81 | 6 | 168 | 2 | 97.6 | 91.6 | 99.7 | 96.6 | 92.6 | 98.7 | # | 32 | TCID50/ml | HIC | 2020-06-26 | 2020-07-23 | wildtype
69 | f43_2_0_0 | f43 | 2 | 0 | 0 | Drain, P.K., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | UK/USA | both | adult, children | symptomatic, screening independent of symptoms | # | 33.2 | 0 | 90 | # | # | 19.4 | 145 | RT-PCR | Cobas | Cobas 6800 | NP | NP | NP | NP | IFU | 255 | 0 | 255 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 3.5 | # | # | # | # | 2.5 | # | # | # | # | # | # | # | # | 39 | 5 | 210 | 1 | 97.5 | 86.8 | 99.9 | 97.7 | 94.7 | 99.2 | # | 32 | TCID50/ml | HIC | 2020-08-17 | 2020-09-28 | wildtype
70 | f46_1_0_0 | f46 | 1 | 0 | 0 | Marti, J.L.G., et al. | Sofia | FIA | yes | 16 | POC | # | no | no | cohort | unclear | USA | clinical | children | symptomatic | # | # | # | 17 | # | # | # | # | RT-PCR | XpertXpress | XpertXpress | MT | AN/MT | MT | AN/MT | IFU | 427 | # | # | # | 427 | 0 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 31 | 1 | 383 | 12 | 72 | 56.3 | 84.7 | 99.7 | 98.6 | 100 | # | # | # | HIC | 2020-03-08 | 2020-09-10 | wildtype
71 | f46_2_0_0 | f46 | 2 | 0 | 0 | Marti, J.L.G., et al. | BD Veritor | DIA | yes | 16 | POC | # | no | no | cohort | unclear | USA | testing site | adult | symptomatic | # | # | 18 | 24 | # | # | # | # | RT-PCR | XpertXpress | XpertXpress | AN | AN/MT | AN | AN/MT | IFU | 32 | # | # | # | 0 | 32 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 30 | NA | NA | 2 | 93.75 | 79.2 | 99.2 | NA | NA | NA | # | # | # | HIC | 2020-03-08 | 2020-09-10 | wildtype
72 | f47_1_0_0 | f47 | 1 | 0 | 0 | Hirotsu, Y., et al., | Lumipulse G | CLEIA | yes | 30 | lab based | 120 | no | yes | cohort | prospective | Japan | clinical | unclear | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | NIID Japan | StepOnePlus Real-Time PCR System | NP | NP | NP | NP | IFU | 1029 | 0 | 36 | # | # | # | # | 36.08 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 0 | 30 | # | # | # | # | # | 0,000853 | # | # | # | # | # | 37 | 0 | 989 | 3 | 92.5 | 79.6 | 98.4 | 100 | 99.6 | 100 | # | # | # | HIC | unclear | unclear | unclear
73 | f50_1_0_0 | f50 | 1 | 0 | 0 | Jaaskelainen, A.E., et al. | Sofia | FIA | yes | 15 | POC | # | yes | yes | case-control | retrospective | Finland | both | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | inhouse | lab developed | NP/nasal | NP | NP/nasal | NP | IFU | 188 | # | # | # | # | # | 20.76 | 21.0567 | 10.74 | 32.49 | 17.0975 | 23.91 | 4.989 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 119 | 0 | 40 | 29 | 80.4 | 73.1 | 86.5 | 100 | 91.2 | 100 | # | # | # | HIC | 2020-04-01 | 2020-11-01 | wildtype
74 | f57_1_0_0 | f57 | 1 | 0 | 0 | Gili, A., et al. | Lumipulse G | CLEIA | yes | 60 | lab based | 120 | no | yes | case-control | unclear | Italy | unclear | unclear | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex SARS-CoV-2 assay | NP | NP | NP | NP | IFU | 226 | # | # | # | # | # | 29 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 86 | 11 | 120 | 9 | 90.5 | 82.8 | 95.6 | 91.6 | 85.5 | 95.7 | 35 | 1.34 | ng/mL | HIC | 2021-01-01 | 2021-01-26 | wildtype
75 | f57_2_0_0 | f57 | 2 | 0 | 0 | Gili, A., et al. | Lumipulse G | CLEIA | yes | 60 | lab based | 120 | no | yes | cohort | unclear | Italy | testing site | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex SARS-CoV-2 assay | NP | NP | NP | NP | IFU | 1738 | # | # | # | # | # | 22 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 90 | 130 | 1518 | 0 | 100 | 96 | 100 | 92.1 | 90.7 | 93.4 | 35 | 1.34 | ng/mL | HIC | 2020-09-10 | 2020-09-15 | wildtype
76 | f58_1_0_0 | f58 | 1 | 0 | 0 | Kruger, L.J., et al. | LumiraDx | FIA | yes | 12 | POC | 3.5 | no | no | cohort | Prospective | Germany | testing site | adult | symptomatic, high risk contact | # | 38.5 | # | # | # | # | 14.2 | 396 | RT-PCR | Allplex; Cobas; TibMolBiol | Allplex, Cobas 6800 or Cobas 8800, TIB Molbiol | NMT | AN/MT | NP | NP | IFU | 761 | 272 | 486 | # | 0 | 761 | # | # | # | # | # | # | # | 7.3 | # | # | # | 3.7 | 9.4 | # | # | 3.9 | # | # | # | # | 3.2 | # | # | # | # | # | # | # | # | 120 | 4 | 611 | 26 | 82.2 | 75 | 88 | 99.3 | 98.3 | 99.7 | # | # | # | HIC | 2020-11-02 | 2021-01-21 | wildtype
77 | f70_1_0_0 | f70 | 1 | 0 | 0 | Lefever, S., et al. | LIAISON | CLIA | yes | # | lab based | # | no | yes | case-control | retrospective | Belgium | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | 37 | # | 2 | 101 | # | # | # | 249 | RT-PCR | Viasure | VIASURE SARS-CoV-2 | NP | NP | NP | NP | IFU | 414 | 178 | 223 | # | # | # | # | # | 10.2 | 39 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 138 | 0 | 210 | 66 | 67.6 | 60.8 | 74 | 100 | 98.3 | 100 | # | # | # | HIC | 2020-11-20 | 2020-12-18 | wildtype
78 | f71_1_0_0 | f71 | 1 | 0 | 0 | Kilic, A., et al. | BD Veritor | DIA | yes | 15 | POC | # | no | yes | cohort | prospective | USA | clinical | adult, children | symptomatic | 46.8 | # | 1 | 98 | # | # | # | 800 | RT-PCR | Simplexa | Simplexa COVID-19 Direct EUA RT-PCR | nasal | AN/MT | NP | NP | IFU | 1384 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 77 | 15 | 1253 | 39 | 66.4 | 57 | 74.9 | 98.8 | 98.1 | 99.3 | >=40 | # | # | HIC | 2020-10-20 | 2020-12-03 | wildtype
79 | f74_1_0_0 | f74 | 1 | 0 | 0 | Asai, N., et al. | Lumipulse G | CLEIA | yes | 30-45 | lab based | # | no | yes | cohort | prospective | Japan | clinical | unclear | symptomatic, high risk contact | # | # | # | # | # | # | # | # | RT-PCR | 2019-nCov CDC EUA Kit | BD MAX system | saliva | saliva | saliva | saliva | IFU | 305 | # | # | # | # | # | 26.6 | # | 15.5 | 36.2 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 18.6 | 135.0478 | 1.14 | 17046.6 | # | # | # | 49 | 4 | 238 | 14 | 77.8 | 65.5 | 87.3 | 98.3 | 95.8 | 99.5 | # | # | # | HIC | 2020-12-01 | 2021-01-31 | wildtype
80 | f77_1_0_0 | f77 | 1 | 0 | 0 | Micocci, M., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | no | cohort | prospective | UK | clinical | adult | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | NP | NP | NP | NP | IFU | 239 | 225 | 14 | # | 0 | 239 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 6 | 9 | 222 | 2 | 75 | 34.9 | 96.8 | 96.1 | 92.7 | 98.2 | # | # | # | HIC | unclear | unclear | unclear
81 | f78_1_0_0 | f78 | 1 | 0 | 0 | NÃ¶rz, D., et al. | Elecsys | ECLIA | yes | 18 | lab based | # | no | no | cohort | retrospective | Germany | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas | Cobas 6800 | NP/OP | NP | NP/OP | NP | non-IFU | 3139 | 79 | 313 | # | # | # | # | # | # | # | # | # | # | 2,95*10^4 | # | # | # | 5,1*10^3 | 3,5*10^6 | # | # | # | # | # | # | # | # | COI | 46 | # | # | # | 6 | 701.3 | # | 236 | 4 | 2743 | 156 | 60.2 | 55.2 | 65.1 | 99.9 | 99.6 | 100 | # | # | # | HIC | 2020-11-01 | 2020-11-30 | wildtype
82 | f82_1_0_0 | f82 | 1 | 0 | 0 | Yin, N., et al. | BD Veritor | DIA | yes | 10-20 | POC | # | no | yes | cohort | retrospective | Belgium | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Cobas | Cobas m2000 | NP | NP | NP | NP | IFU | 177 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | 87.7 | 80 | 92.7 | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-09-01 | wildtype
83 | f86_1_0_0 | f86 | 1 | 0 | 0 | Bruzzone, B., et al. | STANDARD F | FIA | yes | # | POC | # | no | yes | cohort | retrospective | Italy | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | unclear | unclear | unclear | unclear | IFU | 60 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 52 | 0 | NA | 8 | 86.7 | 75.4 | 94.1 | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-12-01 | wildtype
84 | f86_2_0_0 | f86 | 2 | 0 | 0 | Bruzzone, B., et al. | LumiraDx | FIA | yes | # | POC | # | no | yes | cohort | retrospective | Italy | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | unclear | unclear | unclear | unclear | IFU | 23 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 16 | 0 | NA | 7 | 69.6 | 47.1 | 86.8 | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-12-01 | wildtype
85 | f86_3_0_0 | f86 | 3 | 0 | 0 | Bruzzone, B., et al. | FREND | FIA | yes | # | POC | # | no | yes | cohort | retrospective | Italy | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex | unclear | unclear | unclear | unclear | IFU | 60 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 56 | 0 | NA | 4 | 93.3 | 83.8 | 98.2 | NA | NA | NA | # | # | # | HIC | 2020-07-01 | 2020-12-01 | wildtype
86 | f98_1_0_0 | f98 | 1 | 0 | 0 | Brihn, A., et al. | Sofia | FIA | yes | # | POC | # | no | yes | cohort | prospective | USA | clinical | adult | symptomatic, screening independent of symptoms | 56 | # | 16 | 107 | # | # | # | 1126 | RT-PCR | Fulgent | Fulgent COVID-19 RT-PCR | nasal | AN/MT | NP | NP | non-IFU | 2039 | 1732 | 307 | # | # | 2038 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 98 | 12 | 1878 | 51 | 65.8 | 57.6 | 73.3 | 99.4 | 98.9 | 99.7 | # | # | # | HIC | 2020-06-30 | 2020-08-31 | wildtype
87 | k1_1_0_0 | k1 | 1 | 0 | 0 | Wang, H., et.al. | MSD S-PLEX | ECL | yes | 240-300 | lab based | 120 | no | yes | case-control | prospective | USA | clinical | adult, children | unclear | # | # | # | # | # | # | # | # | RT-PCR | Panther Fusion | Hologic Panther fusion assay | NP,mid-turbinate | NP | NP,mid-turbinate | NP | IFU | 200 | # | # | # | # | # | 23.3 | # | # | # | 19.2 | 32.8 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 88 | 7 | 93 | 12 | 88 | 80 | 93.6 | 93 | 86.1 | 97.1 | # | 282 | fg/mL | HIC | 2020-11-13 | 2020-11-15 | wildtype
88 | k1_2_0_0 | k1 | 2 | 0 | 0 | Wang, H., et.al. | MSD S-PLEX | ECL | yes | 240-300 | lab based | 120 | no | yes | case-control | retrospective | USA | clinical | adult, children | unclear | # | # | # | # | # | # | # | # | RT-PCR | Panther Fusion | Hologic Panther fusion assay | NP,mid-turbinate,OP | NP | NP,mid-turbinate,OP | NP | IFU | 250 | # | # | # | # | # | 33.2 | # | # | # | 26 | 36.9 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 51 | 7 | 173 | 19 | 72.9 | 60.9 | 82.8 | 96.1 | 92.2 | 98.4 | # | 282 | fg/mL | HIC | 2020-03-01 | 2020-11-10 | wildtype
89 | k10_1_0_0 | k10 | 1 | 0 | 0 | Leiner, J., et al. | STANDARD F | FIA | yes | # | POC | # | no | yes | cohort | prospective | Germany | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex TM 2019-nCoV | OP, nasal | AN/MT | OP, nasal | AN/MT | non-IFU | 4076 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 491 | 80 | 3208 | 297 | 62.3 | NA | NA | 97.6 | NA | NA | # | # | # | HIC | 2020-10-01 | 2021-01-31 | wildtype
90 | k11_1_0_0 | k11 | 1 | 0 | 0 | Loconsole, D., et al. | Lumipulse G | CLEIA | yes | 50-60 | lab based | 120 | no | yes | cohort | prospective | Italy | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | 52 | # | # | # | 35 | 68 | # | 442 | RT-PCR | TaqPath | TaqPath | NP | NP | NP | NP | non-IFU | 911 | 550 | 361 | # | # | # | # | 22.6409 | # | # | # | # | # | # | # | # | # | # | # | # | # | 4.5 | 0 | 25 | # | # | # | # | # | 1.7888 | 1.34 | 5000 | # | # | # | 205 | 18 | 677 | 11 | 94.9 | 91.1 | 97.4 | 97.4 | 95.9 | 98.5 | # | # | # | HIC | 2020-11-23 | 2020-11-04 | wildtype
91 | k12_1_0_0 | k12 | 1 | 0 | 0 | Mitchell, S. L., et al. | Sofia | FIA | yes | # | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic, high risk contact, screening independent of symptoms | # | 38.6 | 17 | 80 | # | # | # | 106 | RT-PCR | XpertXpress | Cepheid GeneXpert Xpress | nasal | AN/MT | NP | NP | IFU | 144 | 144 | 0 | # | 0 | 144 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2 | 0 | 138 | 4 | 33.3 | NA | NA | 100 | NA | NA | # | # | # | HIC | 2021-01-01 | 2021-03-31 | wildtype
92 | k12_2_0_0 | k12 | 2 | 0 | 0 | Mitchell, S. L., et al. | Sofia | FIA | yes | # | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic, high risk contact, screening independent of symptoms | # | 44.2 | 18 | 83 | # | # | # | 93 | RT-PCR | XpertXpress | Cepheid GeneXpert Xpress | nasal | AN/MT | NP | NP | IFU | 148 | 0 | 148 | # | 0 | 148 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 3.6 | 1 | 7 | # | # | # | # | # | # | # | # | # | # | # | 36 | 0 | 107 | 5 | 87.8 | NA | NA | 100 | NA | NA | # | # | # | HIC | 2021-01-01 | 2021-03-31 | wildtype
93 | k13_1_0_0 | k13 | 1 | 0 | 0 | Montalvo Villalba, M.C., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | cohort | prospective | Cuba | unclear | adult, children | high risk contact | # | # | 0.4 | 96 | # | # | # | # | RT-PCR | STAT‐NAT | STAT‐NAT® COVID‐19 MULTI | NP | NP | NP | NP | IFU | 132 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | # | # | # | # | # | # | # | 52 | 2 | 67 | 11 | 82.5 | 70.9 | 90.9 | 97.06 | 89.9 | 99.5 | 40 | # | # | LMIC | 2021-03-01 | 2021-03-31 | wildtype
94 | k13_2_0_0 | k13 | 2 | 0 | 0 | Montalvo Villalba, M.C., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | cohort | prospective | Cuba | unclear | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | # | 0.4 | 96 | # | # | # | # | RT-PCR | STAT‐NAT | STAT‐NAT® COVID‐19 MULTI | NP | NP | NP | NP | IFU | 59 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | # | # | # | # | # | # | # | 27 | 2 | 23 | 7 | 79.41 | 62.1 | 91.3 | 92 | 74 | 99 | 40 | # | # | LMIC | 2021-03-01 | 2021-03-31 | wildtype
95 | k13_4_0_0 | k13 | 4 | 0 | 0 | Montalvo Villalba, M.C., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | cohort | prospective | Cuba | unclear | adult, children | screening independent of symptoms | # | # | 0.4 | 96 | # | # | # | # | RT-PCR | STAT‐NAT | STAT‐NAT® COVID‐19 MULTI | NP | NP | NP | NP | IFU | 119 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | # | # | # | # | # | # | # | 40 | 9 | 59 | 11 | 78.43 | 64.7 | 88.7 | 86.76 | 76.4 | 93.8 | 40 | # | # | LMIC | 2021-03-01 | 2021-03-31 | wildtype
96 | k14_1_0_0 | k14 | 1 | 0 | 0 | Paul, D., et al. | VITROS | CLIA | yes | 48 | lab based | 130 | yes | yes | cohort | prospective | India | clinical | adult, children | symptomatic, high risk contact | # | 34 | 12 | 87 | # | # | # | 52 | RT-PCR | RealStar | RealStar® SARS-CoV-2 | NP/OP | NP | NP | NP | IFU | 148 | 63 | 85 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2 | # | # | # | 0 | 7 | # | # | # | # | # | # | # | # | # | 72 | 0 | 50 | 26 | 73.47 | 63.59 | 81.88 | 100 | 92.89 | 100 | 35 | # | # | LMIC | unclear | unclear | unclear
97 | k15_1_0_0 | k15 | 1 | 0 | 0 | Sberna, G., et al. | Lumipulse G | CLEIA | yes | # | lab based | 120 | no | yes | cohort | prospective | Italy | both | adult, children | unclear | 47 | # | 1 | 98 | # | # | # | 259 | RT-PCR | Allplex; Easy SARS-CoV-2 WE kit | AllplexTM SARS-CoV-2 Assay, Seegene Inc., Seoul, Easy SARS-CoV-2 WE kit, Diatech Pharma- cogenetics srl, Jesi, Italy | NP | NP | NP | NP | IFU | 513 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 231 | 26 | 212 | 44 | 84 | 79.1 | 88.1 | 89.1 | 84.4 | 92.7 | 40 | >1,34 | pg/mL | HIC | unclear | unclear | unclear
98 | k17_1_0_0 | k17 | 1 | 0 | 0 | Wee, L. E., et al. | BD Veritor | DIA | yes | 12 | POC | # | no | yes | cohort | prospective | Singapore | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | unclear | unclear | non-IFU | 6665 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 29 | 11 | 6623 | 2 | 93.6 | 78.6 | 99.2 | 99.8 | 99.7 | 99.9 | # | # | # | HIC | 2021-06-27 | 2021-08-08 | Delta
99 | k19_1_1_0 | k19 | 1 | 1 | 0 | Pollock, N. R., et al. | MSD S-PLEX | ECL | yes | 300 | lab based | 120 | yes | yes | cohort | prospective | USA | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | 80 | RT-PCR | Panther Fusion | Hologic Panther fusion assay | NP | NP | NP | NP | IFU | 141 | 2 | 99 | 80 | 121 | 20 | 25.8 | 27.2594 | 13.2 | 38.4 | 19.1 | 36.2 | 8.3613 | # | # | # | # | # | # | # | 3 | 4.6979 | 0 | 39 | 2 | 3 | 5.7126 | # | 1.45 | 104290.08 | 0.16 | 2710823.85 | 0.16 | 1189.8498 | 404202.091 | NA | NA | NA | NA | 79 | 70 | 87 | 98 | 87 | 100 | # | 0.36 TCID50/ ml and 2,100 RNA copies/ml for the S-PLEX assay | # | HIC | unclear | unclear | unclear
100 | k19_2_1_0 | k19 | 2 | 1 | 0 | Pollock, N. R., et al. | MSD S-PLEX | ECL | yes | 300 | lab based | 120 | yes | yes | cohort | prospective | USA | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | 67 | RT-PCR | Panther Fusion | Hologic Panther fusion assay | NP | NP | NP | NP | IFU | 85 | 4 | 31 | 0 | 1 | 84 | 26.4 | 28.1163 | 13.2 | 39.6 | 18.95 | 32.65 | 7.8232 | # | # | # | # | # | # | # | 3 | 5.303 | 0 | 33 | 1 | 7 | 7.5394 | # | 0.18 | 57301.38 | 0.16 | 2473839.47 | 0.16 | 57.7894 | 291690.229 | NA | NA | NA | NA | 91 | 77 | 98 | 100 | 93 | 100 | # | 0.36 TCID50/ ml and 2,100 RNA copies/ml for the S-PLEX assay | # | HIC | unclear | unclear | unclear
101 | k2_1_0_0 | k2 | 1 | 0 | 0 | Al-Alawi, S., et al. | PCL | FIA | no | <30 | POC | # | no | yes | cohort | prospective | Oman | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Cobas; Liferiver; Sansuro Biotech | Liferiver Novel Coronavirus (2019-nCoV),Sansure Biotech COVID-19 Nucleic Acid Test Kit,Roche cobas® 6800 | NP and/or throat | NP | NP and/or throat | NP | IFU | 87 | 0 | 87 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | 69.8 | 55.66 | 81.66 | 94.1 | 80.32 | 99.28 | # | # | # | HIC | 2020-06-01 | 2020-07-31 | wildtype
102 | k2_2_0_0 | k2 | 2 | 0 | 0 | Al-Alawi, S., et al. | Sofia | FIA | yes | <30 | POC | # | no | yes | cohort | prospective | Oman | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | Cobas; Liferiver; Sansuro Biotech | Liferiver Novel Coronavirus (2019-nCoV),Sansure Biotech COVID-19 Nucleic Acid Test Kit,Roche cobas® 6800 | NP and/or throat | NP | NP and/or throat | NP | IFU | 78 | 0 | 78 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | 64.3 | 50.36 | 76.64 | 100 | 84.56 | 100 | # | # | # | HIC | 2020-06-01 | 2020-07-31 | wildtype
103 | k20_1_0_0 | k20 | 1 | 0 | 0 | FINDdx | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Germany | testing site | adult | symptomatic, high risk contact | # | 38.5 | 18 | 90 | # | # | # | 396 | RT-PCR | Cobas, RealTime, Allplex, LightMix, RealStar, TibMolibiol | Cobas, Abbott, Allplex, LightMix, RealStar, TibMolibiol | nasal,MT | AN/MT | NP | NP | IFU | 761 | 271 | 487 | # | 0 | 761 | 22.9 | # | # | # | 19.5 | 27.7 | # | # | # | # | # | # | # | # | 3 | # | # | # | 2 | 5 | # | # | # | # | # | # | # | # | # | 120 | 4 | 611 | 26 | 82.2 | 75 | 88 | 99.3 | 98.3 | 99.7 | # | 1,0x10^2 | pfu/ml | HIC | 2020-11-02 | 2020-11-27 | wildtype
104 | k20_2_0_0 | k20 | 2 | 0 | 0 | FINDdx | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Brazil | clinical | adult | symptomatic, high risk contact | # | 36.7 | 19 | 75 | # | # | # | 163 | RT-PCR | 2019-nCov CDC EUA Kit | Lab-developed assay based on the US CDC protocol | nasal | AN/MT | NP | NP | IFU | 251 | 7 | 243 | # | 0 | 249 | 17.43 | # | # | # | 15.1 | 21.6 | # | # | # | # | # | # | # | # | 3 | # | # | # | 2 | 5 | # | # | # | # | # | # | # | # | # | 50 | 2 | 191 | 8 | 86.2 | 74.6 | 93.9 | 99 | 96.3 | 99.9 | # | 5,0x10^2 | pfu/ml | LMIC | 2021-07-22 | 2021-08-02 | Gamma
105 | k3_1_0_0 | k3 | 1 | 0 | 0 | Adreani, J., et al. | Lumipulse G | CLEIA | yes | 35 | lab based | 120 | yes | yes | case-control | retrospective | France | unclear | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | LightCycler 480; Alinity M | LightCycler480 II platform, Alinity m SARS-CoV-2 assay | NP | NP | NP | NP | IFU | 239 | 35 | 204 | # | # | # | 22 | 24.0569 | 13 | 40 | 19 | 29 | 7.0202 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | log10pg/mL | 3.024 | 2.3665 | -2 | 3.699 | 1.3296 | 3.699 | 1.5837 | NA | NA | NA | NA | 87.6 | 81.7 | 93.5 | 72.7 | 63.4 | 82 | # | # | # | HIC | unclear | unclear | unclear
106 | k4_1_0_0 | k4 | 1 | 0 | 0 | Ben Abdelhanin, M., et al. | Elecsys | ECLIA | yes | 18 | lab based | # | no | yes | cohort | prospective | Belgium | unclear | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Allplex | Allplex TM 2019-nCoV | NP | NP | NP | NP | IFU | 225 | # | # | 0 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 81 | 0 | 102 | 42 | 65.85 | 56.76 | 74.16 | 100 | 96.49 | 100 | # | >1 | COI | HIC | 2021-04-01 | 2021-04-30 | Alpha
107 | k6_1_0_0 | k6 | 1 | 0 | 0 | Burdino, E., et al. | LumiraDx | FIA | yes | # | POC | # | no | yes | cohort | prospective | Italy | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Simplexa; Hologic; XpertXpress; TaqPath; Cobas; Allplex | DiaSorin Simplexa® (n = 523, 42.3 %), Panther Hologic® (n = 254, 20.6 %), Cepheid Xpert® (n = 186, 15.0 %), Thermofisher TaqPath RT PCR® (n = 166, 13.4 %), Cobas 6800 Roche® (n = 87, 7.0 %), and Seegene Allplex® 2019 n-CoV Assay (n = 16, 1.3 %) | nasal | AN/MT | NP | NP | non-IFU | 1232 | # | # | 0 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | 90.1 | 86.2 | 93.1 | 99.4 | 98.6 | 99.8 | # | # | # | HIC | 2021-03-01 | 2021-03-31 | Alpha
108 | k7_1_0_0 | k7 | 1 | 0 | 0 | Denina, M., et al. | LumiraDx | FIA | yes | 15 | POC | # | no | yes | cohort | retrospective | Italy | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | 88 | RT-PCR | Simplexa | DiaSorin Simplexa® | nasal | AN/MT | NP | NP | non-IFU | 191 | 36 | 155 | 80 | 191 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 1 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 16 | 14 | 160 | 1 | 94.1 | NA | NA | 91.9 | NA | NA | # | # | # | HIC | 2020-10-23 | 2021-01-23 | wildtype
109 | k8_1_0_0 | k8 | 1 | 0 | 0 | Drain, P., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | USA | clinical | adult, children | high risk contact | # | 38.7 | 0 | 88 | # | # | 17.3 | 140 | RT-PCR | Cobas; TaqPath | cobas 6800 Roche, TaqPath COVID-19 Combo Kit24 Thermofisher | nasal | AN/MT | nasal | AN/MT | non-IFU | 222 | 222 | 0 | # | 22 | 200 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 23 | 0 | 194 | 5 | 82.1 | 63.1 | 93.9 | 100 | 98.1 | 100 | # | # | # | HIC | 2020-11-13 | 2021-03-25 | wildtype
110 | k9_1_0_0 | k9 | 1 | 0 | 0 | Lanari, M., et al. | STANDARD F | FIA | yes | 10-30 | POC | # | no | yes | cohort | prospective | Italy | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | 4.6 | # | 18 | # | # | 4.5 | 180 | RT-PCR | Simplexa; XpertXpress; ELITe | DiaSorin Molecular SimplexaTM, Xpert® Xpress, SARS-CoV-2 ELITe MG- BKit | NP | NP | NP | NP | non-IFU | 381 | # | # | 381 | 381 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 14 | 1 | 354 | 12 | 53.8 | 33.4 | 73.4 | 99.7 | 98.4 | 100 | # | # | # | HIC | 2020-11-01 | 2021-04-09 | wildtype, Alpha
111 | k9_2_0_0 | k9 | 2 | 0 | 0 | Lanari, M., et al. | AFIAS | FIA | yes | 10-30 | POC | # | no | yes | cohort | prospective | Italy | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | 6.8 | # | 18 | # | # | 7.26 | 342 | RT-PCR | Simplexa; XpertXpress; ELITe | DiaSorin Molecular SimplexaTM, Xpert® Xpress, SARS-CoV-2 ELITe MG- BKit | NP | NP | NP | NP | non-IFU | 765 | # | # | 765 | 765 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 51 | 12 | 694 | 8 | 86.4 | 75 | 93.9 | 98.3 | 97.1 | 99.1 | # | # | # | HIC | 2020-11-01 | 2021-04-09 | wildtype, Alpha
112 | s1_1_0_0 | s1 | 1 | 0 | 0 | Agard, A., et al. | Sofia | FIA | yes | 15-20 | POC | # | yes | yes | cohort | prospective | USA | clinical | adult | symptomatic, screening independent of symptoms | # | 56 | # | # | # | # | # | 70 | RT-PCR | Cobas; XpertXpress | Cobas, Xpert Xpress | nasal | AN/MT | NP | NP | IFU | 109 | 72 | 42 | 114 | # | 114 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 13 | 0 | 67 | 29 | 31 | NA | NA | NA | NA | NA | # | # | # | HIC | 2021-02-09 | 2021-05-10 | wildtype, Alpha
113 | s10_1_0_0 | s10 | 1 | 0 | 0 | Fourati, S., et al. | VITROS | CLIA | yes | 48 | lab based | 130 | no | yes | cohort | prospective | France | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Hologic, Alinity | Hologic Aptima, Alinity | NP | NP | NP | NP | IFU | 1758 | 1657 | 76 | # | # | 1758 | 25 | # | 15 | 39 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 102 | 0 | 1614 | 42 | NA | NA | NA | 100 | 99.8 | 100 | # | # | # | HIC | 2020-11-01 | 2020-12-31 | wildtype
114 | s11_1_0_0 | s11 | 1 | 0 | 0 | Freeman, M. C., et al. | Sofia | FIA | yes | 15-20 | POC | # | no | yes | cohort | prospective | USA | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | 5.8 | 0 | 16 | # | # | # | 144 | RT-PCR | XpertXpress | Cepheid GeneXpert Xpress | nasal | AN/MT | nasal | AN/MT | IFU | 249 | # | 249 | # | 249 | # | # | 22.2 | 12.5 | 38.4 | # | # | # | # | # | # | # | # | # | # | # | 2.6 | 1 | 5 | # | # | # | # | # | # | # | # | # | # | # | 36 | 2 | 203 | 8 | 81.8 | 67.3 | 91.8 | 99 | 96.5 | 99.9 | # | # | # | HIC | 2020-11-01 | 2021-04-30 | wildtype, Alpha
115 | s11_2_0_0 | s11 | 2 | 0 | 0 | Freeman, M. C., et al. | Sofia | FIA | yes | 15-20 | POC | # | no | yes | cohort | prospective | USA | clinical | children | symptomatic, high risk contact, screening independent of symptoms | # | 7.3 | 0 | 15 | # | # | # | 61 | RT-PCR | XpertXpress | Cepheid GeneXpert Xpress | nasal | AN/MT | nasal | AN/MT | IFU | 138 | 138 | # | # | 138 | # | # | 26.7 | 18.4 | 38.1 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 9 | 1 | 123 | 5 | 64.3 | 35.1 | 87.2 | 99.2 | 95.6 | 100 | # | # | # | HIC | 2020-11-01 | 2021-04-30 | wildtype, Alpha
116 | s13_1_0_0 | s13 | 1 | 0 | 0 | GarcÃ­a-FernÃ¡ndez, S.,et al. | STANDARD F | FIA | yes | 10-30 | POC | # | no | yes | cohort | prospective | Spain | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | VIASURE SARS-CoV-2 | VIASURE SARS-CoV-2 | NP | NP | NP | NP | IFU | 663 | # | 663 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 105 | 2 | 536 | 20 | 84 | 76.4 | 89.9 | 99.6 | 98.7 | 100 | # | # | # | HIC | 2020-10-14 | 2020-11-18 | wildtype
117 | s15_1_0_0 | s15 | 1 | 0 | 0 | Gourgeon, A. | Sofia | FIA | yes | 15-20 | POC | # | yes | yes | case-control | retrospective | France | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | inhouse, TaqPath, RealStar, Hologic, Alinity | inhouse, TaqPath, RealStar, Hologic, Alinity | NP | NP | NP | NP | IFU | 282 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 78 | 1 | 149 | 54 | 59.1 | 50.2 | 67.6 | 99.3 | 96.3 | 100 | # | # | # | HIC | 2020-03-01 | 2022-01-31 | wildtype, Alpha, Delta
118 | s15_2_0_0 | s15 | 2 | 0 | 0 | Gourgeon, A. | STANDARD F | FIA | yes | 10-30 | POC | # | yes | yes | case-control | retrospective | France | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | inhouse, Altona, RealStar,Aptima, Hologic, Alinity | inhouse, Altona, RealStar,Aptima, Hologic, Alinity | NP | NP | NP | NP | IFU | 300 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 71 | 0 | 150 | 79 | 47.3 | 39.1 | 55.6 | 100 | 97.6 | 100 | # | # | # | HIC | 2020-03-01 | 2022-01-31 | wildtype, Alpha, Delta
119 | s15_3_0_0 | s15 | 3 | 0 | 0 | Gourgeon, A. | Wantai (Beijing) Rapid Test (FIA) | FIA | yes | 15-20 | POC | # | yes | yes | case-control | retrospective | France | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | inhouse, Altona, RealStar,Aptima, Hologic, Alinity | inhouse, Altona, RealStar,Aptima, Hologic, Alinity | NP | NP | NP | NP | IFU | 268 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 94 | 3 | 147 | 24 | 79.7 | 71.3 | 86.5 | 98 | 94.3 | 99.6 | # | # | # | HIC | 2020-03-01 | 2022-01-31 | wildtype, Alpha, Delta
120 | s16_1_0_0 | s16 | 1 | 0 | 0 | Greub, G., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | case-control | retrospective | Switzerland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, BD MAX, GenXpert | Cobas, BD MAX, GenXpert | NP | NP | NP | NP | IFU | NA | # | # | # | # | # | 24.5 | # | 9.3 | 38.5 | 23.8 | 25.2 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 99 | 2 | 198 | 1 | 99 | 94.6 | 100 | 99 | NA | NA | # | # | # | HIC | 2020-11-01 | 2021-02-28 | wildtype
121 | s16_2_0_0 | s16 | 2 | 0 | 0 | Greub, G., et al. | Exdia | FIA | yes | 15-20 | POC | # | no | yes | case-control | retrospective | Switzerland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, VIASURE, GenXpert; Allples; inhouse | Cobas, VIASURE, GenXpert; Allples; inhouse | NP | NP | NP | NP | IFU | NA | # | # | # | # | # | 24.5 | # | 9.3 | 38.5 | 23.8 | 25.2 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 81 | 1 | 199 | 19 | 81 | 71.9 | 88.2 | 99.5 | NA | NA | # | # | # | HIC | 2020-11-01 | 2021-02-28 | wildtype
122 | s16_3_0_0 | s16 | 3 | 0 | 0 | Greub, G., et al. | BD Veritor | DIA | yes | # | POC | # | no | yes | case-control | retrospective | Switzerland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, VIASURE, GenXpert | Cobas, VIASURE, GenXpert | NP | NP | NP | NP | IFU | NA | # | # | # | # | # | 24.5 | # | 9.3 | 38.5 | 23.8 | 25.2 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 71 | 1 | 199 | 29 | 71 | 61.1 | 79.6 | 99.5 | 97.2 | 100 | # | # | # | HIC | 2020-11-01 | 2021-02-28 | wildtype
123 | s17_1_0_0 | s17 | 1 | 0 | 0 | Gunell, M., et al. | mariPOC | FIA | yes | 20-55 | POC | 12.5 | no | yes | cohort | prospective | Finland | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | # | 2 | 81 | # | # | # | 575 | RT-PCR | Chemagic 360 extractor | in house PCR | NP | NP | NP | NP | IFU | 939 | 58 | 881 | # | 129 | 810 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 18 | 0 | 916 | 5 | 78.3 | NA | NA | 100 | NA | NA | # | # | # | HIC | 2021-02-01 | 2021-03-30 | Alpha
124 | s17_2_0_0 | s17 | 2 | 0 | 0 | Gunell, M., et al. | mariPOC | FIA | yes | 20-55 | POC | 12.5 | no | yes | cohort | prospective | Finland | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Chemagic 360 extractor | in house PCR | NP | NP | NP | NP | IFU | 23 | 6 | 17 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 17 | NA | NA | 6 | 73.9 | NA | NA | NA | NA | NA | # | # | # | HIC | 2021-02-01 | 2021-03-30 | Alpha
125 | s18_1_0_0 | s18 | 1 | 0 | 0 | Hirotsu, Y., et al., | Lumipulse G | CLEIA | yes | # | lab based | 120 | yes | yes | case-control | prospective | Japan | clinical | adult, children | unclear | # | 49.2 | 1 | 94 | # | # | # | 78 | RT-PCR | StephOnePlus, XpertXpress | StephOnePlus, XpertXpress | NP | NP | NP | NP | IFU | 154 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 84 | 0 | 63 | 7 | 92.3 | 84.8 | 96.9 | 100 | 94.3 | 100 | # | # | # | HIC | unclear | unclear | unclear
126 | s19_1_0_0 | s19 | 1 | 0 | 0 | Horban, A., et al. | PCL | FIA | no | <30 | POC | # | no | yes | case-control | unclear | Poland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Genesig | Genesig | NP | NP | NP | NP | IFU | 266 | # | # | 266 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 16 | 5 | 221 | 24 | 40 | NA | NA | 97.8 | 94.9 | 99.3 | # | # | # | HIC | unclear | unclear | unclear
127 | s20_1_0_0 | s20 | 1 | 0 | 0 | HÃ¶rber, S., et al. | Atellica | CLIA | yes | 26 | lab based | 1800 | no | yes | cohort | prospective | Germany | testing site | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TaqPath | TaqPath | nasal | AN/MT | NP | NP | IFU | 447 | 426 | 21 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 207 | 1 | 212 | 27 | 88.5 | 83.7 | 92.3 | 99.5 | 97.4 | 99.9 | # | # | # | HIC | 2022-01-16 | 2022-02-28 | Omicron
128 | s21_1_0_0 | s21 | 1 | 0 | 0 | Kirby, J. E., et al. | LumiraDx | FIA | yes | 12 | POC | # | yes | yes | case-control | prospective, retrospective | USA | both | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | Virus culture, RT-PCR | Alinity, Abbott | Alinity, Abbott | NP | NP | NP | NP | IFU | 489 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 124 | 0 | 300 | 65 | NA | NA | NA | NA | NA | NA | # | # | # | HIC | 2021-03-01 | 2021-06-30 | wildtype, Alpha, Delta
129 | s22_1_0_0 | s22 | 1 | 0 | 0 | Kliegr, T., et al. | STANDARD F | FIA | yes | 10-30 | POC | # | yes | yes | case-control | retrospective | Czech Republic | testing site | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | unclear | unclear | non-IFU | 3907 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | 79.5 | 75.7 | 98.7 | 98.3 | 99.1 | NA | # | # | # | HIC | 2021-08-05 | 2021-12-06 | Delta
130 | s23_2_0_0 | s23 | 2 | 0 | 0 | FernÃ¡ndez-Rivas, G., et al. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | case-control | retrospective | Spain | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | 44 | # | 1 | 94 | # | # | # | 616 | RT-PCR | Allplex | Allplex | nasal | AN/MT | nasal | AN/MT | IFU | 732 | 861 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 132 | 0 | 570 | 30 | 81.48 | 74.6 | 87.1 | 100 | 99.4 | 100 | # | # | # | HIC | 2021-01-01 | 2021-01-31 | wildtype
131 | s24_1_0_0 | s24 | 1 | 0 | 0 | Kolesova, O., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | cohort | prospective | Italy | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | 48 | # | 10 | 87 | # | # | # | 52 | RT-PCR | MutaPLEX | MutaPLEX | NP | NP | NP | NP | non-IFU | 110 | # | # | 110 | # | # | # | # | 14 | 40 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 64 | 0 | 34 | 12 | 84.2 | 74 | 91.6 | 100 | NA | NA | # | # | # | HIC | unclear | unclear | unclear
132 | s25_1_0_0 | s25 | 1 | 0 | 0 | KrÃ¼ger, L. J., et al. | Bioeasy | FIA | yes | 12 | POC | # | yes | yes | cross-sectional | prospective | Germany | testing site | adult | symptomatic, high risk contact, screening independent of symptoms | 40 | # | # | # | 30 | 54 | # | 369 | RT-PCR | TibMolBiol, Allplex, RealTime, Roche cobas | TibMolBiol, Allplex, RealTime, Roche cobas | NP/OP | NP | NP/OP | NP | IFU | 727 | 272 | 446 | # | 0 | 728 | # | # | # | # | # | # | # | 6.6 | # | # | # | 4.6 | 8.3 | # | 3 | # | # | # | 2 | 6 | # | # | # | # | # | # | # | # | # | 10 | 49 | 663 | 5 | 66.7 | 38.4 | 88.2 | 93.1 | 91 | 94.9 | # | # | # | HIC | 2020-04-17 | 2021-03-31 | wildtype, Alpha
133 | s26_1_0_0 | s26 | 1 | 0 | 0 | Lau, C. S., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | case-control | retrospective | Singapore | unclear | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | Cobas, GeneXpert | Cobas, GeneXpert | NP | NP | NP | NP | IFU | 350 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 26 | 1 | 288 | 35 | 42.6 | 30 | 55.9 | 99.7 | 98.1 | 100 | # | # | # | HIC | 2020-06-01 | 2021-10-01 | Delta
134 | s27_2_0_0 | s27 | 2 | 0 | 0 | Linssen, J., et al. | HISCL | CLEIA | yes | 17 | lab based | 200 | no | yes | cross-sectional | prospective | Netherlands | testing site | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TibMolbiol, Cobas | TibMolbiol, Cobas | NP | NP | NP | NP | IFU | 541 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 11 | 3 | 522 | 5 | 68.8 | 41.3 | 89 | 99.4 | 98.3 | 99.9 | # | # | # | HIC | 2021-04-01 | 2021-09-30 | Alpha, Delta
135 | s27_3_0_0 | s27 | 3 | 0 | 0 | Linssen, J., et al. | HISCL | CLEIA | yes | 17 | lab based | 200 | no | yes | cross-sectional | prospective | Netherlands | testing site | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | TibMolbiol, Cobas | TibMolbiol, Cobas | NP | NP | NP | NP | IFU | 969 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 50 | 153 | 761 | 5 | 90.9 | 80 | 97 | 83.3 | 80.7 | 85.6 | # | # | # | HIC | 2021-04-01 | 2021-09-30 | Alpha, Delta
136 | s28_1_0_0 | s28 | 1 | 0 | 0 | Mehta, A., et al. | Gazelle | FIA | yes | # | POC | # | no | no | cohort | prospective | India, USA | both | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | 34.7 | 2 | 94 | # | # | # | 570 | RT-PCR | TaqPath | TaqPathTM COVID-19 Combo Kit, Taq Path Covid RT-PCR kit | nasal | AN/MT | nasal | AN/MT | IFU | 1486 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 129 | 4 | 1348 | 5 | 96.3 | NA | NA | 99.7 | NA | NA | # | # | # | HIC/LMIC | 2021-08-23 | 2021-11-12 | Delta
137 | s29_1_0_0 | s29 | 1 | 0 | 0 | MiÅ‚osz, M., et al. | Euroimmun | ELISA | yes | # | lab based | # | no | yes | cohort | prospective | Poland | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | EURORealTime | EURORealTime | NP | NP | NP | NP | IFU | 137 | 35 | 96 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 51 | 1 | 85 | 0 | 100 | NA | NA | 98.84 | NA | NA | # | # | # | HIC | 2020-11-17 | 2020-11-20 | wildtype
138 | s3_1_0_0 | s3 | 1 | 0 | 0 | Alghounaim, M., et al. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | cross-sectional | prospective | Kuwait | testing site | adult | unclear | 40.2 | # | # | # | 32.3 | 47.8 | # | 59 | RT-PCR | TaqPath | TaqPath | nasal | AN/MT | nasal | AN/MT | IFU | 897 | # | 30 | # | # | 897 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 26 | 1 | 836 | 34 | 43.4 | 30.6 | 56.8 | 99.9 | 99.3 | 100 | # | # | # | HIC | unclear | unclear | unclear
139 | s31_1_0_0 | s31 | 1 | 0 | 0 | Navero-Castillejos, J., et al. | BD Veritor | DIA | yes | 12 | POC | # | no | yes | case-control | retrospective | Spain | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | XpertXpress, TaqPath | Xpress, TaqPath | NP | NP | NP | NP | IFU | 130 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 62 | 0 | 50 | 18 | 77.5 | 66.8 | 86.1 | 100 | NA | NA | # | # | # | HIC | 2020-11-01 | 2021-01-31 | wildtype
140 | s32_1_0_0 | s32 | 1 | 0 | 0 | Neukam, K., et al. | BD Veritor | DIA | yes | 12 | POC | # | no | yes | cross-sectional | prospective | Argentina | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | 43 | # | # | # | 33 | 55 | # | 358 | RT-PCR | RealStar | RealStar | NP | NP | NP | NP | IFU | 584 | 476 | 108 | # | 0 | 584 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 22 | 1 | 550 | 11 | 66.7 | 48.2 | 82 | 99.8 | 99 | 100 | # | # | # | LMIC | 2021-06-01 | 2021-06-30 | Gamma
141 | s33_1_0_0 | s33 | 1 | 0 | 0 | Nomoto, H., et al. | Lumipulse G | CLEIA | yes | # | lab based | 120 | no | yes | cohort | prospective | Japan | clinical | adult | symptomatic | # | # | # | # | # | # | # | 67 | RT-PCR | Cobas | Cobas | AN | AN/MT | NP | NP | IFU | 128 | 0 | 128 | # | 0 | 128 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2 | # | # | # | 1 | 4 | # | # | # | # | # | # | # | # | # | 29 | 0 | 96 | 3 | 90.6 | 75 | 98 | 100 | 96.2 | 100 | # | # | # | HIC | 2021-01-21 | 2021-08-17 | wildtype, Alpha, Delta
142 | s33_2_0_0 | s33 | 2 | 0 | 0 | Nomoto, H., et al. | Lumipulse G | CLEIA | yes | # | lab based | 120 | no | yes | cohort | prospective | Japan | clinical | adult | symptomatic | # | # | # | # | # | # | # | 67 | RT-PCR | Cobas | Cobas | NP | NP | NP | NP | IFU | 128 | 0 | 128 | # | 0 | 128 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 2 | # | # | # | 1 | 4 | # | # | # | # | # | # | # | # | # | 31 | 0 | 96 | 1 | 96.9 | 83.8 | 99.9 | 100 | 96.2 | 100 | # | # | # | HIC | 2021-01-21 | 2021-08-17 | wildtype, Alpha, Delta
143 | s35_1_0_0 | s35 | 1 | 0 | 0 | Peck Palmer, O., et al. | ADVIA Centaur | CLIA | NA | 23 | lab based | 240 | no | yes | case-control | retrospective | USA | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | 207 | RT-PCR | XpertXpress | XpertXpress | nasal, NP | NP | nasal, NP | NP | IFU | 347 | 56 | 291 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 232 | 2 | 73 | 40 | 85.3 | 80.5 | 89.3 | 97.3 | 90.7 | 99.7 | # | # | # | HIC | 2021-01-01 | 2021-10-24 | wildtype, Alpha, Delta
144 | s36_1_0_0 | s36 | 1 | 0 | 0 | Petonnet, D., et al. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | case-control | prospective | France | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | 57 | # | # | # | 40 | 68 | # | # | RT-PCR | Cobas, Simplexa, BioFire, NeumoDX | Cobas, Simplexa, BioFire, NeumoDX | NP | NP | NP | NP | IFU | 242 | # | # | # | 0 | 242 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 55 | 0 | 155 | 32 | 63.2 | 52.2 | 73.3 | 100 | 97.6 | 100 | # | # | # | HIC | 2020-12-01 | 2021-01-18 | wildtype
145 | s37_1_0_0 | s37 | 1 | 0 | 0 | Robinson, M. L. | BD Veritor | DIA | yes | 12 | POC | # | no | yes | cohort | prospective | USA | testing site | adult | symptomatic, high risk contact, screening independent of symptoms | 38 | # | # | # | 29 | 50.5 | # | 112 | RT-PCR | RealStar, Cobas; PerkinElmer, NeuMoDx, Aptima | RealStar, Cobas; PerkinElmer, NeuMoDx, Aptima | AN/MT | AN/MT | NP | NP | IFU | 99 | 55 | 44 | # | 0 | 99 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | # | # | # | HIC | 2020-11-17 | 2021-10-01 | wildtype
146 | s38_1_0_0 | s38 | 1 | 0 | 0 | Routsias, J. G., et al. | Wantai (Beijing) Rapid Test (FIA) | FIA | yes | 20 | POC | # | yes | yes | cross-sectional | prospective | Greece | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | LiliF | LiliF | nasal | AN/MT | nasal | AN/MT | IFU | 450 | # | # | 0 | 0 | 450 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 383 | 0 | 50 | 17 | 95.71 | NA | NA | 100 | NA | NA | cT = 35.3 | # | # | HIC | unclear | unclear | unclear
147 | s39_1_0_0 | s39 | 1 | 0 | 0 | Saito, K., et al. | HISCL | CLEIA | yes | 17 | lab based | 200 | no | yes | case-control | prospective | Japan | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | QIAamp | QIAamp Viral RNA Mini Kit | NP | NP | NP | NP | IFU | 115 | # | # | # | # | # | 22.115 | 24.3912 | 17.78 | 36.22 | 19.875 | 27.45 | 8.3587 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | COI | 398.8 | 4191.46 | 2.6 | 38233.6 | 22.8 | 3849.3 | 150.7322 | 26 | 0 | 69 | 20 | NA | NA | NA | NA | NA | NA | # | # | # | HIC | unclear | unclear | unclear
148 | s4_1_0_0 | s4 | 1 | 0 | 0 | Bornemann, L., et al. | Sofia | FIA | yes | 15-20 | POC | # | no | yes | cohort | prospective | Germany | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | 67 | # | 1 | 107 | # | # | # | 4237 | RT-PCR | Viasure; Rida; XpertXperss; LightMix; Allplex | Viasure; Rida; XpertXprss; LightMix; Allplex | NP | NP | NP | NP | IFU | 7863 | # | # | # | 127 | 7750 | 25 | 25.295 | 11 | 41 | # | # | 6.7604 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 236 | 48 | 7440 | 139 | 62.9 | 57.8 | 67.8 | 99.4 | 99.2 | 99.5 | # | # | # | HIC | 2020-11-01 | 2021-03-31 | wildtype, Alpha
149 | s40_1_0_0 | s40 | 1 | 0 | 0 | Salvagno, G. L., et al. | Elecsys | ECLIA | yes | # | lab based | 300 | no | yes | cohort | prospective | Italy | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | 38 | # | # | # | 24 | 58 | # | 69 | RT-PCR | RealStar | Realstar | NP | NP | NP | NP | non-IFU | 160 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 47 | 0 | 50 | 63 | 42.7 | 33.3 | 52.5 | 100 | 93 | 100 | # | # | # | HIC | 2021-08-16 | 2021-09-15 | Delta
150 | s41_1_0_0 | s41 | 1 | 0 | 0 | Salvetti, S., et al. | Lumipulse G | CLEIA | yes | # | lab based | 120 | no | yes | cohort | prospective | Italy | clinical | unclear | screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | ELITe MGB, XpertXpress | ELITe MGB, XpertXpress | NP | NP | NP | NP | non-IFU | 2659 | 2023 | 636 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 106 | 176 | 2367 | 10 | 91.4 | 84.7 | 95.8 | 93.1 | 92 | 94 | # | # | # | HIC | 2021-04-26 | 2021-06-22 | Alpha
151 | s44_1_0_0 | s44 | 1 | 0 | 0 | Sitoe, N., et al. | LumiraDx | FIA | yes | 12 | POC | # | yes | yes | cross-sectional | prospective | Mozambique | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | 36.23 | # | # | # | # | # | # | 381 | RT-PCR | Abbott m200 | Abbott | nasal | AN/MT | OP | OP | IFU | 696 | 121 | 575 | # | 0 | 696 | 25.1 | # | 3.7 | 35.9 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 301 | 13 | 297 | 85 | 78 | 73.5 | 82 | 95.8 | 92.9 | 97.7 | # | # | # | LMIC | 2021-04-01 | 2022-01-31 | Alpha, Delta
152 | s48_1_0_0 | s48 | 1 | 0 | 0 | Uster, S. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | case-control | prospective | Switzerland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, Liat rRT-PCR | Cobas, Liat rRT-PCR | NP | NP | NP | NP | IFU | 208 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 122 | 0 | 47 | 39 | 75.8 | 68.4 | 82.2 | 100 | 92.5 | 100 | # | # | # | HIC | 2020-10-04 | 2021-02-17 | wildtype, Alpha
153 | s48_2_0_0 | s48 | 2 | 0 | 0 | Uster, S. | LIAISON | CLIA | yes | 45 | lab based | # | no | yes | case-control | prospective | Switzerland | clinical | unclear | unclear | # | # | # | # | # | # | # | # | RT-PCR | Cobas, Liat rRT-PCR | Cobas, Liat rRT-PCR | NP | NP | NP | NP | IFU | 137 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 91 | 0 | 18 | 28 | 76.5 | 67.8 | 83.8 | 100 | 81.5 | 100 | # | # | # | HIC | 2022-01-23 | 2022-03-02 | Omicron
154 | s49_1_0_0 | s49 | 1 | 0 | 0 | Wells, C. R., et al. | BD Veritor | DIA | yes | 12 | POC | # | no | yes | cohort | prospective | USA | testing site | adult | screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | unclear | unclear | unclear | unclear | non-IFU | 409 | # | # | # | # | 409 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | NA | # | # | # | # | # | # | # | 3 | 1 | 396 | 9 | NA | NA | NA | NA | NA | NA | # | # | # | HIC | 2020-11-22 | 2021-01-17 | wildtype
155 | s51_1_0_0 | s51 | 1 | 0 | 0 | Yu, K., et al. | STANDARD F | FIA | yes | 10-30 | POC | # | no | yes | case-control | prospective | Korea | clinical | unclear | symptomatic, high risk contact, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | PowerChek | PowerChek SARS-CoV-2 and Influenza A&B Multiplex Real-Time PCR Kit | NP | NP | NP | NP | non-IFU | 554 | 237 | 317 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 147 | 36 | 299 | 72 | 67.1 | 60.5 | 73.3 | NA | NA | NA | # | # | # | HIC | 2021-07-01 | 2021-10-10 | Delta
156 | s6_1_0_0 | s6 | 1 | 0 | 0 | Cattelan, A. M., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Italy | clinical | adult, children | symptomatic, high risk contact, screening independent of symptoms | # | 61 | # | # | 43 | 79 | # | 105 | RT-PCR | One Step Real Time kit | One Step Real Time kit | nasal | AN/MT | NP | NP | IFU | 282 | 32 | 250 | # | # | # | # | 26 | # | # | 20.3 | 31.7 | 5,7 | # | # | # | # | # | # | # | # | 9 | # | # | 2.5 | 15.5 | # | # | # | # | # | # | # | # | # | 174 | 3 | 51 | 54 | 76.3 | 70.3 | 81.7 | 94.4 | 84.6 | 98.8 | # | # | # | HIC | 2021-01-01 | 2021-05-31 | wildtype, Alpha
157 | s7_1_0_0 | s7 | 1 | 0 | 0 | ÄŒernila, M., et al. | Sofia | FIA | yes | 15-20 | POC | # | no | yes | cohort | prospective | Slowenia | clinical | unclear | symptomatic | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | NP | NP | NP | NP | non-IFU | 132 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 9 | 3 | 119 | 1 | 90 | 55.5 | 99.75 | 97.52 | 92.93 | 99.49 | # | # | # | HIC | unclear | unclear | unclear
158 | s7_2_0_0 | s7 | 2 | 0 | 0 | ÄŒernila, M., et al. | Sofia | FIA | yes | 15-20 | POC | # | no | yes | cohort | prospective | Slowenia | clinical | unclear | symptomatic, screening independent of symptoms | # | # | # | # | # | # | # | # | RT-PCR | unclear | unclear | NP | NP | NP | NP | non-IFU | 672 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 524 | 6 | 113 | 29 | 94.8 | 92.6 | 96.5 | 95 | 89.3 | 98.1 | # | # | # | HIC | unclear | unclear | unclear
159 | s9_1_0_0 | s9 | 1 | 0 | 0 | Elli, S., et al. | LumiraDx | FIA | yes | 12 | POC | # | no | yes | cohort | prospective | Italy | clinical | adult | symptomatic, high risk contact, screening independent of symptoms | 72 | # | # | # | # | # | # | 118 | RT-PCR | GeneFinder, Xpert Xpress | GeneFinder, Xpert Xpress | nasal | AN/MT | nasal | AN/MT | IFU | 234 | 174 | # | 234 | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | # | 13 | 15 | 181 | 25 | 34.2 | 19.6 | 51.4 | 92.3 | 87.7 | 95.7 | # | # | # | HIC | 2020-11-11 | 2020-12-08 | wildtype
Note on AE1: seda: leave only the columns used in the analysis

## List of parameters extracted
List of parameters extracted from studies.
Clinical study data |  | CT values across TP and FN |  | Subgroup analysis by symptom status and duration |  | Subgroup analysis by ct-value |  | Subgroup analysis by age |  | Subgroup analysis by viral load
General | Test assessed | General | Author | General | Author | General | Author | General | Author |  | Author
| Technology |  | Test assessed |  | Test assessed |  | Test assessed |  | Test assessed | General | Test assessed
| Connectivity |  | TP vs FN |  | Study location |  | Study location |  | Study location |  | Study location
| Tournaround Time |  | Number of Participants | Accuracy - asymptomatic | Number of Participants | Accuracy - Ct Value | Number of Participants | Accuracy - children (<18 years) | Number of Participants | Accuracy - viral load >5 log10 copies/ml | Number of Participants
| Thoughput | Ct values | CT value (median) |  | True positives |  | Ct value |  | True positives |  | True positives
| Expected location |  | CT value (mean) |  | False positives |  | ct_label |  | False positives |  | False positives
| Head to Head? |  | CT value (lower range) |  | True negatives |  | True positives |  | True negatives |  | True negatives
| Peer reviewed? |  | CT value (upper range) |  | False negatives |  | False positives |  | False negatives |  | False negatives
| Type of study |  | CT value (IQR 1) |  | Sensitivity |  | True negatives |  | Sensitivity |  | Sensitivity
| Retro/Prosp? |  | CT value (IQR 3) |  | Lower 95 CI sensitivity |  | False negatives |  | Lower 95 CI sensitivity |  | Lower 95 CI sensitivity
| Study location |  | CT value (STDev) |  | Upper 95 CI sensitivity |  | Sensitivity |  | Upper 95 CI sensitivity |  | Upper 95 CI sensitivity
| Type of location |  |  |  | Specificity |  | Lower 95 CI sensitivity |  | Specificity |  | Specificity
| Study Population |  |  |  | Lower 95 CI specificity |  | Upper 95 CI sensitivity |  | Lower 95 CI specificity |  | Lower 95 CI specificity
| Screening criteria |  |  |  | Upper 95 CI specificity |  | Specificity |  | Upper 95 CI specificity |  | Upper 95 CI specificity
| Age (Median) |  |  |  | Viral load median |  | Lower 95 CI specificity |  | Comment |  | Comment
| Age (mean) |  |  |  | Viral load mean |  | Upper 95 CI specificity | Accuracy - adults (> or = 18 years) | Number of Participants | Accuracy - viral load inclduing >6 log10 copies/ml | Number of Participants
| Age (lower Range) |  |  |  | CT value (median) |  |  |  | True positives |  | True positives
| Age (upper Range) |  |  |  | CT value (mean) |  |  |  | False positives |  | False positives
| Age (IQR 1) |  |  | Accuracy - symptomatic | Number of Participants |  |  |  | True negatives |  | True negatives
| Age (IQR 3) |  |  |  | True positives |  |  |  | False negatives |  | False negatives
| Age (STDev) |  |  |  | False positives |  |  |  | Sensitivity |  | Sensitivity
| Gender (Number of Female) |  |  |  | True negatives |  |  |  | Lower 95 CI sensitivity |  | Lower 95 CI sensitivity
| Reference Standard |  |  |  | False negatives |  |  |  | Upper 95 CI sensitivity |  | Upper 95 CI sensitivity
| Type of RT-PCR |  |  |  | Sensitivity |  |  |  | Specificity |  | Specificity
| Type of RT-PCR_notes |  |  |  | Lower 95 CI sensitivity |  |  |  | Lower 95 CI specificity |  | Lower 95 CI specificity
| Type of sample (AG) |  |  |  | Upper 95 CI sensitivity |  |  |  | Upper 95 CI specificity |  | Upper 95 CI specificity
| sample_adjusted |  |  |  | Specificity | [5 empty] | Comment
| Type of sample (PCR) |  |  |  | Lower 95 CI specificity |  |  |  |  | Accuracy - viral load including >7 log copies/ml | Number of Participants
| Testing procedure |  |  |  | Upper 95 CI specificity | [5 empty] | True positives
| Why non-IFU* |  |  |  | Viral load median | [5 empty] | False positives
| Why non-IFU_notes* |  |  |  | Viral load mean | [5 empty] | True negatives
| Mentioned if followed IFU |  |  |  | CT value (median) | [5 empty] | False negatives
| Transport medium |  |  |  | CT value (mean) | [5 empty] | Sensitivity
| Sample condition |  |  | Accuracy - Duration of Symptoms <= 7 days | Number of Participants | [5 empty] | Lower 95 CI sensitivity
Clinical data | Population size |  |  |  | True positives | [5 empty] | Upper 95 CI sensitivity
| # of asymptomatic |  |  |  | False positives | [5 empty] | Specificity
| # of symptomatic |  |  |  | True negatives | [5 empty] | Lower 95 CI specificity
| # of hospitalized |  |  |  | False negatives | [5 empty] | Upper 95 CI specificity
| # of children (<18 years) |  |  |  | Sensitivity | [5 empty] | Comment
| # of adults (> or = 18 years) |  |  |  | Lower 95 CI sensitivity
| CT value (median) |  |  |  | Upper 95 CI sensitivity
| CT value (mean) |  |  |  | Specificity
| CT value (lower range) |  |  |  | Lower 95 CI specificity
| CT value (upper range) |  |  |  | Upper 95 CI specificity
| CT value (IQR 1) |  |  | Accuracy - Duration of Symptoms > 7 days | Number of Participants
| CT value (IQR 3) |  |  |  | True positives
| CT Value (STDev) |  |  |  | False positives
| Viral load (median) |  |  |  | True negatives
| Viral load (mean) |  |  |  | False negatives
| Viral load (lower Range) |  |  |  | Sensitivity
| Viral load (upper Range) |  |  |  | Lower 95 CI sensitivity
| Viral load (IQR 1) |  |  |  | Upper 95 CI sensitivity
| Viral load (IQR 3) |  |  |  | Specificity
| Viral load (STDev) |  |  |  | Lower 95 CI specificity
| Duration of symptoms (median) |  |  |  | Upper 95 CI specificity
Duration of symptoms (mean)
Duration of symptoms (lower range)
Duration of symptoms (upper range)
Duration of symptoms (IQR 1)
Duration of symptoms (IQR 3)
Duration of symptoms (STDev)
Antigen levels (quantity vs COI)
Antigen levels (Median)
Antigen levels (Mean)
Antigen levels (lower Range)
Antigen levels (upper Range)
Antigen levels (IQR 1)
Antigen levels (IQR 3)
Antigen levels (STDev)
Accuracy | True positives
False positives
True negatives
False negatives
Sensitivity
Lower 95 CI sensitivity
Upper 95 CI sensitivity
Specificity
Lower 95 CI specificity
Analytical | Upper 95 CI specificity
Comment
LOD by Ct value
LOD by viral load
viral load UNIT
test-specific values (eg COI for Standard F)
Comment

## Investigated tests
Expected location | Test | Technology | Sensitivity from MA | Specificity from MA | Sensitivity independent from MA | Specificity independnent from MA | Through-put [tests/hour] | Turnaround time | Numbre of paper | Numbre of samples
POC | BD Veritor | Colorimetric Immunochromatography assay | 84% | 100% | 73.9% | 99.4% |  | 15 | 17 | 11578
| RapidTesta | Colorimetric Immunochromatography assay | NP 68,6% Nasal 57,1% | NP 68,6% Nasal 57,1% |  |  |  | 10 | 1 | 1127
| mariPOC | Fluorescence immunoassay | 92.3% | 100% |  |  | 12.5 | 20 | 2 | 1020
| Exdia | Fluorescence immunochromatographicassay(FICA) | 97.1% | 99.6% |  |  |  | 15-20 | 2 | 532
| Bioeasy | Fluorescence immunochromatographicassay(FICA) | 91.71% | 100% |  |  |  | 10 | 5 | 1864
| PCL COVID-19 Ag | Fluorescence immunochromatographicassay(FICA) | 100% | 97.78% |  |  |  | 10 | 2 | 353
| Standard F | Fluorescence immunochromatographicassay(FICA) | N/A | N/A | 65.5% | 98.1% |  | 30 | 17 | 19617
| Sofia | Fluorescence immunochromatographicassay(FICA) | 96.7% | 100% | 75,3%% | 99,1%% |  | 15 | 18 | 20970
| Gazelle COVID-19 | Fluorescence immunochromatographicassay(FICA) | N/A | N/A |  |  |  | N/A | 1 | 1486
| WANTAI SARS-CoV-2 Antigen Rapid Test (FIA) | Fluorescence immunochromatographicassay(FICA) | 95.5% | 99% |  |  |  | N/A | 2 | 718
| ACON FIAflex | Fluorescence immunochromatographicassay(FICA) | N/A | N/A |  |  |  | 15 | 1 | 347
| AFIAS | Fluorescence immunochromatographicassay(FICA) | N/A | N/A |  |  |  | 12 | 2 | 846
| ichroma | Fluorescence immunochromatographicassay(FICA) | 90.9% | 98,70% |  |  |  | 12 | 1 | 232
| FREND | Microfluidic fluorescence immunoassay | 94.12% | 100% |  |  |  | <3 | 2 | 170
| LumiraDx | Microfluidic fluorescence immunoassay | 97.69% | 96% | 81.1% | 97.3% | 3.5 | 12 | 22 | 9836
lab based | S-PLEX, MesoScale | Electrochemiluminescence immunoassays (ECLIA) | N/A | N/A |  |  | 120 | <20 | 3 | 846
| Elecsys | Electrochemiluminescence immunoassay (ECLIA) | N/A | N/A | 65.6% | 98.3% | 300 | 15-18 | 9 | 5586
| HISCL SARS-COV-2 AG ASSAY KIT | Chemiluminescent enzyme immunoassay (CLEIA) | 100% | 100% |  |  | 200 | 17 | 2 | 1625
| SARS-CoV-2 Antigen (CoV2Ag) Assay on an Atellica Solution | Chemiluminescent enzyme immunoassay (CLEIA) | 100% | 84.97% |  |  | 200 | 26 | 1 | 447
| SARS-CoV-2 Antigen (CoV2Ag) Assay on an ADVIA Centaur | Chemiluminescent enzyme immunoassay (CLEIA) | N/A | N/A |  |  | 240 | 18 | 1 | 347
| Lumipulse G | Chemiluminescent enzyme immunoassay (CLEIA) | 91.7% | 98.5% | 86,5%% | 96.4% | 120 | <35 | 13 | 8953
| Liaison | Chemiluminescent Immunoassay | 93,1% NS 95,4% NPS | 100%NS 99,7%NPS | 62,5%% | 99,6%% | 136 | 17 | 9 | 4087
| VITROS | Chemiluminescent Immunoassay | 75.4% | 100% | 76.8% | 99.7% | 130 | 15 | 5 | 2774
| Euroimmun | Enzyme-linked immunosorbent assay (ELISA) | 95.5% | 99.2% |  |  |  | 15-30 | 2 | 547