Table S1 The cone and collision energy for the detection of α-AAA, PA, α-AASA, P6C, D3-α-AAA, D9-PA and 6-oxo-PIP.

| Instrument model |  |  | Waters XEVO TQD |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Detection mode |  |  | positive |  |  |  |
| Source temperature |  |  | 150℃ |  |  |  |
| Capillary voltage |  |  | 1KV |  |  |  |
| Desolvation Temp |  |  | 250℃ |  |  |  |
| Desolvation gas flow rate |  |  | 800L/hr |  |  |  |
| Cone gas flow rate |  |  | 50 L/hr |  |  |  |
| name | Q1(m/z) | Q3(m/z) |  | Dwell (s) | Cone(V) | CE(V) |
| α-AAA | 274.28 | 98.05 |  | 0.035 | 30 | 22 |
|  |  | 172.17* |  | 0.035 | 30 | 15 |
| PA | 186.17 | 57.04 |  | 0.035 | 35 | 20 |
|  |  | 83.95* |  | 0.035 | 35 | 18 |
| α-AASA | 202.2 | 81.96* |  | 0.035 | 20 | 18 |
|  |  | 128.29 |  | 0.035 | 20 | 15 |
| P6C | 184.25 | 81.97* |  | 0.035 | 35 | 18 |
|  |  | 128.03 |  | 0.035 | 35 | 12 |
| D3-AAA | 277.31 | 101.1 |  | 0.035 | 30 | 22 |
|  |  | 175.1* |  | 0.035 | 30 | 15 |
| D9-PA | 195.17 | 93.06* |  | 0.035 | 30 | 18 |
|  |  | 139.14 |  | 0.035 | 30 | 15 |
| 6-oxo-PIP | 200.11 | 98.01* |  | 0.15 | 25 | 20 |
|  |  | 144.05 |  | 0.15 | 25 | 12 |

In table“*”label represent the ions for quantification

Table S2 The imprecision for the detection of α-AAA, PA, α-AASA, P6C, D3-α-AAA, D9-PA and 6-oxo-PIP in plamsa, DBS, and urine

Plasma：

|  | α-AAA | PA | AASA+P6C | 6-oxo-PIP |
| --- | --- | --- | --- | --- |
| inter-assay | 2.5%-3.6% | 1.9%-8.8% | 1.8%-5.0% | 2.4%-4.1% |
| intra-assay | 3.6%-4.3% | 5.7%-10.8% | 5.9%-11.0% | 4.7%-8.9% |

DBS：

|  | α-AAA | PA | AASA+P6C | 6-oxo-PIP |
| --- | --- | --- | --- | --- |
| inter-assay | 2.2%-3.1% | 2.3%-3.5% | 2.5%-3.7% | 2.3%-4.5% |
| intra-assay | 2.2%-4.2% | 3.0%-5.8% | 2.1%-4.1% | 6.4%-6.7% |

Urine：

|  | α-AAA | PA | AASA+P6C | 6-oxo-PIP |
| --- | --- | --- | --- | --- |
| inter-assay | 0.9%-1.8% | 0.8%-2.4% | 0.7%-1.6% | 1.8%-5.4% |
| intra-assay | 1.5%-1.8% | 3.2%-4.6% | 2.1%-2.8% | 5.8%-7.3% |

Fig.S1. The classic extract ion chromatogram of α-AASA, P6C, PA, α-AAA, d9-PA, d3-AAA, and 6-oxo-PIP in quality controls (A) and PDE-ALDH7A1 patient (B).

Fig.S2. The correlation between the concentration of α-AASA-P6C, PA, α-AAA, 6-oxo-PIP and age (A), pyridoxine dose (B), duration of treatment (C), respectively. r and p from linear regression is shown. The r and P values are on the right side of each graph, and from top to bottom are plasma, DBS, urine, and DUS.
Fig.S3. The difference of α-AASA-P6C, PA, α-AAA, 6-oxo-PIP levels among different developmental levels. ns p>0.05; *p < 0.05.