Supplementary Materials

Table of Contents

Tables	2

Table S1. Characteristics of Included Studies	2

Table S2. Pulse Oximetry Study Bias by Skin Pigmentation	8

Table S3. Pulse Rate Study Bias by Skin Pigmentation	13

Table S4. Heterogeneity Pulse Rate Bias by Skin Tone	15

Table S5. Sensitivity Analyses of Pulse Rate Bias by Skin Pigmentation	16

Table S6. Pulse Oximetry Sensitivity Analysis with Crooks et al. (2022) Removed	17

Table S7. Shi et al. (2022)28 Manuscripts Not Included	18

Table S8. Coding Skin Tone, Race, and Ethnicity to Low, Medium, and High Skin Pigmentation Categories	19

Table S9. Excluded Manuscripts	20

Figures	27

Figure S1. Funnel Plot Pulse Oximetry	27

Figure S2. Funnel Plot Pulse Rate	28

Materials	29

Search Report: PPG and Skin of Color	29

Data Extraction	34

Additional Details on Method for Statistical Analysis	35

# Tables

## Table S1. Characteristics of Included Studies

| Study (Author, Year) | Sample Size | Test Device Evaluated | Reference Device | SaO2 (%) Range | Participant Population | Research Setting | Skin Pigmentation Method |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Pulse Oximetry |  |  |  |  |  |  |  |
| Abrams 20031 | 200 | Nellcor N-200 | Radiometer ABL-520 | Not reported | Medical- Adult patients with cirrhosis | Inpatient | Race- Black and White |
| Adler 19982 | 298 | Nellcor D-25 | 4-wavelength spectro-photometer, or co-oximeter (Radiometer OSM3) | 50 to 99 | Medical- Adult emergency room patients | Inpatient | Skin Tone- Munsell color system categorized into light, medium, or dark |
| Andrist 20223 | 1,061 | Not Reported | Not Reported | Not Reported | Medical- Pediatric | Inpatient | Race |
| Barker 20234 | 75 | Masimo SET® pulse oximeters with RD-SET® sensors | Radiometer ABL-835 Flex CO-Oximeter | 70 to 100 | Mixed- Healthy and mild systemic disease | Research- Lab | Race |
| Bickler 20055 | 23 | -Nellcor N-595 with Nellcor OxiMax A finger probe -Novametrix 513s models (two types) -Nonin Onyx models (two types) | Radiometer OSM3 | 60 to 100 | Healthy- non-smoking | Research- Lab | Ethnicity- light (Northern European) and dark (African-American) |
| Bothma 19966 | 100 | Simed S100e Nihon Koden Ohmeda 3740 | IL482 Co-oximeter System | 87·8 to 99·2 | Medical- critically ill adult patients | Inpatient | Skin Tone- EEL reflectance spectrophotometer. All participants had dark skin tone |
| Burnett 20227 | 46,253 | Unspecified Nellcor and Masimo devices | GEMStat Premier 3000 | Not Reported | Medical- patients receiving anesthesia | Inpatient | Race and Ethnicity |
| Crooks 20228 | 2,997 | Not reported | Not reported | Not reported | Medical- COVID-19 | Inpatient | Race |
| Ebmeier 20189 | 394 | Marquette Rac-4A monitors with Masimo sensors and Philips IntelliVue MP70 monitors with Philips Adult Reusable SpO2 sensors | Radiometer ABL 800 FLEX arterial blood gas analyser | Not reported | Medical- multiple conditions | Inpatient | Ethnicity |
| Fawzy 202210 | 1,216 | Not reported | ABL825, ABL827, or ABL90 blood gas analyzers | Not reported | Medical- COVID-19 | Inpatient | Race and Ethnicity |
| Feiner 200711 | 36 | Nellcor N-595 (OxiMax A adhesive probe) Nellcor N-595 (a clip-type probe) Masimo Radical (clip probe) Masimo Radical (adhesive disposable probe) Nonin 9700 (clip-type probe) Nonin 9700 (disposable adhesive probe) | Radiometer OSM3 multiwavelength oximeter | 60 to 100 | Healthy- non-smoking | Research- Lab | Ethnicity- light (Caucasian), intermediate (Hispanic, Indian, Filipino, Vietnamese), and dark (African American) categories |
| Foglia 201712 | 36 | Nellcor Oximax (Covidien) Masimo Rainbow SET Radical 7 | Siemens Rapidlab 1265 | 60 to 92 | Medical- Infants with cyanotic congenital heart disease and oxygen saturation <90% | Inpatient | Skin Tone- Munsell Color System |
| Jubran 199013 | 54 | Nellcor pulse oximeter with disposable or reusable probes Ohmeda-Biox3700 pulse oximeter with reusable probe | CO-oximetry | Not reported | Medical- Critically ill, ventilator-dependent patients | Inpatient | Race- Black and White |
| McGovern 199614 | 8 | Ohmeda 3700 | IL 482 Co-oximeter | Not reported | Medical- Adults with stable condition with severe COPD | Research- Lab | Race- All White |
| Muñoz 200815 | 846 | Minolta Pulsox-7 | IL 682 co-oximeter | Not reported | Medical- Adults under assessment for long-term home oxygen therapy | Outpatient | Race- All White |
| Pilcher 202016 | 400 | Carescape B450 monitor with Nellcor probe GE Dash 3000 Masimo Radical 7 Masimo SET Quartz (unspecified) Masimo SET Quartz Q400 Nonin 2120 Nonin 2140 Nonin Avant (unspecified) Nonin Avant 4000 Nonin Avant 9700 Nonin Lifesense Medair Novametrix Model 512 Ohmeda Biox 3700E with a GE TruSignal or Nellcor probe Philips Intellivue MP70 with a GE TruSignal Nellcor or Philips probe Welch Allyn with a Nellcor probe | Radiometer ABL800 | 72 to 100 | Medical- Hospitalized adult patients | Inpatient and Outpatient | Skin Tone- Fitzpatrick Scale categorized into light, medium, dark |
| Ruppel 202317 | 774 | Not reported | Not reported | Not reported | Medical- cardiac catheterization | Inpatient | Race |
| Sudat 202318 | 8,735 | Not reported | Not reported | Not reported | Medical- multiple conditions | Inpatient | Race |
| Thrush 199419 | 25 | Critikon Dinamap Plus Model 8700 Critikon Oxyshuttle Ohmeda 3700 Catalyst Research MiniOx IV | IL482 co-oximeter | 80 to 100 | Healthy- non-smoking adults | Research- Lab | Race- All White |
| Valbuena 202220 | 372 | Not reported | Not reported, blood gas analysis | Not reported | Medical- Adult patients with respiratory failure or COVID-19 | Inpatient | Race and Ethnicity- White, Black, Hispanic, and Asian |
| Vesoulis 202121 | 294 | Nellcor SpO2 module with Neonatal-Adult MAX-N adhesive SpO2 sensor (Covidien) (used with either Philips IntelliVue MP70 or MX800 monitors) | Radiometer ABL800 Flex | Not reported | Medical- Preterm infants at neonatal intensive care unit | Inpatient | Race- White and Black |
| Wiles 202222 | 194 | Nellcor reusable SpO2 probes or Mindray disposable SpO2 probes (GE Healthcare B1x5 M/P monitor) | RAPIDpoint 500 analyser (Siemens Healthcare GmbH) | Not reported | Medical- Adult patients with COVID-19 pneumonitis | Inpatient | Race- Asian, Black, White, and Other |
| Zeballos 199123 | 33 | Hewlett-Packard HP-47201A Ohmeda Biox IIA | IL282 co-oximeter | Not reported | Healthy- non-smoking volunteers | Research- Lab | Race- All Black |
| Pulse Rate |  |  |  |  |  |  |  |
| Bent et al. (2020)24 | 53 | Empatica E4 Apple Watch Fitbit Charge Garmin Vivosmart 3 Xiaomi Miband Biovotion Everion | BittiumFaros 180, Bittium Inc. | NA | Healthy | Research- Lab | Skin Tone- Fitzpatrick Scale 1 (n = 7) 2 (n = 8) 3 (n = 10) 4 (n = 9) 5 (n = 9) 6 (n = 10) |
| Nelson et al. (2019)25 | 1 | Apple Watch 3 Fitbit Charge 2 | Vrije Universiteit Ambulatory Monitoring System | NA | Healthy | Research- Real World | Skin Tone- Fitzpatrick Scale 1 (n = 1) |
| Sanudo 201926 | 45 | Apple Watch (version not reported) | Polar Chest Strap | NA | Healthy | Research- Lab | Skin Tone- Fitzpatrick Scale 2 (n = 15) 3 (n = 15) 4 (n = 15) |
| Chow 202027 | 40 | Garmin Vivosmart HR+ Xiaomi Mi Band 2 | Polar H7 Chest Strap | NA | Healthy | Research- Lab | Ethnicity and Skin Tone- East Asian, Fitzpatrick Scale 3 and 4 |

## Table S2. Pulse Oximetry Study Bias by Skin Pigmentation

| Study | Device | Sample | Data Pairs | Mean Bias | SE | Skin Pigmentation Group |
| --- | --- | --- | --- | --- | --- | --- |
| Adler et al. 1998 | Nellcor D-25 | 145 | 145 | 2·50 | 0·38 | light |
| Andrist et al. 2022 | Not Specified | 878 | 7018 | 3·50 | 0·17 | light |
| Barker et al. 2023 | Masimo SET¬Æ pulse oximeters | 36 | 3982 | -0·05 | 0·23 | light |
| Bickler et al. 2005 | Nonin Onyx | 10 | 200 | -0·21 | 0·47 | light |
| Bickler et al. 2005 | Novametrix 513 | 10 | 200 | 1·15 | 0·60 | light |
| Bickler et al. 2005 | Nellcor D-25 | 10 | 220 | -0·27 | 0·68 | light |
| Burnett et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 22089 | 70722 | 0·20 | 0·04 | light |
| Crooks et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 3946 | 3946 | 3·20 | 13·24 | light |
| Ebmeier et al. 2018 | Philips (IntelliVue MP70) | 98 | 98 | -0·17 | 0·22 | light |
| Ebmeier et al. 2018 | Masimo (GE Marquette Rac-4A) | 142 | 142 | 1·07 | 0·16 | light |
| Fawzy et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 460 | 460 | 0·75 | 0·09 | light |
| Feiner et al. 2007 | Masimo Radical (Clip) | 12 | 263 | 0·21 | 0·52 | light |
| Feiner et al. 2007 | Nellcor N-595 (clip) | 12 | 263 | 0·36 | 0·50 | light |
| Feiner et al. 2007 | Nonin 9700 (clip) | 12 | 263 | -0·95 | 0·45 | light |
| Feiner et al. 2007 | Masimo Radical (adhesive) | 12 | 263 | -0·64 | 0·68 | light |
| Feiner et al. 2007 | Nellcor N-595 & OxiMax (adhesive) | 12 | 263 | 1·04 | 0·47 | light |
| Feiner et al. 2007 | Nonin 9700 (adhesive) | 12 | 263 | 0·52 | 0·41 | light |
| Foglia et al. 2017 | Masimo Radical (Rainbow) | 21 | 21 | 0·20 | 0·83 | light |
| Foglia et al. 2017 | Nellcor OxiMax | 19 | 19 | 3·00 | 1·15 | light |
| Jubran et al. 1990 | Nellcor/Ohmeda Biox 3700 | 25 | 55 | 2·20 | 0·36 | light |
| McGovern et al. 1996 | IL 482 Co-oximeter | 8 | 8 | -1·7 | 1·03 | light |
| Munoz et al. 2008 | Minolta pulse oximeter | 846 | 846 | 0·47 | 0·08 | light |
| Pilcher et al. 2020 | 14 models | 242 | 242 | -1·2 | 0·11 | light |
| Ruppel et al. 2023 | Not Specified- Hospital Standard Pulse Oximeter | 573 | 573 | 0·88 | 0·13 | light |
| Sudat et al. 2023 | Not Specified- Hospital Standard Pulse Oximeter | 10514 | 35127 | 1·53 | 0·05 | light |
| Thrush et al. 1994 | Dinamap Plus 8700 | 22 | 88 | 0·00 | 0·65 | light |
| Thrush et al. 1994 | Dinamap Plus 8700 | 22 | 88 | 1·00 | 0·43 | light |
| Thrush et al. 1994 | Dinamap Plus 8700 | 22 | 88 | 1·00 | 0·43 | light |
| Thrush et al. 1994 | Dinamap Plus 8700 | 22 | 88 | 1·00 | 0·22 | light |
| Thrush et al. 1994 | Ohmeda Oxyshuttle | 22 | 88 | 2·00 | 0·65 | light |
| Thrush et al. 1994 | Ohmeda Oxyshuttle | 22 | 88 | 1·00 | 0·65 | light |
| Thrush et al. 1994 | Ohmeda Oxyshuttle | 22 | 88 | 0·00 | 0·43 | light |
| Thrush et al. 1994 | Ohmeda Oxyshuttle | 22 | 88 | 1·00 | 0·22 | light |
| Thrush et al. 1994 | MiniOx IV | 22 | 88 | 1·00 | 0·65 | light |
| Thrush et al. 1994 | MiniOx IV | 22 | 88 | 0·00 | 0·87 | light |
| Thrush et al. 1994 | MiniOx IV | 22 | 88 | 1·00 | 0·43 | light |
| Thrush et al. 1994 | MiniOx IV | 22 | 88 | 0·00 | 0·43 | light |
| Thrush et al. 1994 | Ohmeda 3700 | 22 | 88 | 1·00 | 1·08 | light |
| Thrush et al. 1994 | Ohmeda 3700 | 22 | 88 | 0·00 | 0·65 | light |
| Thrush et al. 1994 | Ohmeda 3700 | 22 | 88 | 2·00 | 0·43 | light |
| Thrush et al. 1994 | Ohmeda 3700 | 22 | 88 | 1·00 | 0·43 | light |
| Valbuena et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 186 | 675 | 0·30 | 0·54 | light |
| Vesoulis et al. 2021 | nelcor MAX-N | 170 | 2342 | 0·72 | 0·68 | light |
| Wiles et al. 2022 | Nelcor GE B1x5 M/P | 135 | 4197 | -0·28 | 0·09 | light |
| Adler et al. 1998 | Nellcor D-25 | 105 | 105 | 2·80 | 0·51 | medium |
| Burnett et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 8916 | 29344 | -0·41 | 0·08 | medium |
| Crooks et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 246 | 246 | 5·10 | 14·74 | medium |
| Ebmeier et al. 2018 | Philips (IntelliVue MP70) | 89 | 89 | -0·45 | 0·22 | medium |
| Ebmeier et al. 2018 | Masimo (GE Marquette Rac-4A) | 48 | 48 | -0·11 | 0·26 | medium |
| Fawzy et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 278 | 278 | 0·79 | 0·14 | medium |
| Feiner et al. 2007 | Masimo Radical (clip) | 7 | 154 | 0·99 | 0·85 | medium |
| Feiner et al. 2007 | Nellcor N-595 (clip) | 7 | 154 | 0·64 | 0·65 | medium |
| Feiner et al. 2007 | Nonin 9700 (clip) | 7 | 154 | 0·03 | 0·78 | medium |
| Feiner et al. 2007 | Masimo Radical (adhesive) | 7 | 154 | -0·34 | 0·80 | medium |
| Feiner et al. 2007 | Nellcor N-595 & OxiMax (adhesive) | 7 | 154 | 0·83 | 0·95 | medium |
| Feiner et al. 2007 | Nonin 9700 (adhesive) | 7 | 154 | 0·51 | 1·07 | medium |
| Pilcher et al. 2020 | 14 models | 157 | 157 | -1·10 | 0·13 | medium |
| Valbuena et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 135 | 481 | 0·27 | 0·53 | medium |
| Wiles et al. 2022 | Not Specified | 34 | 1241 | 0·33 | 0·19 | medium |
| Abrams et al. 2002 | Nellcor N200 | 16 | 16 | 3·60 | 0·33 | dark |
| Adler et al. 1998 | Nellcor D-25 | 34 | 34 | 2·20 | 0·63 | dark |
| Andrist et al. 2022 | Not Specified | 183 | 2005 | 4·30 | 0·37 | dark |
| Barker et al. 2023 | Masimo SET¬Æ pulse oximeters | 39 | 3201 | -0·20 | 0·23 | dark |
| Bickler et al. 2005 | Nonin Onyx | 11 | 119 | 0·15 | 0·47 | dark |
| Bickler et al. 2005 | Novametrix 513 | 11 | 119 | 1·83 | 0·67 | dark |
| Bickler et al. 2005 | Nellcor 5-595 | 11 | 209 | 1·29 | 0·64 | dark |
| Bothma et al. 1996 | Simed S100e | 100 | 100 | 1·20 | 0·19 | dark |
| Bothma et al. 1996 | Nihon Koden | 100 | 100 | 0·85 | 0·25 | dark |
| Bothma et al. 1996 | Ohmeda 3740 (finger) | 100 | 100 | 0·55 | 0·20 | dark |
| Bothma et al. 1996 | Ohmeda 3740 (ear) | 100 | 100 | -1·00 | 0·24 | dark |
| Burnett et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 5177 | 16011 | -0·60 | 0·13 | dark |
| Crooks et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 151 | 151 | 5·40 | 15·99 | dark |
| Ebmeier et al. 2018 | IntelliVue MP70 | 8 | 8 | -2·24 | 0·61 | dark |
| Ebmeier et al. 2018 | GE Marquette Rac-4A | 4 | 4 | -1·10 | 0·20 | dark |
| Fawzy et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 478 | 478 | 1·16 | 0·13 | dark |
| Feiner et al. 2007 | Masimo Radical (clip) | 17 | 372 | 2·14 | 0·73 | dark |
| Feiner et al. 2007 | Nellcor N-595 (clip) | 17 | 371 | 2·04 | 0·52 | dark |
| Feiner et al. 2007 | Nonin 9700 (clip) | 17 | 372 | -0·63 | 0·29 | dark |
| Feiner et al. 2007 | Masimo Radical (adhesive) | 17 | 372 | -0·24 | 0·70 | dark |
| Feiner et al. 2007 | Nellcor N-595 (Oximax, adhesive) | 17 | 359 | 2·13 | 0·60 | dark |
| Feiner et al. 2007 | Nonin 9700 (adhesive) | 17 | 371 | 1·26 | 0·52 | dark |
| Foglia et al. 2017 | Masimo Radical | 14 | 14 | 1·60 | 1·28 | dark |
| Foglia et al. 2017 | Nellcor Oximax | 14 | 14 | 5·40 | 1·36 | dark |
| Jubran et al. 1990 | Nellcor/Ohmeda Biox 3700 | 29 | 43 | 3·30 | 0·50 | dark |
| Ruppel et al. 2023 | Not Specified- Hospital Standard Pulse Oximeter | 201 | 201 | 2·61 | 0·22 | dark |
| Sudat et al. 2023 | Not Specified- Hospital Standard Pulse Oximeter | 2616 | 8626 | 2·45 | 0·14 | dark |
| Valbuena et al. 2022 | Not Specified- Hospital Standard Pulse Oximeter | 51 | 195 | 1·70 | 1·44 | dark |
| Vesoulis et al. 2021 | Nellcor MAX-N | 124 | 2044 | 1·73 | 0·84 | dark |
| Wiles et al. 2022 | Not Specified | 19 | 599 | 0·75 | 0·33 | dark |
| Zeballos et al. 1991 | Hewlett-Packard 47201A | 33 | 33 | -0·40 | 0·30 | dark |
| Zeballos et al. 1991 | Ohmeda Biox IIA | 33 | 33 | 2·10 | 0·35 | dark |
| Zeballos et al. 1991 | Hewlett-Packard 47201A | 11 | 11 | -0·80 | 0·51 | dark |
| Zeballos et al. 1991 | Ohmeda Biox IIA | 11 | 11 | 3·50 | 0·84 | dark |
| Zeballos et al. 1991 | Hewlett-Packard 47201A | 22 | 22 | -4·80 | 1·47 | dark |
| Zeballos et al. 1991 | Ohmeda Biox IIA | 22 | 22 | 9·80 | 1·47 | dark |

## Table S3. Pulse Rate Study Bias by Skin Pigmentation

| Study | Device | Sample | Data Pairs | Mean Bias | SE | Skin Pigmentation Group |
| --- | --- | --- | --- | --- | --- | --- |
| Bent et al. 2020 | Empatica E4 | 15 | 41655 | -0·63 | 4·71 | light |
| Bent et al. 2020 | Fitbit Charge 2 | 15 | 3127 | -4·06 | 3·80 | light |
| Bent et al. 2020 | Biovotion Everion | 14 | 48367 | -4·82 | 5·57 | light |
| Bent et al. 2020 | Apple Watch 4 | 15 | 3841 | -0·38 | 2·45 | light |
| Bent et al. 2020 | Xiaomi Miband 3 | 15 | 6724 | -2·96 | 4·67 | light |
| Bent et al. 2020 | Garmin Vivosmart 3 | 15 | 11188 | -3·47 | 3·30 | light |
| Nelson et al. 2019 | Fitbit Charge 2 | 1 | 1446 | -3·47 | 6·17 | light |
| Nelson et al. 2019 | Apple Watch 3 | 1 | 394 | -180 | 7·40 | light |
| Sanudo et al. 2019 | Apple Watch (Version not reported) | 15 | 15 | -1·50 | 0·86 | light |
| Bent et al. 2020 | Apple Watch 4 | 19 | 4263 | -0·19 | 2·28 | medium |
| Bent et al. 2020 | Fitbit Charge 2 | 19 | 3763 | -4·09 | 4.44 | medium |
| Bent et al. 2020 | Garmin Vivosmart 3 | 19 | 12988 | -5·49 | 4·24 | medium |
| Bent et al. 2020 | Xiaomi Miband 3 | 19 | 6288 | -3·77 | 5·19 | medium |
| Bent et al. 2020 | Empatica E4 | 19 | 48843 | -1·53 | 5·34 | medium |
| Bent et al. 2020 | Biovotion Everion | 15 | 48312 | -5·16 | 6·01 | medium |
| Chow et al. 2020 | Garmin Vivosmart HR+ | 80 | 86440 | -0·55 | 1·13 | medium |
| Chow et al. 2020 | Xiaomi Mi Band 2 | 80 | 86440 | -1·35 | 1·69 | medium |
| Sanudo et al. 2019 | Apple Watch (Version not reported) | 30 | 30 | -1·50 | 0·51 | medium |
| Bent et al. 2020 | Xiaomi Miband 3 | 16 | 6922 | -2·93 | 4·66 | dark |
| Bent et al. 2020 | Biovotion Everion | 16 | 49194 | -4·87 | 6·35 | dark |
| Bent et al. 2020 | Garmin Vivosmart 3 | 18 | 12313 | -3·28 | 3·01 | dark |
| Bent et al. 2020 | Apple Watch 4 | 19 | 4631 | -0·39 | 2·22 | dark |
| Bent et al. 2020 | Fitbit Charge 2 | 19 | 3512 | -5·42 | 4·07 | dark |
| Bent et al. 2020 | Empatica E4 | 19 | 45744 | -0·98 | 4·56 | dark |

## Table S4. Heterogeneity Pulse Rate Bias by Skin Tone

| Device Type | Skin Tone Category | Level | Percent of Total Variance | I2 | τ2 |
| --- | --- | --- | --- | --- | --- |
| Pulse Rate | Light | Level 1 | 89·00% | - | 9·62 |
|  |  | Level 2 | 11·00% | 11 | 1·19 |
|  |  | Level 3 | 0·00% | 0 | 0·00 |
|  | Medium | Level 1 | 74·99% | - | 3·24 |
|  |  | Level 2 | 25·01% | 25·01 | 1·08 |
|  |  | Level 3 | 0·00% | 0 | 0·00 |
|  | Dark* | Level 1 | 86·30% | - | 13·65 |
|  |  | Level 2 | 13·70% | 13·7 | 2·17 |
|  |  | Level 3 | - | - | - |
|  | Combined | Level 1 | 69·34% | - | 4·24 |
|  |  | Level 2 | 3·90% | 20·29 | 0·24 |
|  |  | Level 3 | 26·76% | 0·87 | 1·64 |
| Pulse Oximetry | Light | Level 1 | 2·55% | - | 0·03 |
|  |  | Level 2 | 51·01% | 51·01 | 0·61 |
|  |  | Level 3 | 46·44% | 46·44 | 0·55 |
|  | Medium | Level 1 | 4·69% | - | 0·05 |
|  |  | Level 2 | 14·14% | 14·14 | 0·16 |
|  |  | Level 3 | 81·16% | 81·16 | 0·94 |
|  | Dark | Level 1 | 1·53% | - | 0·08 |
|  |  | Level 2 | 98·46% | 98·62 | 5·48 |
|  |  | Level 3 | 0·00% | 0 | 0·00 |
|  | Combined | Level 1 | 1·45% | - | 0·04 |
|  |  | Level 2 | 84·02% | 84·02 | 2·37 |
|  |  | Level 3 | 14·53% | 14·53 | 0·41 |

Note: Rho = 0·9, * = Pulse Rate only had one study for dark skin tone and was a two-level model

## Table S5. Sensitivity Analyses of Pulse Rate Bias by Skin Pigmentation

| ⍴ | Device Type | Skin Pigmentation Category | Unit | Pooled Mean Bias (95% CI) | Pooled SD (SE) | 95% LoA | Arms % | Overall I2 (between and within heterogeneity) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0·3 | Pulse Oximetry | Light | percent | 0·70 (0·16 to 1·24) | 3·90 (1·36) | -6·93 to 8·34 | 3·96** | 97·52% (67·25% and 30·27%) |
|  |  | Medium | percent | 0·29 (-0·63 to 1·21) | 4·71 (1·71) | -9·93 to 9·51 | 4·71** | 95·32% (94·51% and 0·81%) |
|  |  | Dark | percent | 1·27 (0·61 to 1·94) | 3·96 (1·30) | -6·48 to- 9·03 | 4·15** | 98·04 (0·00% and 98·04%) |
| 0·3 | Pulse Rate | Light | bpm | -1·54 (-2·57 to -0·52) | 7·54 (2·13) | -16·32* to 13·23* | – | 0·00% (0·00% and 0·00%) |
|  |  | Medium | bpm | -1·34 (-3·94 to 1·25) | 9·05 (1·75) | -19·07* to 16·39* | – | 0·00% (0·00% and 0·00%) |
|  |  | Dark | bpm | -1·43 (-16·49 to 13·6) | 16·89 (1·31) | -34·54* to 31·68* | – | 0·00% (NA and 0·00%) |
| 0·6 | Pulse Oximetry | Light | percent | 0·70 (0·17 to 1·24) | 3·90 (1·36) | -6·94 to 8·34 | 3·96** | 97·50% (58·83% and 38·67%) |
|  |  | Medium | percent | 0·29 (-0·64 to 1·21) | 4·71 (1·71) | -8·93 to 9·51 | 4·71** | 95·35% (92·49% and 2·86%) |
|  |  | Dark | percent | 1·27 (0·59 to 1·95) | 3·96 (1·30) | -6·49 to 9·03 | 4·15** | 98·26% (0·00% and 98·26%) |
| 0·6 | Pulse Rate | Light | bpm | -1·43 (-3·34 to 0·47) | 7·54 (2·13) | -16·21* to 13·34* | – | 0·00% (0·00% and 0·00%) |
|  |  | Medium | bpm | -1·28 (-5·05 to 2·49) | 9·05 (1·75) | -19·01* to 16·45* | – | 0·00% (0·00% and 0·00%) |
|  |  | Dark | bpm | -0·50 (-20·72 to 19·72) | 16·89 (1·31) | -33·61* to 32·61* | – | 0·00% (NA and 0·00%) |

Note: ⍴ was used in CHE models to pool both mean bias and SD; Bold = statistically significant; * = exceeds ANSI Standards for Pulse Rate; ** = exceeds FDA Guidance for Pulse Oximetry. Arms = accuracy root mean square, LoA = limits of agreement, CI = confidence interval, SD = standard deviation, SE = standard error.

## Table S6. Pulse Oximetry Sensitivity Analysis with Crooks et al. (2022) Removed

| Skin Tone Category | Pooled Mean Bias (95% CI) | Pooled SD (SE) | 95% LoA | Arms % | Overall I2 (between and within heterogeneity) |
| --- | --- | --- | --- | --- | --- |
| Light | 0·70 (0·17 to 1·22) | 2·91 (1·15) | -5·02 to 6·41 | 3·00 | 97·51% (46·5% and 51·01%) |
| Medium | 0·27 (-0·65 to 1·19) | 2·90 (1·28) | -5·41 to 5·95 | 2·91 | 95·63% (81·45% and 14·18%) |
| Dark | 1·26 (0·58 to 1·94) | 3·16 (1·17) | -4·93 to 7·46 | 3·40* | 98·51% (0·00% and 98·51%) |
| Combined | 0·82 (0·29 to 1·35) | 2·85 (1·13) | -4·77 to 6·41 | 2·97 | 98·59% (14·37% and 84·23%) |

Note: ⍴= 0.9  was used in CHE models to pool both mean bias and SD; Bold = statistically significant; * = exceeds FDA Guidance for Pulse Oximetry. Arms = accuracy root mean square, LoA = limits of agreement, CI = confidence interval, SD = standard deviation, SE = standard error.

## Table S7. Shi et al. (2022)28 Manuscripts Not Included

| Author | Reason |
| --- | --- |
| Avant 1997 | Could not confirm if bias was calculated as test - reference or reference - test |
| Hinkelbein 2006 | Examined nail polish and not skin tone |
| Hinkelbein 2007 | Examined nail polish and not skin tone |

## Table S8. Coding Skin Tone, Race, and Ethnicity to Low, Medium, and High Skin Pigmentation Categories

| Skin pigmentation measurement methods | The number of classification categories as reported | Low (light) skin pigmentation | Medium skin pigmentation | High (dark) skin pigmentation |
| --- | --- | --- | --- | --- |
| Fitzpatrick scale | Three categories | ‘Light (Type I to Type II)’, or ‘light (score of 1 or 2)’ | ‘Medium (Type III to Type IV)’, or ‘medium (score of 3 or 4)’ | ‘Dark (Type V to Type VI)’, or ‘dark (score of 5 or 6)’ |
| Munsell color system | Three categories | ‘Light’ | ‘Medium’ | ‘Dark’ |
| Using race and ethnicity to indicate skin pigmentation | Three categories | ‘Light (Caucasian)’ | ‘Intermediate (Hispanic, Indian, Filipino, Vietnamese)’ | ‘Dark (African American)’ |
| Objective quantification using a reflectance spectrophotometer | One category |  |  | ‘Dark pigmentation’ |

Note: Adapted from Shi et al. (2022)28

## Table S9. Excluded Manuscripts

| Study | Reason For Exclusion |
| --- | --- |
| Banerjee et al. 2020 | Abstract |
| Bhoyar et al. 2010 | Abstract |
| Blanchet et al. 2021 | Abstract |
| Foglia et al. 2016 | Abstract |
| Kapur et al. 2011 | Abstract |
| Nickel et al. 2018 | Abstract |
| Sathiyakumar et al. 2011 | Abstract |
| Vybornova et al. 2021 | Abstract |
| Cui et al. 1990 | Benchtop testing, no human subjects |
| Fallow et al. 2013 | Benchtop testing, no human subjects |
| Mark et al. 2021 | Benchtop testing, no human subjects |
| Avant et al. 1997 | Cannot confirm how mean bias was calculated |
| Arnold et al. 2019 | Collected race, but did not report on it |
| Gillinov et al. 2017 | Collected race, but did not report on it |
| Huynh et al. 2021 | Collected race, but did not report on it |
| Reddy et al. 2018 | Collected race, but did not report on it |
| Seshadri et al. 2020 | Collected race, but did not report on it |
| Hochstadt et al. 2020 | Collected skin tone but did not report data categorized by skin tone |
| Webster et al. 2021 | Collected skin tone but did not report data categorized by skin tone |
| Giggins et al. 2022 | Collected skin tone but did not report data categorized by skin tone |
| Stove et al. 2023 | Combined Fitzpatrick Skin Scale Types 1-3 into one category |
| Blondel et al. 2021 | Could not obtain full text |
| Gutmann et al. 2003 | Could not obtain full text |
| Housset et al. 1957 | Could not obtain full text |
| Lee et al. 1992 | Could not obtain full text |
| Yan et al. 2017 | Device uses PPG and ECG |
| Scardulla et al. 2018 | Device uses PPG and ECG |
| Mantri et al. 2022 | Device utilized audio and PPG |
| Jang et al. 2020 | Did not comment on primary endpoints of interest |
| Nikseresht et al. 2021 | Did not comment on primary endpoints of interest |
| Velichkovska et al. 2022 | Did not comment on primary endpoints of interest |
| Mukherjee et al. 2018 | Did not comment on primary endpoints of interest |
| Murphy et al. 2018 | Did not comment on primary endpoints of interest |
| Gibney et al. 2010 | Did not comment on primary endpoints of interest |
| Huang et al. 2021 | Did not comment on primary endpoints of interest |
| Nachman et al. 2022 | Did not comment on primary endpoints of interest |
| Hussain et al. 2009 | Did not discuss PPG and/or skin tone |
| Kashimutt et al. 2014 | Did not discuss PPG and/or skin tone |
| Bonafide et al. 2020 | Did not discuss PPG and/or skin tone |
| Lipchak et al. 2021 | Did not discuss PPG and/or skin tone |
| Ojeda et al. 2020 | Did not discuss PPG and/or skin tone |
| Rahnama'i et al. 2006 | Did not discuss PPG and/or skin tone |
| Rodriguez-Quinonez et al. 2012 | Did not discuss PPG and/or skin tone |
| Sartor et al. 2021 | Did not discuss PPG and/or skin tone |
| Vandeput et al. 1990 | Did not discuss PPG and/or skin tone |
| Bangash et al. 2022 | Does not provide required statistics |
| Henry et al. 2022 | Does not provide required statistics |
| Hermand et al. 2019 | Does not provide required statistics |
| Lee et al. 1993 | Does not provide required statistics |
| Puranen et al. 2020 | Does not provide required statistics |
| Ray et al. 2021 | Does not provide required statistics |
| Ries et al. 1989 | Does not provide required statistics |
| Shcherbina et al. 2017 | Does not provide required statistics |
| Sjoding et al. 2020 | Does not provide required statistics |
| Wiles et al. 2022 | Does not provide required statistics |
| Wong et al. 2021 | Does not provide required statistics |
| Brooks et al. 2020 | Does not provide required statistics |
| Ceci et al. 1988 | Does not provide required statistics |
| Escourrou et al. 1990 | Does not provide required statistics |
| Espinosa et al. 2020 | Does not provide required statistics |
| Etiwy et al. 2019 | Does not provide required statistics |
| Gabrielczyk et al. 1988 | Does not provide required statistics |
| Harris et al. 2019 | Does not provide required statistics |
| Harris et al. 2016 | Does not provide required statistics |
| Harskamp et al. 2021 | Does not provide required statistics |
| Horton et al. 2017 | Does not provide required statistics |
| Menghini et al. 2019 | Does not provide required statistics |
| Pasadyn et al. 2019 | Does not provide required statistics |
| Ries et al. 1985 | Does not provide required statistics |
| Ross et al. 2014 | Does not provide required statistics |
| Schallom et al. 2018 | Does not provide required statistics |
| Smyth et al. 1986 | Does not provide required statistics |
| Spierer et al. 2015 | Does not provide required statistics |
| Stewart et al. 1991 | Does not provide required statistics |
| Valbuena et al. 2022 | Does not provide required statistics |
| Wallen et al. 2016 | Does not provide required statistics |
| Wang et al. 1985 | Does not provide required statistics |
| Pipek et al. 2021 | Does not provide required statistics |
| Harrison et al. 2011 | Duplicate record of another record this search |
| Harrison et al. 2011 | Duplicate record of another record this search |
| Pipek et al. 2021 | Duplicate record of another record this search |
| Ralston et al. 1991 | Duplicate record of another record this search |
| Ralston et al. 1991 | Duplicate record of another record this search |
| Wong et al. 2022 | Duplicate record of another record this search |
| Sjoding et al. 2021 | Duplicate record of another record this search |
| Adler et al. 1998 | Duplicate record of another record this search |
| Avant et al. 1997 | Duplicate record of another record this search |
| Bickler et al. 2005 | Duplicate record of another record this search |
| Bothma et al. 1996 | Duplicate record of another record this search |
| Brooks et al. 2020 | Duplicate record of another record this search |
| Ebmeier et al. 2018 | Duplicate record of another record this search |
| Escourrou et al. 1990 | Duplicate record of another record this search |
| Feiner et al. 2007 | Duplicate record of another record this search |
| Foglia et al. 2017 | Duplicate record of another record this search |
| Gabrielczyk et al. 1988 | Duplicate record of another record this search |
| Harris et al. 2016 | Duplicate record of another record this search |
| Harris et al. 2019 | Duplicate record of another record this search |
| Harskamp et al. 2021 | Duplicate record of another record this search |
| Hinkelbein et al. 2006 | Duplicate record of another record this search |
| Hinkelbein et al. 2007 | Duplicate record of another record this search |
| Jubran et al. 1990 | Duplicate record of another record this search |
| Lee et al. 1993 | Duplicate record of another record this search |
| Ries et al. 1985 | Duplicate record of another record this search |
| Ries et al. 1989 | Duplicate record of another record this search |
| Ross et al. 2014 | Duplicate record of another record this search |
| Schallom et al. 2018 | Duplicate record of another record this search |
| Smyth et al. 1986 | Duplicate record of another record this search |
| Stewart et al. 1991 | Duplicate record of another record this search |
| Valbuena et al. 2021 | Duplicate record of another record this search |
| Vesoulis et al. 2021 | Duplicate record of another record this search |
| Wiles et al. 2021 | Duplicate record of another record this search |
| Zeballos et al. 1991 | Duplicate record of another record this search |
| Wong et al. 2022 | Erratum |
| Sjoding et al. 2021 | Erratum |
| Bent et al. 2021 | Erratum |
| Whitehead-Clarke et al. 2021 | Letter |
| Okunlola et al. 2022 | Meta-analysis |
| Shi et al. 2022 | Meta-analysis |
| Garrett et al. 2023 | No skin tone, race, or ethnicity |
| Bai et al. 2018 | No skin tone, race, or ethnicity |
| Bellenger et al. 2021 | No skin tone, race, or ethnicity |
| Benedetti et al. 2021 | No skin tone, race, or ethnicity |
| Berryhill et al. 2020 | No skin tone, race, or ethnicity |
| Boudreaux et al. 2018 | No skin tone, race, or ethnicity |
| Cadmus-Bertram et al. 2017 | No skin tone, race, or ethnicity |
| Cao et al. 2022 | No skin tone, race, or ethnicity |
| de Zambotti et al. 2016 | No skin tone, race, or ethnicity |
| Dooley et al. 2017 | No skin tone, race, or ethnicity |
| Gorny et al. 2017 | No skin tone, race, or ethnicity |
| Held et al. 2022 | No skin tone, race, or ethnicity |
| Hettiarachchi et al. 2019 | No skin tone, race, or ethnicity |
| Kinnunen et al. 2020 | No skin tone, race, or ethnicity |
| Kroll et al. 2016 | No skin tone, race, or ethnicity |
| Mendelson et al. 1988 | No skin tone, race, or ethnicity |
| Miller et al. 2022 | No skin tone, race, or ethnicity |
| Muggeridge et al. 2021 | No skin tone, race, or ethnicity |
| Nuuttila et al. 2021 | No skin tone, race, or ethnicity |
| Renevey et al. 2018 | No skin tone, race, or ethnicity |
| Sarhaddi et al. 2022 | No skin tone, race, or ethnicity |
| Schubert et al. 2018 | No skin tone, race, or ethnicity |
| Schuurmans et al. 2020 | No skin tone, race, or ethnicity |
| Sen-Gupta et al. 2019 | No skin tone, race, or ethnicity |
| Stucky et al. 2021 | No skin tone, race, or ethnicity |
| Wang et al. 2017 | No skin tone, race, or ethnicity |
| List et al. 1991 | Non-English text |
| Dyer et al. 2020 | Not a primary study (commentary on another study) |
| Ferrari et al. 2022 | Not a primary study (commentary on another study) |
| Holder et al. 2022 | Not a primary study (commentary on another study) |
| Norton et al. 2022 | Not a primary study (commentary on another study) |
| Papaioannou et al. 2017 | Not a primary study (commentary on another study) |
| Philip et al. 2021 | Not a primary study (commentary on another study) |
| Ploen et al. 2016 | Oral presentation |
| Todd et al. 2021 | Oral presentation |
| Mukherjee et al. 2020 | Prototype device |
| Harrison et al. 2011 | Published in textbok, not peer reviewed |
| Królak et al. 2020 | Published in textbok, not peer reviewed |
| Sjoding et al. 2021 | Reply |
| Tobin et al. 2022 | Reply |
| Colvonen et al. 2021 | Response |
| Agache et al. 1994 | Review Paper |
| Aldallal et al. 2019 | Review paper |
| Fine et al. 2021 | Review Paper |
| Hay et al. 1987 | Review Paper |
| Hunasikatti et al. 2021 | Review Paper |
| Kamal et al. 1989 | Review paper |
| Knight et al. 2022 | Review paper |
| Ralston et al. 1991 | Review paper |
| Ray et al. 2021 | Review paper |
| Taylor et al. 1986 | Review paper |
| Tobin et al. 2022 | Review Paper |
| Ajmal et al. 2021 | Simulation, no human subjects |
| Flynn et al. 2021 | Simulation, no human subjects |
| Chatterjee et al. 2019 | Simulation, no human subjects |
| Fine et al. 2020 | Simulation, no human subjects |
| Hinkelbein et al. 2007 | Studied the effect of nail polish and not skin tone |
| Hinkelbein et al. 2006 | Studied the effect of nail polish and not skin tone |
| Emery et al. 1987 | Study noted significant difficulty in classifying subject skin tone |
| Baek et al. 2018 | Studying pulse oximetry but reference is not blood gas |
| Mendelson et al. 1987 | Studying pulse oximetry but reference is not blood gas |
| Olmo Arroyo et al. 2017 | Studying pulse oximetry but reference is not blood gas |
| Bickler et al. 2013 | Studying pulse oximetry but reference is not blood gas |
| Cahan et al. 1990 | Studying pulse oximetry but reference is not blood gas |
| Smith et al. 2019 | Studying pulse oximetry but reference is not blood gas |
| Stell et al. 2022 | Studying pulse oximetry but reference is not blood gas |
| Witting et al. 2008 | Studying pulse oximetry but reference is not blood gas |
| Schröder et al. 2023 | Studying pulse oximetry but reference is not blood gas |
| Ahmed et al. 2018 | Studying pulse rate but reference is not ECG |
| Preejith et al. 2016 | Studying pulse rate but reference is not ECG |
| Spaccarotella et al. 2022 | Studying pulse rate but reference is not ECG |
| Powell et al. 2013 | Survey data, not directly studying PPG |

# Figures

a.b.

c.

## Figure S1. Funnel Plots Pulse Oximetry

A. Light Skin Pigmentation, B. Medium Skin Pigmentation, C. Dark Skin Pigmentation. These plot individual study effect estimates against their respective precision, represented by the standard error. On these plots, the vertical axis shows the effect size, while the horizontal axis shows the standard error. When there is no bias, the plot looks like an inverted funnel; larger, more precise studies cluster towards the top, closer to the true effect size, and smaller, less precise studies spread towards the bottom, indicating a wider spread in effect estimates. Asymmetry in the funnel plot can hint at potential publication bias, but it might also indicate other biases or study heterogeneity.

a.b.

c.

## Figure S2. Funnel Plot Pulse Rate

A. Light Skin Pigmentation, B. Medium Skin Pigmentation, C. Dark Skin Pigmentation. These plot individual study effect estimates against their respective precision, represented by the standard error. On these plots, the vertical axis shows the effect size, while the horizontal axis shows the standard error. When there is no bias, the plot looks like an inverted funnel; larger, more precise studies cluster towards the top, closer to the true effect size, and smaller, less precise studies spread towards the bottom, indicating a wider spread in effect estimates. Asymmetry in the funnel plot can hint at potential publication bias, but it might also indicate other biases or study heterogeneity.

# Materials

## Search Report: PPG and Skin of Color

Whitney A Townsend, MLIS

whitneyt@umich.edu

All searches run from inception of the database to 04/15/2022 unless otherwise noted

All searches run without limits unless otherwise noted

Original PubMed search strategy translated and adapted to other databases using http://sr-accelerator.com/#/polyglot and the searcher's discretion.

Please cite:

SR Accelerator Polyglot search translation tool29.

Databases

- PubMed (PubMed.gov)

- Embase.com (including Embase Classic)

- Scopus.com

- Web of Science Core Collection (SCI-EXPANDED, SSCI, A&HCI, CPCI-S, CPCI-SSH, BKCI-S, BKCI-SSH, ESCI, CCR-EXPANDED)

- CINAHL Complete (Ebsco)

- ClinicalTrials.gov

Total results before deduplication: 8582

Total results after initial DistillerSR deduplication:  5082 (3500 duplicates removed)

SEARCH NOTES:

Antiquated, non-standard, and potentially offensive terms for racial and ethnic groups were included in search strategies in order to conduct a sensitive, comprehensive search for relevant studies. Language has changed over time and varies geographically and by community, so a thorough search must include the terms that have been used in past and present literature. The authors recognize and acknowledge the inappropriate and harmful nature of these terms.

PubMed - 1548 on 4/15/22

(("Race Factors"[MeSH Terms] OR "Skin Pigmentation"[MeSH Terms] OR Racial Groups [MeSH Terms] OR "Indigenous Peoples"[MeSH Terms] OR "Arabs"[MeSH Terms] OR Jews[mesh] OR Roma[mesh] OR "Ethnicity"[MeSH Terms] OR

pigment*[tiab] OR racial[tiab] OR race[tiab] OR biracial[tiab] OR bi-racial[tiab] OR multiracial[tiab] OR multi-racial[tiab] OR "multiple race"[tiab] OR "multiple races"[tiab] OR racism[tiab] OR racist[tiab] OR African*[tiab] OR Black[tiab] OR Blacks[tiab] OR negro*[tiab] OR Native American*[tiab] OR Indian*[tiab] OR Indigenous[tiab] OR First Nation People*[tiab] OR "Native People"[tiab] OR Inuit*[tiab] OR Arab[tiab] OR Arabs[tiab] OR Arabic[tiab] OR "Middle East*"[tiab] OR Asian*[tiab] OR Asiatic*[tiab] OR Hispanic*[tiab] OR Latin*[tiab]

OR Hawaiian*[tiab] OR Pacific Islander*[tiab] OR Aborigin*[tiab] OR "people of color"[tiab] OR "people of colour"[tiab] OR BIPOC[tiab] OR "skin of color"[tiab] OR "skin of colour"[tiab] OR "brown skin"[tiab] OR "black skin"[tiab] OR "dark skin"[tiab] OR "dark skinned"[tiab] OR "darker skin"[tiab] OR melanin[tiab] OR ethnic*[tiab] OR skin tone*[tiab] OR skin color*[tiab] OR skin colour*[tiab] OR "skin phototypes"[tiab] OR hyperpigmentation[tiab] OR hypopigmentation[tiab] OR dyschromia[tiab]) AND (Photoplethysmography[MeSH Terms] OR PPG[tiab] OR

Photoplethysmogra*[tiab] OR Plethysmogra*[tiab] OR Optometric*[tiab] OR "Apple Watch"[tiab] OR Alivecor[tiab] OR Kardiamobile[tiab] OR Pulse oximeter*[tiab] OR pulse-ox*[tiab] OR Cuffless[tiab] OR "Empatica E4"[tiab] OR "Finger probe*"[tiab] OR Fitbit[tiab] OR "Smart watch*"[tiab] OR smartwatch*[tiab] OR Xiaomi[tiab] OR Biovotion[tiab] OR "Basis peak"[tiab] OR "Polar M600"[tiab] OR Misfit[tiab] OR Withings[tiab] OR Mio[tiab] OR Samsung[tiab] OR PulseON[tiab] OR TomTom[tiab] OR Empatica[tiab] OR Optoelectronic[tiab] OR "reflection index*"[tiab] OR "stiffness index*"[tiab])) NOT (animals[mesh] NOT humans[mesh])

Embase - 2652 on 4/15/22

((Race/exp OR "Skin Pigmentation"/exp OR "Ancestry Group"/exp OR "Ethnic Group"/exp OR Jew/exp OR Ethnicity/exp OR Racism/exp OR pigment*:ti,ab OR racial:ti,ab OR race:ti,ab OR biracial:ti,ab OR bi-racial:ti,ab OR multiracial:ti,ab OR multi-racial:ti,ab OR "multiple race":ti,ab OR "multiple races":ti,ab OR racism:ti,ab OR racist:ti,ab OR African*:ti,ab OR Black:ti,ab OR Blacks:ti,ab OR negro*:ti,ab OR "Native American*":ti,ab OR Indian*:ti,ab OR Indigenous:ti,ab OR "First Nation People*":ti,ab OR "Native People":ti,ab OR Inuit*:ti,ab OR Arab:ti,ab OR Arabs:ti,ab OR Arabic:ti,ab OR "Middle East*":ti,ab OR Asian*:ti,ab OR Asiatic*:ti,ab OR Hispanic*:ti,ab OR Latin*:ti,ab OR Hawaiian*:ti,ab OR "Pacific Islander*":ti,ab OR Aborigin*:ti,ab OR "people of color":ti,ab OR "people of colour":ti,ab OR BIPOC:ti,ab OR "skin of color":ti,ab OR "skin of colour":ti,ab OR "brown skin":ti,ab OR "black skin":ti,ab OR "dark skin":ti,ab OR "dark skinned":ti,ab OR "darker skin":ti,ab OR melanin:ti,ab OR ethnic*:ti,ab OR "skin tone*":ti,ab OR "skin color*":ti,ab OR "skin colour*":ti,ab OR "skin phototypes":ti,ab OR hyperpigmentation:ti,ab OR hypopigmentation:ti,ab OR dyschromia:ti,ab) AND ('photoelectric plethysmography'/exp OR PPG:ti,ab OR Photoplethysmogra*:ti,ab OR Plethysmogra*:ti,ab OR Optometric*:ti,ab OR "Apple Watch":ti,ab OR Alivecor:ti,ab OR Kardiamobile:ti,ab OR "Pulse oximeter*":ti,ab OR pulse-ox*:ti,ab OR Cuffless:ti,ab OR "Empatica E4":ti,ab OR "Finger probe*":ti,ab OR Fitbit:ti,ab OR "Smart watch*":ti,ab OR smartwatch*:ti,ab OR Xiaomi:ti,ab OR Biovotion:ti,ab OR "Basis peak":ti,ab OR "Polar M600":ti,ab OR Misfit:ti,ab OR Withings:ti,ab OR Mio:ti,ab OR Samsung:ti,ab OR PulseON:ti,ab OR TomTom:ti,ab OR Empatica:ti,ab OR Optoelectronic:ti,ab OR "reflection index*":ti,ab OR "stiffness index*":ti,ab)) NOT ([animals]/lim NOT [humans]/lim)

Scopus - 2654 on 4/15/22

( TITLE-ABS-KEY ( pigment* )  OR  TITLE-ABS-KEY ( racial )  OR  TITLE-ABS-KEY ( race )  OR  TITLE-ABS-KEY ( biracial )  OR  TITLE-ABS-KEY ( bi-racial )  OR  TITLE-ABS-KEY ( multiracial )  OR  TITLE-ABS-KEY ( multi-racial )  OR  TITLE-ABS-KEY ( {multiple race} )  OR  TITLE-ABS-KEY ( {multiple races} )  OR  TITLE-ABS-KEY ( racism )  OR  TITLE-ABS-KEY ( racist )  OR  TITLE-ABS-KEY ( african* )  OR  TITLE-ABS-KEY ( black  W/3  skin )  OR  TITLE-ABS-KEY ( negro* )  OR  TITLE-ABS-KEY ( "Native American*" )  OR  TITLE-ABS-KEY ( indian* )  OR  TITLE-ABS-KEY ( indigenous )  OR  TITLE-ABS-KEY ( "First Nation People*" )  OR  TITLE-ABS-KEY ( {Native People} )  OR  TITLE-ABS-KEY ( inuit* )  OR  TITLE-ABS-KEY ( arab )  OR  TITLE-ABS-KEY ( arabs )  OR  TITLE-ABS-KEY ( arabic )  OR  TITLE-ABS-KEY ( "Middle East*" )  OR  TITLE-ABS-KEY ( asian* )  OR  TITLE-ABS-KEY ( asiatic* )  OR  TITLE-ABS-KEY ( hispanic* )  OR  TITLE-ABS-KEY ( latin* )  OR  TITLE-ABS-KEY ( hawaiian* )  OR  TITLE-ABS-KEY ( "Pacific Islander*" )  OR  TITLE-ABS-KEY ( aborigin* )  OR  TITLE-ABS-KEY ( {people of color} )  OR  TITLE-ABS-KEY ( {people of colour} )  OR  TITLE-ABS-KEY ( bipoc )  OR  TITLE-ABS-KEY ( {skin of color} )  OR  TITLE-ABS-KEY ( {skin of colour} )  OR  TITLE-ABS-KEY ( {brown skin} )  OR  TITLE-ABS-KEY ( {black skin} )  OR  TITLE-ABS-KEY ( {dark skin} )  OR  TITLE-ABS-KEY ( {dark skinned} )  OR  TITLE-ABS-KEY ( {darker skin} )  OR  TITLE-ABS-KEY ( melanin )  OR  TITLE-ABS-KEY ( ethnic* )  OR  TITLE-ABS-KEY ( {skin tone} )  OR  TITLE-ABS-KEY ( {skin tones} )  OR  TITLE-ABS-KEY ( {skin color} )  OR  TITLE-ABS-KEY ( {skin colors} )  OR  TITLE-ABS-KEY ( {skin colour} )  OR  TITLE-ABS-KEY ( {skin colours} )  OR  TITLE-ABS-KEY ( {skin phototypes} )  OR  TITLE-ABS-KEY ( hyperpigmentation )  OR  TITLE-ABS-KEY ( hypopigmentation )  OR  TITLE-ABS-KEY ( dyschromia ) )

AND

( TITLE-ABS-KEY ( ppg )  OR  TITLE-ABS-KEY ( photoplethysmogra* )  OR  TITLE-ABS-KEY ( plethysmogra* )  OR  TITLE-ABS-KEY ( optometric* )  OR  TITLE-ABS-KEY ( {Apple Watch} )  OR  TITLE-ABS-KEY ( alivecor )  OR  TITLE-ABS-KEY ( kardiamobile )  OR  TITLE-ABS-KEY ( "Pulse oximeter*" )  OR  TITLE-ABS-KEY ( pulse-ox* )  OR  TITLE-ABS-KEY ( cuffless )  OR  TITLE-ABS-KEY ( "Empatica E4" )  OR  TITLE-ABS-KEY ( "Finger probe*" )  OR  TITLE-ABS-KEY ( fitbit )  OR  TITLE-ABS-KEY ( "Smart watch*" )  OR  TITLE-ABS-KEY ( smartwatch* )  OR  TITLE-ABS-KEY ( xiaomi )  OR  TITLE-ABS-KEY ( biovotion )  OR  TITLE-ABS-KEY ( "Basis peak" )  OR  TITLE-ABS-KEY ( "Polar M600" )  OR  TITLE-ABS-KEY ( withings )  OR  TITLE-ABS-KEY ( samsung )  OR  TITLE-ABS-KEY ( pulseon )  OR  TITLE-ABS-KEY ( tomtom )  OR  TITLE-ABS-KEY ( empatica )  OR  TITLE-ABS-KEY ( optoelectronic )  OR  TITLE-ABS-KEY ( "reflection index" )  OR  TITLE-ABS-KEY ( "stiffness index" ) )

Web of Science Core Collection - 1543 on 4/15/22

((TS="Race Factors" OR TS="Skin Pigmentation" OR TS="Racial Groups" OR TS="Indigenous Peoples" OR TS=Arabs OR TS=Jews OR TS=Roma OR TS=Ethnicity OR (TI=pigment* OR AB=pigment*) OR (TI=racial OR AB=racial) OR (TI=race OR AB=race) OR (TI=biracial OR AB=biracial) OR (TI=bi-racial OR AB=bi-racial) OR (TI=multiracial OR AB=multiracial) OR (TI=multi-racial OR AB=multi-racial) OR (TI="multiple race" OR AB="multiple race") OR (TI="multiple races" OR AB="multiple races") OR (TI=racism OR AB=racism) OR (TI=racist OR AB=racist) OR (TI=African* OR AB=African*) OR ((TI=Black OR AB=Black) AND (TI=skin OR AB=skin)) OR (TI=negro* OR AB=negro*) OR (TI="Native American*" OR AB="Native American*") OR (TI=Indian* OR AB=Indian*) OR (TI=Indigenous OR AB=Indigenous) OR (TI="First Nation People*" OR AB="First Nation People*") OR (TI="Native People" OR AB="Native People") OR (TI=Inuit* OR AB=Inuit*) OR (TI=Arab OR AB=Arab) OR (TI=Arabs OR AB=Arabs) OR (TI=Arabic OR AB=Arabic) OR (TI="Middle East*" OR AB="Middle East*") OR (TI=Asian* OR AB=Asian*) OR (TI=Asiatic* OR AB=Asiatic*) OR (TI=Hispanic* OR AB=Hispanic*) OR (TI=Latin* OR AB=Latin*) OR (TI=Hawaiian* OR AB=Hawaiian*) OR (TI="Pacific Islander*" OR AB="Pacific Islander*") OR (TI=Aborigin* OR AB=Aborigin*) OR (TI="people of color" OR AB="people of color") OR (TI="people of colour" OR AB="people of colour") OR (TI=BIPOC OR AB=BIPOC) OR (TI="skin of color" OR AB="skin of color") OR (TI="skin of colour" OR AB="skin of colour") OR (TI="brown skin" OR AB="brown skin") OR (TI="black skin" OR AB="black skin") OR (TI="dark skin" OR AB="dark skin") OR (TI="dark skinned" OR AB="dark skinned") OR (TI="darker skin" OR AB="darker skin") OR (TI=melanin OR AB=melanin) OR (TI=ethnic* OR AB=ethnic*) OR (TI="skin tone*" OR AB="skin tone*") OR (TI="skin color*" OR AB="skin color*") OR (TI="skin colour*" OR AB="skin colour*") OR (TI="skin phototypes" OR AB="skin phototypes") OR (TI=hyperpigmentation OR AB=hyperpigmentation) OR (TI=hypopigmentation OR AB=hypopigmentation) OR (TI=dyschromia OR AB=dyschromia)) AND (ALL=Photoplethysmography OR (TI=PPG OR AB=PPG) OR (TI=Photoplethysmogra* OR AB=Photoplethysmogra*) OR (TI=Plethysmogra* OR AB=Plethysmogra*) OR (TI=Optometric* OR AB=Optometric*) OR (TI="Apple Watch" OR AB="Apple Watch") OR (TI=Alivecor OR AB=Alivecor) OR (TI=Kardiamobile OR AB=Kardiamobile) OR (TI="Pulse oximeter*" OR AB="Pulse oximeter*") OR (TI=pulse-ox* OR AB=pulse-ox*) OR (TI=Cuffless OR AB=Cuffless) OR (TI="Empatica E4" OR AB="Empatica E4") OR (TI="Finger probe*" OR AB="Finger probe*") OR (TI=Fitbit OR AB=Fitbit) OR (TI="Smart watch*" OR AB="Smart watch*") OR (TI=smartwatch* OR AB=smartwatch*) OR (TI=Xiaomi OR AB=Xiaomi) OR (TI=Biovotion OR AB=Biovotion) OR (TI="Basis peak" OR AB="Basis peak") OR (TI="Polar M600" OR AB="Polar M600") OR (TI=Misfit OR AB=Misfit) OR (TI=Withings OR AB=Withings) OR (TI=Mio OR AB=Mio) OR (TI=Samsung OR AB=Samsung) OR (TI=PulseON OR AB=PulseON) OR (TI=TomTom OR AB=TomTom) OR (TI=Empatica OR AB=Empatica) OR (TI=Optoelectronic OR AB=Optoelectronic) OR (TI="reflection index*" OR AB="reflection index*") OR (TI="stiffness index*" OR AB="stiffness index*")))

CINAHL Complete - 360 on 4/15/22

(((MH "Race Factors+") OR (MH "Skin Pigmentation+") OR (MH "Racial Groups+") OR (MH "Indigenous Peoples+") OR (MH Arabs+) OR (MH Jews+) OR (MH Roma+) OR (MH Ethnicity+) OR (TI pigment* OR AB pigment*) OR (TI racial OR AB racial) OR (TI race OR AB race) OR (TI biracial OR AB biracial) OR (TI bi-racial OR AB bi-racial) OR (TI multiracial OR AB multiracial) OR (TI multi-racial OR AB multi-racial) OR (TI "multiple race" OR AB "multiple race") OR (TI "multiple races" OR AB "multiple races") OR (TI racism OR AB racism) OR (TI racist OR AB racist) OR (TI African* OR AB African*) OR (TI Black OR AB Black) OR (TI Blacks OR AB Blacks) OR (TI negro* OR AB negro*) OR (TI "Native American*" OR AB "Native American*") OR (TI Indian* OR AB Indian*) OR (TI Indigenous OR AB Indigenous) OR (TI "First Nation People*" OR AB "First Nation People*") OR (TI "Native People" OR AB "Native People") OR (TI Inuit* OR AB Inuit*) OR (TI Arab OR AB Arab) OR (TI Arabs OR AB Arabs) OR (TI Arabic OR AB Arabic) OR (TI "Middle East*" OR AB "Middle East*") OR (TI Asian* OR AB Asian*) OR (TI Asiatic* OR AB Asiatic*) OR (TI Hispanic* OR AB Hispanic*) OR (TI Latin* OR AB Latin*) OR (TI Hawaiian* OR AB Hawaiian*) OR (TI "Pacific Islander*" OR AB "Pacific Islander*") OR (TI Aborigin* OR AB Aborigin*) OR (TI "people of color" OR AB "people of color") OR (TI "people of colour" OR AB "people of colour") OR (TI BIPOC OR AB BIPOC) OR (TI "skin of color" OR AB "skin of color") OR (TI "skin of colour" OR AB "skin of colour") OR (TI "brown skin" OR AB "brown skin") OR (TI "black skin" OR AB "black skin") OR (TI "dark skin" OR AB "dark skin") OR (TI "dark skinned" OR AB "dark skinned") OR (TI "darker skin" OR AB "darker skin") OR (TI melanin OR AB melanin) OR (TI ethnic* OR AB ethnic*) OR (TI "skin tone*" OR AB "skin tone*") OR (TI "skin color*" OR AB "skin color*") OR (TI "skin colour*" OR AB "skin colour*") OR (TI "skin phototypes" OR AB "skin phototypes") OR (TI hyperpigmentation OR AB hyperpigmentation) OR (TI hypopigmentation OR AB hypopigmentation) OR (TI dyschromia OR AB dyschromia)) AND ((MH Photoplethysmography+) OR (TI PPG OR AB PPG) OR (TI Photoplethysmogra* OR AB Photoplethysmogra*) OR (TI Plethysmogra* OR AB Plethysmogra*) OR (TI Optometric* OR AB Optometric*) OR (TI "Apple Watch" OR AB "Apple Watch") OR (TI Alivecor OR AB Alivecor) OR (TI Kardiamobile OR AB Kardiamobile) OR (TI "Pulse oximeter*" OR AB "Pulse oximeter*") OR (TI pulse-ox* OR AB pulse-ox*) OR (TI Cuffless OR AB Cuffless) OR (TI "Empatica E4" OR AB "Empatica E4") OR (TI "Finger probe*" OR AB "Finger probe*") OR (TI Fitbit OR AB Fitbit) OR (TI "Smart watch*" OR AB "Smart watch*") OR (TI smartwatch* OR AB smartwatch*) OR (TI Xiaomi OR AB Xiaomi) OR (TI Biovotion OR AB Biovotion) OR (TI "Basis peak" OR AB "Basis peak") OR (TI "Polar M600" OR AB "Polar M600") OR (TI Misfit OR AB Misfit) OR (TI Withings OR AB Withings) OR (TI Mio OR AB Mio) OR (TI Samsung OR AB Samsung) OR (TI PulseON OR AB PulseON) OR (TI TomTom OR AB TomTom) OR (TI Empatica OR AB Empatica) OR (TI Optoelectronic OR AB Optoelectronic) OR (TI "reflection index*" OR AB "reflection index*") OR (TI "stiffness index*" OR AB "stiffness index*")))

ClinicalTrials.gov - 185 on 4/15/22

Other Terms: Photoplethysmography

GoogleScholar

Skin Tone, Skin Pigmentation, Race, Ethnicity, Photoplethysmography, pulse oximetry, pulse rate, accuracy, bias.

DRAFT SEARCH METHODS:

The following databases were searched from inception to March 15, 2022 in order to identify relevant articles, trials, or meeting abstracts describing photoplethysmography and skin of color: PubMed.gov, Elsevier Embase (including Embase Classic), Elsevier Scopus, Web of Science Core Collection (SCI-EXPANDED; SSCI; A&HCI; CPCI-S; CPCI-SSH; BKCI-S; BKCI-SSH; ESCI; CCR-EXPANDED), CINAHL Complete, GoogleScholar, and ClinicalTrials.gov. Each search utilized controlled vocabulary whenever possible in combination with relevant keywords. No limits were applied to the search. A set of sentinel articles were identified before the search process and were used to generate search terms and test the effectiveness of the strategies in each database. Reference tracking was performed on highly relevant articles. Original search strategies were developed in Pubmed and translated as appropriate to the other databases using the Systematic Review Accelerator Polyglot tool. Citations were deduplicated using DistillerSR's deduplication tools.

## Data Extraction

We developed a comprehensive data extraction form based on the guidelines from the

Cochrane Handbook for Systematic Reviews of Interventions30. A consort chart of the literature search is shown in Figure 1. Studies that met inclusion criteria were included in the meta-analysis (a total of 27 original studies). Two reviewers (BWN, SS) independently extracted data from published reports and pulled participant-level data from Bent et al. supplementary materials.24 A third reviewer (MB, HG) adjudicated any differences between the initial two reviewers and resolved any disagreements by checking manuscripts. All articles were verified with a quality check by the first author. A standardized data extraction system was developed to extract the study characteristics (Authors, Year, Sample size, study setting (inpatient, outpatient, research setting- lab, research setting- real world, other), Skin tone measurements (Skin pigmentation measurement method), sample size by skin tone categorized as light, medium, dark, race31 (sample size by race groups), ethnicity (sample size by ethnicity groups), description of study population (medical vs. healthy), device type (medical, non-medical), reference device (electrocardiogram, arterial blood gas), measurement method (transmittance, reflectance), device location (finger, wrist, ear, forehead, chest, other (If no location was identified in a paper, “other” was selected as the location). and results (mean bias with SD, SE , 95% CI for race, ethnicity and skin tone). For one paper, we used open access to data in supplementary materials to calculate mean bias and SE by each skin pigmentation category. We used skin tone, race, and ethnicity to categorize participants into light, medium, and dark skin pigmentation categories that were defined by race, ethnicity and skin tone.

## Additional Details on Method for Statistical Analysis

Following Shi. et. al. we applied following methods to adjust standard deviation of bias, when repeated measures were taken within a study.

- adjusted SD2= reported SD2 * [(the total number of repeated measures-1) / (the total number of repeated measures - the average number of replications per participant)]

Subsequently the sampling variance of mean bias for such a scenario is taken as (adjusted SD2 / the number of participants).

In order to pool SD across studies, adopting Tipton and Shuster 2017, we used the below transformation to stabilize and normalize the estimates of SD for each study, before feeding into the CHE models.

- point estimate of log(adjusted σ2) = log(adjusted SD2) + 1 / (the number of participants-1)

- sampling variance of log(adjusted σ2) =  2 / (the number of participants - 1)

When a study has only 1 subject, the study was excluded from CHE modeling to pool SD. Antilog was taken to report the pooled SD from the CHE model.

References

1.	Abrams, G. A., Sanders, M. K. & Fallon, M. B. Utility of pulse oximetry in the detection of arterial hypoxemia in liver transplant candidates. Liver Transpl. 8, 391–396 (2002).

2.	Adler, J. N., Hughes, L. A., Vivilecchia, R. & Camargo, C. A., Jr. Effect of skin pigmentation on pulse oximetry accuracy in the emergency department. Acad. Emerg. Med. 5, 965–970 (1998).

3.	Andrist, E., Nuppnau, M., Barbaro, R. P., Valley, T. S. & Sjoding, M. W. Association of Race With Pulse Oximetry Accuracy in Hospitalized Children. JAMA Netw Open 5, e224584 (2022).

4.	Barker, S. J. & Wilson, W. C. Racial effects on Masimo pulse oximetry: a laboratory study. J. Clin. Monit. Comput. 37, 567–574 (2023).

5.	Bickler, P. E., Feiner, J. R. & Severinghaus, J. W. Effects of skin pigmentation on pulse oximeter accuracy at low saturation. Anesthesiology 102, 715–719 (2005).

6.	Bothma, P. A. et al. Accuracy of pulse oximetry in pigmented patients. S. Afr. Med. J. 86, 594–596 (1996).

7.	Burnett, G. W. et al. Self-reported Race/Ethnicity and Intraoperative Occult Hypoxemia: A Retrospective Cohort Study. Anesthesiology 136, 688–696 (2022).

8.	Crooks, C. J. et al. Pulse oximeter measurements vary across ethnic groups: an observational study in patients with COVID-19. Eur. Respir. J. 59, (2022).

9.	Ebmeier, S. J. et al. A two centre observational study of simultaneous pulse oximetry and arterial oxygen saturation recordings in intensive care unit patients. Anaesth. Intensive Care 46, 297–303 (2018).

10.	Fawzy, A. et al. Racial and Ethnic Discrepancy in Pulse Oximetry and Delayed Identification of Treatment Eligibility Among Patients With COVID-19. JAMA Intern. Med. 182, 730–738 (2022).

11.	Feiner, J. R., Severinghaus, J. W. & Bickler, P. E. Dark skin decreases the accuracy of pulse oximeters at low oxygen saturation: the effects of oximeter probe type and gender. Anesth. Analg. 105, S18–S23 (2007).

12.	Foglia, E. E. et al. The Effect of Skin Pigmentation on the Accuracy of Pulse Oximetry in Infants with Hypoxemia. J. Pediatr. 182, 375–377.e2 (2017).

13.	Jubran, A. & Tobin, M. J. Reliability of pulse oximetry in titrating supplemental oxygen therapy in ventilator-dependent patients. Chest 97, 1420–1425 (1990).

14.	McGovern, J. P., Sasse, S. A., Stansbury, D. W., Causing, L. A. & Light, R. W. Comparison of oxygen saturation by pulse oximetry and co-oximetry during exercise testing in patients with COPD. Chest 109, 1151–1155 (1996).

15.	Muñoz, X. et al. Accuracy and reliability of pulse oximetry at different arterial carbon dioxide pressure levels. Eur. Respir. J. 32, 1053–1059 (2008).

16.	Pilcher, J. et al. A multicentre prospective observational study comparing arterial blood gas values to those obtained by pulse oximeters used in adult patients attending Australian and New Zealand hospitals. BMC Pulm. Med. 20, 7 (2020).

17.	Ruppel, H. et al. Evaluating the Accuracy of Pulse Oximetry in Children According to Race. JAMA Pediatr. 177, 540–543 (2023).

18.	Sudat, S. E. K. et al. Racial Disparities in Pulse Oximeter Device Inaccuracy and Estimated Clinical Impact on COVID-19 Treatment Course. Am. J. Epidemiol. 192, 703–713 (2023).

19.	Thrush, D. & Hodges, M. R. Accuracy of pulse oximetry during hypoxemia. South. Med. J. 87, 518–521 (1994).

20.	Valbuena, V. S. M. et al. Racial bias and reproducibility in pulse oximetry among medical and surgical inpatients in general care in the Veterans Health Administration 2013-19: multicenter, retrospective cohort study. BMJ 378, e069775 (2022).

21.	Vesoulis, Z., Tims, A., Lodhi, H., Lalos, N. & Whitehead, H. Racial discrepancy in pulse oximeter accuracy in preterm infants. J. Perinatol. 42, 79–85 (2022).

22.	Wiles, M. D. et al. The effect of patient ethnicity on the accuracy of peripheral pulse oximetry in patients with COVID-19 pneumonitis: a single-centre, retrospective analysis. Anaesthesia 77, 143–152 (2022).

23.	Zeballos, R. J. & Weisman, I. M. Reliability of noninvasive oximetry in black subjects during exercise and hypoxia. Am. Rev. Respir. Dis. 144, 1240–1244 (1991).

24.	Bent, B., Goldstein, B. A., Kibbe, W. A. & Dunn, J. P. Investigating sources of inaccuracy in wearable optical heart rate sensors. NPJ Digit Med 3, 18 (2020).

25.	Nelson, B. W. & Allen, N. B. Accuracy of Consumer Wearable Heart Rate Measurement During an Ecologically Valid 24-Hour Period: Intraindividual Validation Study. JMIR Mhealth Uhealth 7, e10828 (2019).

26.	Sañudo, B., De Hoyo, M., Muñoz-López, A., Perry, J. & Abt, G. Pilot Study Assessing the Influence of Skin Type on the Heart Rate Measurements Obtained by Photoplethysmography with the Apple Watch. J. Med. Syst. 43, 195 (2019).

27.	Chow, H.-W. & Yang, C.-C. Accuracy of Optical Heart Rate Sensing Technology in Wearable Fitness Trackers for Young and Older Adults: Validation and Comparison Study. JMIR Mhealth Uhealth 8, e14707 (2020).

28.	Shi, C. et al. The accuracy of pulse oximetry in measuring oxygen saturation by levels of skin pigmentation: a systematic review and meta-analysis. BMC Med. 20, 267 (2022).

29.	Clark, J. M. et al. Improving the translation of search strategies using the Polyglot Search Translator: a randomized controlled trial. J. Med. Libr. Assoc. 108, 195–207 (2020).

30.	Higgins, J. P. T. & Deeks, J. J. Selecting studies and collecting data. in Cochrane Handbook for Systematic Reviews of Interventions 151–185 (John Wiley & Sons, Ltd, 2008).

31.	Revisions to the Standards for the Classification of Federal Data on Race and Ethnicity. The White House https://obamawhitehouse.archives.gov/omb/fedreg_1997standards.