Supplementary material 05

Table 4. Results assessing filter aerosol penetration, airflow resistance and. The experimental group refers to the different methods included. The control group refers to masks not subjected to any disinfection process. Thus, the closer to the control value, the more advantageous the disinfection method.

| 1. Outcome: Filter aerosol penetration | Experimental |  |  | Control |  |  | Additional information |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 1.1 Liquid Hydrogen Peroxide |  |  |  |  |  |  |  |
| Bergman et al 201037 | 1.33 | 0.4617 | 18 | 1.16 | 0.245 | 18 | 3%, 30 min, 03 cycles |
| Viscusi et al 200760 | 0.75 | 0.242 | 4 | 0.7 | 0.267 | 10 | 3%, 30 min, 01 cycle |
| Viscusi et al 200760 | 0.71 | 0.148 | 4 | 0.7 | 0.267 | 10 | 6%, 30 min, 01 cycle |
| 1.2 Hydrogen Peroxide Vapor |  |  |  |  |  |  |  |
| Bergman et al 201037 | 0.97 | 0.17 | 36 | 1.16 | 0.245 | 18 | Clarus® R, 03 cycles |
| Bergman et al 201037 | 5.16 | 5.3133 | 36 | 1.16 | 2.45 | 18 | STERRAD® 100S, 03 cycles |
| Viscusi et al 200760 | 0.75 | 0.166 | 4 | 0.7 | 0.267 | 10 | STERRAD® 100S, 01 cycle |
| Viscusi et al 200760 | 0.5 | 0.213 | 4 | 0.7 | 0.267 | 10 | STERRAD® NX, 01 cycle |
| Viscusi et al 200961 | 0.7 | 0.4133 | 18 | 0.87 | 0.5083 | 18 | STERRAD® 100S, 01 cycle |
| 1.3 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Bergman et al 201037 | 0.92 | 0.2533 | 18 | 1.16 | 0.245 | 18 | 254 nm, 15 min, 03 cycles |
| Lindsley et al 201546 | 2.66 | 1.4435 | 16 | 2.38 | 1.2192 | 16 | 254 nm, 120 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 3.29 | 1.6252 | 16 | 2.72 | 1.5125 | 16 | 254 nm, 240 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 2.98 | 1.7514 | 16 | 2.33 | 1.3365 | 16 | 254 nm, 470 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 2.76 | 1.5111 | 16 | 2.29 | 1.523 | 16 | 254 nm, 950 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Viscusi et al 200760 | 0.57 | 0.152 | 4 | 0.7 | 0.267 | 10 | 254 nm, 30 min, 01 cycle |
| Viscusi et al 200760 | 0.79 | 0.082 | 4 | 0.7 | 0.267 | 10 | 254 nm, 480 min, 01 cycle |
| Viscusi et al 200961 | 0.93 | 0.4967 | 18 | 0.87 | 0.5083 | 18 | Sterilgard III, 01 cycle |
| 1.4 Ethylene Oxide |  |  |  |  |  |  |  |
| Bergman et al 201037 | 1.02 | 0.255 | 36 | 1.16 | 0.245 | 18 | Amsco® Eagle® 3017, 03 cycles |
| Viscusi et al 200760 | 1.29 | 0.35 | 4 | 0.7 | 0.267 | 10 | EtO 3M 4XL, 01 cycle |
| Viscusi et al 200760 | 1.2 | 0.729 | 4 | 0.7 | 0.267 | 10 | EtO 3M 5XL, 01 cycle |
| Viscusi et al 200961 | 0.78 | 0.455 | 18 | 0.87 | 0.5083 | 18 | Steri-Vac 5XL, 01 cycle |
| 1.5 Dry Heat |  |  |  |  |  |  |  |
| Viscusi et al 200760 | 0.84 | 0.258 | 4 | 0.7 | 0.267 | 10 | oven, 80ºC, 60 min, 01 cycle |
| 1.6 Moist Heat |  |  |  |  |  |  |  |
| Bergman et al 201037 | 0.97 | 0.2433 | 18 | 1.16 | 0.245 | 18 | 60ºC, 80% RH, 30 min, 03 cycles |
| 1.7 Isopropanol solution |  |  |  |  |  |  |  |
| Viscusi et al 200760 | 21.6 | 1.337 | 4 | 0.7 | 0.267 | 10 | 1 minute submersion, 01 cycle |
| Viscusi et al 200760 | 17.8 | 5.508 | 4 | 0.7 | 0.267 | 10 | 1 second submersion, 01 cycle |
| 1.8 Microwave |  |  |  |  |  |  |  |
| Bergman et al 201037 | 1.05 | 0.2567 | 18 | 1.16 | 0.245 | 18 | 2 min exposure, 03 cycles |
| Viscusi et al 200760 | 1.13 | 0.662 | 4 | 0.7 | 0.267 | 10 | 2 min exposure, 01 cycle |
| Viscusi et al 200760 | 1.77 | 0.617 | 4 | 0.7 | 0.267 | 10 | 4 min exposure, 01 cycle |
| Viscusi et al 200961 | 0.73 | 0.48 | 15 | 0.98 | 0.572 | 15 | 2 min exposure, 01 cycle |
| 1.9 Sodium hypochlorite |  |  |  |  |  |  |  |
| Bergman et al 201037 | 1.35 | 0.205 | 18 | 1.16 | 0.245 | 18 | 0.6%, 30 min sub, 03 cycles |
| Viscusi et al 200760 | 0.68 | 0.124 | 4 | 0.7 | 0.267 | 10 | 0.525% in tap water, 01 cycle |
| Viscusi et al 200760 | 3.79 | 3.446 | 4 | 0.7 | 0.267 | 10 | 5.25% in tap water, 01 cycle |
| Viscusi et al 200961 | 0.58 | 0.36 | 18 | 0.87 | 0.5083 | 18 | 0.6%, 30 min sub, 01 cycle |
| 1.10 Autoclave |  |  |  |  |  |  |  |
| Viscusi et al 200760 | 18.7 | 5.263 | 4 | 0.7 | 0.267 | 10 | 121°C (15 psi), 15 min, 01 cycle |
| Viscusi et al 200760 | 34.4 | 9.963 | 4 | 0.7 | 0.267 | 10 | 121°C (15 psi), 30 min, 01 cycle |
| 1.11 Bar soap and water |  |  |  |  |  |  |  |
| Viscusi et al 200760 | 38.8 | 1.269 | 4 | 0.7 | 0.267 | 10 | 02 min, 01 cycle |
| Viscusi et al 200760 | 34.9 | 4.599 | 4 | 0.7 | 0.267 | 10 | 20 min, 01 cycle |
| 2. Filter airflow resistance | Experimental |  |  | Control |  |  | Additional information |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 2.1 Liquid Hydrogen Peroxide |  |  |  |  |  |  |  |
| Bergman et al 201037 | 9.78 | 5.833 | 18 | 10.45 | 3.333 | 18 | 6%, 30 min, 03 cycles |
| 2.2 Hydrogen Peroxide Vapor |  |  |  |  |  |  |  |
| Bergman et al 201037 | 10.62 | 4.833 | 36 | 10.45 | 3.333 | 18 | Clarus® R, 03 cycles |
| Bergman et al 201037 | 10.58 | 3.833 | 36 | 10.45 | 3.333 | 18 | STERRAD® 100S, 03 cycles |
| Viscusi et al 200961 | 7.68 | 0.93 | 18 | 7.52 | 6.567 | 18 | STERRAD® 100S, 01 cycle |
| 2.3 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Bergman et al 2010 | 11.17 | 6.167 | 18 | 10.45 | 3.333 | 18 | 254 nm, 15 min, 03 cycles |
| Lindsley et al 201546 | 11.8 | 3.9553 | 16 | 11.83 | 4.1971 | 16 | 254 nm, 120 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 11.85 | 3.8176 | 16 | 11.87 | 3.8495 | 16 | 254 nm, 240 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 11.66 | 3.7878 | 16 | 11.66 | 3.9936 | 16 | 254 nm, 470 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201546 | 12.08 | 3.8758 | 16 | 12.08 | 3.8757 | 16 | 254 nm, 950 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Viscusi et al 200961 | 7.98 | 5.667 | 18 | 7.52 | 6.567 | 18 | Sterilgard III, 01 cycle |
| 2.4 Ethylene Oxide |  |  |  |  |  |  |  |
| Bergman et al 201037 | 11.47 | 3.667 | 36 | 10.45 | 3.333 | 18 | Amsco® Eagle® 3017, 03 cycles |
| Viscusi et al 200961 | 7.5 | 0.44 | 18 | 7.52 | 6.567 | 18 | Steri-Vac 5XL, 01 cycle |
| 2.5 Moist Heat |  |  |  |  |  |  |  |
| Bergman et al 201037 | 10.18 | 0.15 | 18 | 10.45 | 3.333 | 18 | 60ºC, 80% RH, 30 min, 03 cycles |
| 2.6 Microwave |  |  |  |  |  |  |  |
| Bergman et al 201037 | 10.63 | 4.667 | 18 | 10.45 | 3.333 | 18 | 2 min exposure, 03 cycles |
| Viscusi et al 200961 | 7.44 | 4.648 | 15 | 7.8 | 758 | 15 | 2 min exposure, 01 cycle |
| 2.7 Sodium hypochlorite |  |  |  |  |  |  |  |
| Bergman et al 201037 | 10.02 | 0.5 | 18 | 10.45 | 3.333 | 18 | 0.6%, 30 min submersion, 03 cycles |
| Viscusi et al 200961 | 8.0 | 8.567 | 18 | 7.52 | 6.567 | 18 | 0.6%, 30 min submersion, 01 cycle |
| 3. Filtration efficiency | Experimental |  |  | Control |  |  | Additional information (provided when necessary) |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 3.1 Plasma Vapor Hydrogen Peroxide (PVHP) |  |  |  |  |  |  |  |
| Cai and Floyd 202068 | 96.85 | 1.7 | 10 | 97.0 | 0.7 | 10 | PVHP |
| Cramer et al 202071 | 99.67 | 875 | 16 | 98.59 | 142 | 20 | SteraMist-BIT, 01 cycle The Gerson 2130 model was not considered in the analysis, as there was no reference (control) for this mask model. |
| Cramer et al 202071 | 99.33 | 0.03 | 8 | 98.37 | 1.267 | 12 | SteraMist-BIT, 05 cycles Only the models 3M 1860 and KC/Halyard 46767 were considered in the analysis, as they were the only ones tested in 05 cycles and with reference data (control) |
| Cramer et al 202071 | 99.16 | 1.233 | 12 | 98.3 | 0.15 | 16 | SteraMist-BIT, 10 cycles Only the models 3M 1860, KC/Halyard 46767 and 3M 8210 were considered in the analysis, as they were the only ones tested in 10 cycles and with reference data (control). |
| 3.2 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Liao et al 202042 | 95.0 | 1.59 | 3 | 96.52 | 1.37 | 30 | 254 nm, 8 W, 01 cycle |
| Liao et al 202042 | 96.78 | 0.8 | 3 | 96.52 | 1.37 | 30 | 254 nm, 8 W, 10 cycles |
| 3.3 Dry Heat |  |  |  |  |  |  |  |
| Daeschler et al 202027 | 97.2 | 3.61 | 18 | 96.2 | 1.57 | 12 | 70°C, 60 min, 05 cycles |
| Daeschler et al 202027 | 97.6 | 1.5 | 16 | 96.2 | 1.57 | 12 | 70°C, 60 min, 10 cycles |
| Liao et al 202042 | 96.67 | 0.65 | 3 | 96.52 | 1.37 | 30 | 75 °C, 30 min, 01 cycle |
| Liao et al 202042 | 97.25 | 0.34 | 3 | 96.52 | 1.37 | 30 | 75 °C, 30 min, 10 cycles |
| Pascoe et al 202052 | 98.19 | 0.08 | 3 | 98.35 | 0.1 | 3 | 70°C, 90 min, 01 cycle |
| Pascoe et al 202052 | 98.07 | 0.08 | 3 | 98.35 | 0.1 | 3 | 70°C, 90 min, 03 cycles |
| 3.4 Moist heat |  |  |  |  |  |  |  |
| Anderegg et al 202026 | 99.12 | 8.544 | 5 | 99.29 | 8.347 | 5 | 85°C, 60-85%RH, 05 cycles |
| Daeschler et al 202027 | 97.0 | 0.94 | 12 | 96.2 | 1.57 | 12 | 70°C, 50% RH, 05 cycles |
| Daeschler et al 202027 | 96.3 | 1.89 | 12 | 96.2 | 1.57 | 12 | 70°C, 50% RH, 10 cycles |
| Liao et al 202042 | 95.16 | 0.73 | 3 | 96.52 | 1.37 | 30 | 100°C, 10 min, 01 cycle |
| Liao et al 202042 | 80.65 | 2.97 | 3 | 96.52 | 1.37 | 30 | 100°C, 10 min, 10 cycles |
| 3.5 Ethanol |  |  |  |  |  |  |  |
| Grinshpun et al 202031 | 80.2 | 1.7 | 3 | 96.9 | 1.8 | 3 | soaking in 70% ethanol for 2 h, 01 cycle |
| Liao et al 202042 | 56.33 | 03.03 | 3 | 96.52 | 1.37 | 30 | 75%, immersion and air dry, 01 cycle |
| 3.6 Microwave |  |  |  |  |  |  |  |
| Pascoe et al 202052 | 98.28 | 0.06 | 3 | 98.35 | 0.1 | 3 | MGS, 90 sec, 03 cycles |
| Pascoe et al 202052 | 98.21 | 0.03 | 3 | 98.35 | 0.1 | 3 | MGS+oil, 90 sec, 01 cycle |
| Pascoe et al 202052 | 98.19 | 0.07 | 3 | 98.35 | 0.1 | 3 | MGS+oil, 90 sec, 03 cycles |
| Pascoe et al 202052 | 97.36 | 0.96 | 3 | 98.35 | 0.1 | 3 | MGS 90 sec, 01 cycle |
| 3.7 Sodium hypochlorite |  |  |  |  |  |  |  |
| Liao et al 202042 | 73.11 | 7.32 | 3 | 96.52 | 1.37 | 30 | chlorine-based (2%), light spray and air dry, 01 cycle |
| 3.8 Autoclave |  |  |  |  |  |  |  |
| Grinshpun et al 202031 | 95.5 | 1.1 | 3 | 96.9 | 1.8 | 3 | 121ºC, 15 psi, 30 min, 01 cycle |
| Grinshpun et al 202031 | 96.0 | 2.1 | 3 | 96.9 | 1.8 | 3 | 121ºC, 15 psi, 30 min, 05 cycles |
| Grinshpun et al 202031 | 96.0 | 1.9 | 3 | 96.9 | 1.8 | 3 | soiled, 121ºC, 15 psi, 30 min, 01 cycle |
| Grinshpun et al 202031 | 93.0 | 1.8 | 3 | 96.9 | 1.8 | 3 | soiled, 121ºC, 15 psi, 30 min, 05 cycles |
| Harskamp et al 202032 | 93.43 | 1.1333 | 14 | 97.17 | 3.333 | 8 | 17 min at 121°C with 34 min total cycle time, 01 cycle |
| Harskamp et al 202032 | 87.3 | 1.1667 | 16 | 97.17 | 3.333 | 8 | 17 min at 121°C with 34 min total cycle time, 02 cycles |
| Harskamp et al 202032 | 94.4 | 1.3 | 4 | 99.3 | 0.3 | 4 | 17 min at 121°C with 34 min total cycle time, 03 cycles Only the 3M1862 model was considered in the analysis, as it was the only one that maintained the characteristics after two test cycles. |
| 3.9 Chlorine dioxide |  |  |  |  |  |  |  |
| Cai and Floyd 202068 | 85.65 | 2.15 | 10 | 97.0 | 0.7 | 10 | ClO₂ |