Supplementary material 05

Table 4. Results assessing filter aerosol penetration, airflow resistance and filtration efficiency. The experimental group refers to the different methods included. The control group refers to masks not subjected to any disinfection process. Thus, the closer to the control value, the more advantageous the disinfection method.

| 1. Outcome: Filter aerosol penetration | Experimental |  |  | Control |  |  | Additional information |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 1.1 Liquid Hydrogen Peroxide |  |  |  |  |  |  |  |
| Bergman et al 201042 | 1.33 | 0.4617 | 18 | 1.16 | 0.245 | 18 | 3%, 30 min, 03 cycles |
| Viscusi et al 200765 | 0.75 | 0.242 | 4 | 0.7 | 0.267 | 10 | 3%, 30 min, 01 cycle |
| Viscusi et al 200765 | 0.71 | 0.148 | 4 | 0.7 | 0.267 | 10 | 6%, 30 min, 01 cycle |
| 1.2 Hydrogen Peroxide Vapor |  |  |  |  |  |  |  |
| Bergman et al 201042 | 0.97 | 0.17 | 36 | 1.16 | 0.245 | 18 | Clarus® R, 03 cycles |
| Bergman et al 201042 | 5.16 | 5.3133 | 36 | 1.16 | 2.45 | 18 | STERRAD® 100S, 03 cycles |
| Viscusi et al 200765 | 0.75 | 0.166 | 4 | 0.7 | 0.267 | 10 | STERRAD® 100S, 01 cycle |
| Viscusi et al 200765 | 0.5 | 0.213 | 4 | 0.7 | 0.267 | 10 | STERRAD® NX, 01 cycle |
| Viscusi et al 200966 | 0.7 | 0.4133 | 18 | 0.87 | 0.5083 | 18 | STERRAD® 100S, 01 cycle |
| 1.3 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Bergman et al 201042 | 0.92 | 0.2533 | 18 | 1.16 | 0.245 | 18 | 254 nm, 15 min, 03 cycles |
| Lindsley et al 201551 | 2.66 | 1.4435 | 16 | 2.38 | 1.2192 | 16 | 254 nm, 120 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 3.29 | 1.6252 | 16 | 2.72 | 1.5125 | 16 | 254 nm, 240 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 2.98 | 1.7514 | 16 | 2.33 | 1.3365 | 16 | 254 nm, 470 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 2.76 | 1.5111 | 16 | 2.29 | 1.523 | 16 | 254 nm, 950 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Viscusi et al 200765 | 0.57 | 0.152 | 4 | 0.7 | 0.267 | 10 | 254 nm, 30 min, 01 cycle |
| Viscusi et al 200765 | 0.79 | 0.082 | 4 | 0.7 | 0.267 | 10 | 254 nm, 480 min, 01 cycle |
| Viscusi et al 200966 | 0.93 | 0.4967 | 18 | 0.87 | 0.5083 | 18 | Sterilgard III, 01 cycle |
| 1.4 Ethylene Oxide |  |  |  |  |  |  |  |
| Bergman et al 201042 | 1.02 | 0.255 | 36 | 1.16 | 0.245 | 18 | Amsco® Eagle® 3017, 03 cycles |
| Viscusi et al 200765 | 1.29 | 0.35 | 4 | 0.7 | 0.267 | 10 | EtO 3M 4XL, 01 cycle |
| Viscusi et al 200765 | 1.2 | 0.729 | 4 | 0.7 | 0.267 | 10 | EtO 3M 5XL, 01 cycle |
| Viscusi et al 200966 | 0.78 | 0.455 | 18 | 0.87 | 0.5083 | 18 | Steri-Vac 5XL, 01 cycle |
| 1.5 Dry Heat |  |  |  |  |  |  |  |
| Viscusi et al 200765 | 0.84 | 0.258 | 4 | 0.7 | 0.267 | 10 | oven, 80ºC, 60 min, 01 cycle |
| 1.6 Moist Heat |  |  |  |  |  |  |  |
| Bergman et al 201042 | 0.97 | 0.2433 | 18 | 1.16 | 0.245 | 18 | 60ºC, 80% RH, 30 min, 03 cycles |
| 1.7 Isopropanol solution |  |  |  |  |  |  |  |
| Viscusi et al 200765 | 21.6 | 1.337 | 4 | 0.7 | 0.267 | 10 | 1 minute submersion, 01 cycle |
| Viscusi et al 200765 | 17.8 | 5.508 | 4 | 0.7 | 0.267 | 10 | 1 second submersion, 01 cycle |
| 1.8 Microwave |  |  |  |  |  |  |  |
| Bergman et al 201042 | 1.05 | 0.2567 | 18 | 1.16 | 0.245 | 18 | 2 min exposure, 03 cycles |
| Viscusi et al 200765 | 1.13 | 0.662 | 4 | 0.7 | 0.267 | 10 | 2 min exposure, 01 cycle |
| Viscusi et al 200765 | 1.77 | 0.617 | 4 | 0.7 | 0.267 | 10 | 4 min exposure, 01 cycle |
| Viscusi et al 200966 | 0.73 | 0.48 | 15 | 0.98 | 0.572 | 15 | 2 min exposure, 01 cycle |
| 1.9 Sodium hypochlorite |  |  |  |  |  |  |  |
| Bergman et al 201042 | 1.35 | 0.205 | 18 | 1.16 | 0.245 | 18 | 0.6%, 30 min sub, 03 cycles |
| Viscusi et al 200765 | 0.68 | 0.124 | 4 | 0.7 | 0.267 | 10 | 0.525% in tap water, 01 cycle |
| Viscusi et al 200765 | 3.79 | 3.446 | 4 | 0.7 | 0.267 | 10 | 5.25% in tap water, 01 cycle |
| Viscusi et al 200966 | 0.58 | 0.36 | 18 | 0.87 | 0.5083 | 18 | 0.6%, 30 min sub, 01 cycle |
| 1.10 Autoclave |  |  |  |  |  |  |  |
| Viscusi et al 200765 | 18.7 | 5.263 | 4 | 0.7 | 0.267 | 10 | 121°C (15 psi), 15 min, 01 cycle |
| Viscusi et al 200765 | 34.4 | 9.963 | 4 | 0.7 | 0.267 | 10 | 121°C (15 psi), 30 min, 01 cycle |
| 1.11 Bar soap and water |  |  |  |  |  |  |  |
| Viscusi et al 200765 | 38.8 | 1.269 | 4 | 0.7 | 0.267 | 10 | 02 min, 01 cycle |
| Viscusi et al 200765 | 34.9 | 4.599 | 4 | 0.7 | 0.267 | 10 | 20 min, 01 cycle |
| 2. Filter airflow resistance | Experimental |  |  | Control |  |  | Additional information |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 2.1 Liquid Hydrogen Peroxide |  |  |  |  |  |  |  |
| Bergman et al 201042 | 9.78 | 5.833 | 18 | 10.45 | 3.333 | 18 | 6%, 30 min, 03 cycles |
| 2.2 Hydrogen Peroxide Vapor |  |  |  |  |  |  |  |
| Bergman et al 201042 | 10.62 | 4.833 | 36 | 10.45 | 3.333 | 18 | Clarus® R, 03 cycles |
| Bergman et al 201042 | 10.58 | 3.833 | 36 | 10.45 | 3.333 | 18 | STERRAD® 100S, 03 cycles |
| Viscusi et al 200966 | 7.68 | 0.93 | 18 | 7.52 | 6.567 | 18 | STERRAD® 100S, 01 cycle |
| 2.3 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Bergman et al 2010 | 11.17 | 6.167 | 18 | 10.45 | 3.333 | 18 | 254 nm, 15 min, 03 cycles |
| Lindsley et al 201551 | 11.8 | 3.9553 | 16 | 11.83 | 4.1971 | 16 | 254 nm, 120 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 11.85 | 3.8176 | 16 | 11.87 | 3.8495 | 16 | 254 nm, 240 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 11.66 | 3.7878 | 16 | 11.66 | 3.9936 | 16 | 254 nm, 470 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Lindsley et al 201551 | 12.08 | 3.8758 | 16 | 12.08 | 3.8757 | 16 | 254 nm, 950 Jcmˉ², 01 cycle All models of N95 FFRs were considered together in the analysis. |
| Viscusi et al 200966 | 7.98 | 5.667 | 18 | 7.52 | 6.567 | 18 | Sterilgard III, 01 cycle |
| 2.4 Ethylene Oxide |  |  |  |  |  |  |  |
| Bergman et al 201042 | 11.47 | 3.667 | 36 | 10.45 | 3.333 | 18 | Amsco® Eagle® 3017, 03 cycles |
| Viscusi et al 200966 | 7.5 | 0.44 | 18 | 7.52 | 6.567 | 18 | Steri-Vac 5XL, 01 cycle |
| 2.5 Moist Heat |  |  |  |  |  |  |  |
| Bergman et al 201042 | 10.18 | 0.15 | 18 | 10.45 | 3.333 | 18 | 60ºC, 80% RH, 30 min, 03 cycles |
| 2.6 Microwave |  |  |  |  |  |  |  |
| Bergman et al 201042 | 10.63 | 4.667 | 18 | 10.45 | 3.333 | 18 | 2 min exposure, 03 cycles |
| Viscusi et al 200966 | 7.44 | 4.648 | 15 | 7.8 | 758 | 15 | 2 min exposure, 01 cycle |
| 2.7 Sodium hypochlorite |  |  |  |  |  |  |  |
| Bergman et al 201042 | 10.02 | 0.5 | 18 | 10.45 | 3.333 | 18 | 0.6%, 30 min submersion, 03 cycles |
| Viscusi et al 200966 | 8.0 | 8.567 | 18 | 7.52 | 6.567 | 18 | 0.6%, 30 min submersion, 01 cycle |
| 3. Filtration efficiency | Experimental |  |  | Control |  |  | Additional information (provided when necessary) |
| Study or subgroup | Mean | SD | Total | Mean | SD | Total |  |
| 3.1 Plasma Vapor Hydrogen Peroxide (PVHP) |  |  |  |  |  |  |  |
| Cai and Floyd 202073 | 96.85 | 1.7 | 10 | 97.0 | 0.7 | 10 | PVHP |
| Cramer et al 202076 | 99.67 | 875 | 16 | 98.59 | 142 | 20 | SteraMist-BIT, 01 cycle The Gerson 2130 model was not considered in the analysis, as there was no reference (control) for this mask model. |
| Cramer et al 202076 | 99.33 | 0.03 | 8 | 98.37 | 1.267 | 12 | SteraMist-BIT, 05 cycles Only the models 3M 1860 and KC/Halyard 46767 were considered in the analysis, as they were the only ones tested in 05 cycles and with reference data (control) |
| Cramer et al 202076 | 99.16 | 1.233 | 12 | 98.3 | 0.15 | 16 | SteraMist-BIT, 10 cycles Only the models 3M 1860, KC/Halyard 46767 and 3M 8210 were considered in the analysis, as they were the only ones tested in 10 cycles and with reference data (control). |
| 3.2 Ultraviolet Germicidal Irradiation |  |  |  |  |  |  |  |
| Liao et al 202047 | 95.0 | 1.59 | 3 | 96.52 | 1.37 | 30 | 254 nm, 8 W, 01 cycle |
| Liao et al 202047 | 96.78 | 0.8 | 3 | 96.52 | 1.37 | 30 | 254 nm, 8 W, 10 cycles |
| 3.3 Dry Heat |  |  |  |  |  |  |  |
| Daeschler et al 202032 | 97.2 | 3.61 | 18 | 96.2 | 1.57 | 12 | 70°C, 60 min, 05 cycles |
| Daeschler et al 202032 | 97.6 | 1.5 | 16 | 96.2 | 1.57 | 12 | 70°C, 60 min, 10 cycles |
| Liao et al 202047 | 96.67 | 0.65 | 3 | 96.52 | 1.37 | 30 | 75 °C, 30 min, 01 cycle |
| Liao et al 202047 | 97.25 | 0.34 | 3 | 96.52 | 1.37 | 30 | 75 °C, 30 min, 10 cycles |
| Pascoe et al 202057 | 98.19 | 0.08 | 3 | 98.35 | 0.1 | 3 | 70°C, 90 min, 01 cycle |
| Pascoe et al 202057 | 98.07 | 0.08 | 3 | 98.35 | 0.1 | 3 | 70°C, 90 min, 03 cycles |
| 3.4 Moist heat |  |  |  |  |  |  |  |
| Anderegg et al 202031 | 99.12 | 8.544 | 5 | 99.29 | 8.347 | 5 | 85°C, 60-85%RH, 05 cycles |
| Daeschler et al 202032 | 97.0 | 0.94 | 12 | 96.2 | 1.57 | 12 | 70°C, 50% RH, 05 cycles |
| Daeschler et al 202032 | 96.3 | 1.89 | 12 | 96.2 | 1.57 | 12 | 70°C, 50% RH, 10 cycles |
| Liao et al 202047 | 95.16 | 0.73 | 3 | 96.52 | 1.37 | 30 | 100°C, 10 min, 01 cycle |
| Liao et al 202047 | 80.65 | 2.97 | 3 | 96.52 | 1.37 | 30 | 100°C, 10 min, 10 cycles |
| 3.5 Ethanol |  |  |  |  |  |  |  |
| Grinshpun et al 202036 | 80.2 | 1.7 | 3 | 96.9 | 1.8 | 3 | soaking in 70% ethanol for 2 h, 01 cycle |
| Liao et al 202047 | 56.33 | 03.03 | 3 | 96.52 | 1.37 | 30 | 75%, immersion and air dry, 01 cycle |
| 3.6 Microwave |  |  |  |  |  |  |  |
| Pascoe et al 202057 | 98.28 | 0.06 | 3 | 98.35 | 0.1 | 3 | MGS, 90 sec, 03 cycles |
| Pascoe et al 202057 | 98.21 | 0.03 | 3 | 98.35 | 0.1 | 3 | MGS+oil, 90 sec, 01 cycle |
| Pascoe et al 202057 | 98.19 | 0.07 | 3 | 98.35 | 0.1 | 3 | MGS+oil, 90 sec, 03 cycles |
| Pascoe et al 202057 | 97.36 | 0.96 | 3 | 98.35 | 0.1 | 3 | MGS 90 sec, 01 cycle |
| 3.7 Sodium hypochlorite |  |  |  |  |  |  |  |
| Liao et al 202047 | 73.11 | 7.32 | 3 | 96.52 | 1.37 | 30 | chlorine-based (2%), light spray and air dry, 01 cycle |
| 3.8 Autoclave |  |  |  |  |  |  |  |
| Grinshpun et al 202036 | 95.5 | 1.1 | 3 | 96.9 | 1.8 | 3 | 121ºC, 15 psi, 30 min, 01 cycle |
| Grinshpun et al 202036 | 96.0 | 2.1 | 3 | 96.9 | 1.8 | 3 | 121ºC, 15 psi, 30 min, 05 cycles |
| Grinshpun et al 202036 | 96.0 | 1.9 | 3 | 96.9 | 1.8 | 3 | soiled, 121ºC, 15 psi, 30 min, 01 cycle |
| Grinshpun et al 202036 | 93.0 | 1.8 | 3 | 96.9 | 1.8 | 3 | soiled, 121ºC, 15 psi, 30 min, 05 cycles |
| Harskamp et al 202037 | 93.43 | 1.1333 | 14 | 97.17 | 3.333 | 8 | 17 min at 121°C with 34 min total cycle time, 01 cycle |
| Harskamp et al 202037 | 87.3 | 1.1667 | 16 | 97.17 | 3.333 | 8 | 17 min at 121°C with 34 min total cycle time, 02 cycles |
| Harskamp et al 202037 | 94.4 | 1.3 | 4 | 99.3 | 0.3 | 4 | 17 min at 121°C with 34 min total cycle time, 03 cycles Only the 3M1862 model was considered in the analysis, as it was the only one that maintained the characteristics after two test cycles. |
| 3.9 Chlorine dioxide |  |  |  |  |  |  |  |
| Cai and Floyd 202073 | 85.65 | 2.15 | 10 | 97.0 | 0.7 | 10 | ClO₂ |