Visual-Servo Micromanipulation Platform for Patterned and Complex Organoid Biofabrication

Authorship: Mingsi Tong1, #, Gang Huang1, #, Xinghu Yu2, #, Weiyang Lin1, Yuchen Chen1, Yang Wu1, Shijia Cheng1, Songlin Zhuang2, *, Meng Jiang1, *, Huijun Gao1, *

# Equal contribution

*Corresponding authors

Affiliations

1 Research Institute of Intelligent Control and Systems, Harbin Institute of Technology, Harbin, 150001, China

2 Yongjiang Laboratory, Ningbo, 315202, China

Correspondence should be addressed to:

Research Institute of Intelligent Control and Systems, Harbin Institute of Technology, Harbin, 150001, China

Key words:

biofabrication, organoids, micromanipulation, bioengineering, visual servo

1. Cell density statistics in the pipette

To analyze the density of cells in the pipette, a region of 150 pixels by 100 pixels is taken as the regions of interest (ROI). We took the first image when there were no cells as the background template. This region is mean-filtered to avoid interference. The cell density was detected in real time and was obtained by differencing the density from the background and then segmenting it by threshold. The area of the obtained binary image was considered the cell cluster area and compared with the area of the ROI as the final density statistical value.

2. Automated organoid manufacturing platform

The hardware setup of the proposed system (Supplementary Figures 3a-3b) mainly includes the establishment of feedback information, motion control for the micromanipulator and the motorized stage, pressure control for the cell extrusion, and temperature and CO2 control for providing a suitable cultivation environment. As the controller, the host computer is responsible for the regulation of the whole system operation. The host computer takes images from the microscope (as the sensor) and analyzes them to update the system status in real time. The obtained system status is fed into the individual algorithms, and the output of the algorithm is transmitted to the corresponding slave computer controller and actuator. For example, the micromanipulator controller will obtain a set of desired 3D coordinates. The position of the micromanipulator is controlled according to the desired value, and the current actual coordinates are fed to the host computer. The motorized stage controller obtains a set of desired 2D coordinates, where it controls and feeds back the position of the motorized stage. The motorized oil pump controller obtains information about the speed and direction of the motor. The speed of the oil pump is controlled according to the desired value, and cell movement is further controlled. The temperature control unit and carbon dioxide control unit obtain the desired temperature and carbon dioxide concentration, respectively. They maintain the system environment according to the expected value combined with the feedback information of the temperature sensor and carbon dioxide sensor. The temperature was generally 4°C during the operation, while it was 37°C during the observation. The carbon dioxide concentration was maintained at 5%.

The process of automated cell seeding is illustrated in the Supplementary Figure 3c. Once the equipment is prepared, for first-time use, both the motion stage and the micromanipulator require precalibration. After this initial calibration, the pipette, which served as the micromanipulator's end effector, is positioned within the field of view. Real-time tracking of the pipette tip's position through imaging can mitigate the impact of end position errors on motion precision. Then, the pipette is moved into the reservoir dish, after which the cells are sequentially aspirated. When the cell loading is completed, the pipette is moved to the incubation dish, where the microwells array for organoid culture is located. The process of cell seeding in microwells array mainly includes microwell location, contact detection and cell distribution control.

- The establishment of three-dimensional spatial coordinates of the system

The precision of cell placement and the formation of specific patterns are critical factors in organoid self-organization. In general, micromanipulation systems that can achieve high throughput, high precision, and a large travel range simultaneously are difficult. The open-loop precalibration method is used for mapping multiple coordinate systems. However it cannot eliminate errors, which may be amplified during large travel movements, thus impacting the system's accuracy. Visual servo control offers an effective solution to this problem.

System motion calibration. To control the cell placement and distribution, we need to establish the 3D coordinates of the end effector with respect to the microwells array. Due to the motion displacement limitation of the micromanipulator, planar motion is mainly realized by the motorized stage. The purpose of planar motion is mainly to ensure positioning accuracy when moving the micromanipulator from one microwell to the next. The motorized stage is precalibrated to obtain the transformation matrix with respect to the microwells coordinate system  (Supplementary Figure 5a). Because the microwells array is placed in a round dish, it is difficult to ensure that the angle is fixed. To avoid multipoint calibration each time, we obtained  by calibrating only the rotation angle of the microwell relative to the microscope image coordinate according to the coordinate transformation of the fixed microscope with respect to the motorized stage . Then, we calculate the transformation matrix with respect to the microwells coordinate system .

For the coordinate transformation of the microscope with respect to the motorized stage , the motorized stage is moved beyond four points within the microscope's field of view. During this movement, we record the coordinates in both the motorized stage and the image. The coordinate transformation matrix is then calculated using the least squares method. This parameter needs to be calibrated only once after installing the motorized stage.

For the coordinate transformation of the microwell with respect to the microscope, the rotation change matrix is first obtained (Supplementary Figure 5a), and its declension with respect to the image can be measured automatically by the program . The scaling ratio transformation between the microwell and the microscope coordinate system can be obtained directly according to the transformation relationship between the motorized stage and the microscope. The current microwell position is , and the target position is . This value can be obtained from the previously designed parameters. Then, the motion instructions for the motorized stage can be calculated according to the following equation:

Visual servo control for planar motion. To further improve the accuracy, we use precalibration combined with visual feedback. This process relies primarily on three-dimensional visual positioning and motion controller design (Supplementary Figure 4c).With this method, the distance between the end effector and the microwells can be robustly guaranteed during each seeding.

For visual localization, to mitigate interference from complex backgrounds and significant changes in object features in relief phase contrast imaging (RP) mode, we trained a neural network to recognize and locate target objects (Supplementary Figure 4d). Microwells array positioning is essential for avoiding disturbing the flow field due to pipette misplacement, which consequently affects cell distribution. To achieve this goal, we collected image stacks of microwells under various conditions and constructed datasets. These include microwells under various polarization angles and lighting conditions in RPs, datasets of arrays with the pipette tip at different heights in the field of view, and datasets of with different cell distributions in the microwells, among others. After labeling the data, we trained an object detection network to obtain a model capable of accurately recognizing and locating microwells under diverse environments and conditions.

We used an efficient real-time object detector (RTMDet) and evaluated the proposed model's performance on a validation dataset. The model with variant sizes and the template matching method were compared (Supplementary Figure 5e). These metrics include the IoU mAP and algorithm runtime. RTMDet with a large model (RTMDet_L) achieved a 0.96 mAP (runtime= 1.3 s, IoU = 0.95), while RTMDet with a small model (RTMDet_S) provided a 0.93 mAP (runtime= 0.4 s, IoU = 0.95). Although traditional visual detection algorithms can achieve faster runtimes (0.04 s), their detection performances cannot meet these requirements (IoU = 0.24). The performance of RTMDet_S is not as good as that of RTMDet_L, but it achieves a better balance between detection speed and performance. To compensate for the impact of visual detection failure, we employed a multistep detection strategy to re-detect the results adjusted by visual feedback at each step and compared the distance difference between the two detection results and the target point. If this value decreases continuously, the detection results are regarded as effective; otherwise, the position obtained by precalibration calculation is used.

We further evaluated the displacement error of our proposed method for microwells array locations. A typical example is shown in Supplementary Figure 4g. For coarse positioning (open-loop precalibration method), the error between the end effector and the microwell center gradually increases with the number of rows and columns. The deviation of the open-loop control method mainly occurs because of the angular error between the microwells array coordinates and the microscope coordinates, which increases with distance. For the visual-servo control method, this amplified deviation can be immediately compensated without accumulation. In a large travel range, a consistent deviation can be maintained for each microwell within a reasonable range, making this method significantly superior to the precalibration method. The hybrid visual-servo method can reduce the average error, while only 1 second seconds longer per microwell than does the open-loop control method.

Visual servo control for vertical motion. For the acquisition of the Z-axis depth, under a bright-field microscope, the general method is based on blur similarity matching. The accuracy of this method is determined by the depth of the microscope field; thus, there is a large depth error for low-magnification lenses. In our system, it is more important to obtain the height relative to the microwells array. We designed a contact detection algorithm to obtain the micropipette height relative to the microwells height (Supplementary Figure 4h). We first positioned and moved the microwell to the center of the field of view (FOV). The micromanipulator is gradually lowered, and the pipette deformation is recorded simultaneously. The algorithm can judge the contact moment between the tip and the microwell through the deformation. We set the contact position to zero, and the depth can be calculated in real time according to the physical model and the position information of the pipette tip. Once the pipette tip is in the appropriate position, the extrusion of cells is controlled by the movement of the oil pump. When the cells reach the target distribution, the pipette moves to the next position, and the system starts a new round of seeding.

- Construction of the observer for different constraint strategies

For different constrained control strategies, the corresponding observer implementation algorithm is as follows (Supplementary Figure 5a):

For 1D control, the cell density is constrained. First, the cell intensity is calculated in the defined region of interest (ROI). The results are subsequently compared with the desired values to provide feedback information. Then, it serves as the input for the controller, ensuring that the controlled density is close to the desired number.

For 2D control, both cell densities and boundaries are constrained simultaneously. The method of constraining the cell density is similar to that used for 1D control. However, the ROI used for cell density calculation dynamically changes. Moreover, to realize boundary control, we detect the boundary of the cell cluster after the binarization process in the boundary ROI. To avoid interference from individual discrete cells, we also calculate the contour circle with 95% of the intensity region and input it into the controller as the actual boundary. The cell density constraint is accompanied by the boundary constraint.

For 3D control, the cell spatial distribution is controlled. This includes constraints on the cell density as well as the boundary. Based on the obtained 3D intensity values of the cell distribution, the polar coordinate system is established with the center of the intensity distribution as the origin. The accuracy of the observations is determined by subdividing the angles. Here, we take two-subdivision as an example, where each concentric circle is discretized into semiarc regions. By counting the density of cells in each region, we project the intensity values of cells with various radii to a single point. This process results in a one-dimensional curve that describes the distribution. By comparing the intensity values layer by layer in each concentric circle with the desired value, we can effectively implement constraints on the cell distribution.

- Modeling of microfluidic movement in pipette

During cell aspiration and dispensing (Supplementary Figure 5b), the change in volume through the pump movement is denoted by:

where  is the initial volume of the oil segment. is the displacement of the culture medium column inside the pipette.  denotes the sectional area of the pipette.  is the displacement of the syringe piston. is the sectional area of the oil pump piston. Based on Boyle’s law:

where  is the pressure of the oil after the displacement of the pump.  is the initial pressure, which is approximately equal to the atmospheric pressure. Dividing the resultant equation by  results in:

where  is the velocity of the culture medium inside the micropipette. As , the simplified equation can be obtained:

The flow velocity in the pipette is relatively slow and the compressibility of the liquid is negligible. When the oil pump pressure is applied to the culture medium, the fluid motion can be calculated based on the Hagen-Poiseuille equation

where  is the volumetric flow rate.  is the medium pressure gradient in the pipette.  is the inside radius of the pipette.  is the length of the pipette.  is the length of the culture medium.  and  are the dynamic viscosity of the culture medium and the oil, respectively. Substituting the equation above and dividing the resultant equation by  results in:

In general, the liquid pressure is equal to the atmospheric pressure. Then,

The dynamics of fluid motion can be rewritten as

From the aforementioned equation, it is clear that at the microscale, the piston displacement and the liquid displacement within the pipette do not maintain a linear relationship. The compression of the liquid and the elasticity of the tube introduce a nonlinear factor. Consequently, precise control of the liquid droplet cannot rely on open-loop control, wherein the movement of the piston in the pump is used to infer the movement in the pipette. Instead, feedback control of the liquid position should be introduced to achieve precise movement control.

- Microfluidic controller design

Assuming that the cell movement in the pipette tip is nearly consistent with fluid flow, fluid feedback can be achieved through tracking cell movement. Based on the microfluidic motion model, we designed a corresponding controller (Supplementary Figure 5c):

The system state equation is modified in the following form:

where .  is a coefficient to be calibrated experimentally. represents the bounded uncertain nonlinearity.

Sliding mode controller design. Let ,where  is the desired position. The sliding surface is designed based on e:

Designing the controller

where . Construct the Lyapunov function . Dividing the resultant equation by  results in:

When ,. According to Lasalle's invariant set principle, the closed-loop system is asymptotically stable, where . To reduce the chattering of the controller during control, we use the saturation function instead of the sign function.

In the actual situation, there are several cases of parameter perturbation. A general control scheme suffers from fluctuations. Therefore, an adaptive sliding mode controller was designed to estimate the system parameters. The dynamic fluid motion is rewritten as follows:

where

Let , where  is the desired position. The sliding surface is designed based on e:

Designing the controller

where . Construct the Lyapunov function . Dividing the resultant equation by  results in:

where . To achieve accurate parameter estimation, we define the Lyapunov function as follows:

Furthermore, we obtain the following function:

Design the adaptive scheme:

Then . When , . According to Lasalle's invariant set principle, the closed-loop system is asymptotically stable, where .

PD Controller design. The state equations for the cell density control are:

Let

Through the Laplace transform

For the error value

Microfluidic control is fundamental for achieving quantitative cell seeding. The cell distribution in the pipette was uneven, showing consistent trends but random local values. Therefore, based on microfluidic feedback, we further achieved closed-loop control via real-time monitoring of cell density and the specific control scheme to regulate the cell spatial distribution curve.

To compare the performance of the proposed method against fluid disturbance under the same conditions, a simulation experiment was conducted. When the dynamic characteristics and parameters of the system model was fixed (Supplementary Figures 6a-b), all these control schemes can achieve the desired accuracy (for which the steady-state positioning errors are less than 5%). The PD controller may overshoot to achieve a faster response, which will cause the cell to be resucked into the pipette and exert additional flow shear force. When considering the external fluid disturbance and the model perturbation caused by the cell density change (Supplementary Figures 6c-d), the variable structure robust control had better performance in terms of efficiency and accuracy requirements.

- Cell distribution curve evaluation

To compare the cell distribution curves obtained under different control strategies, the deviations from the ideal curve and the self-consistency were measured. For the error measurement of the ideal curve, we use the root mean squared error (RMSE), which is calculated as follows：

To evaluate the consistency of the distribution curves, we first fit a mean curve  according to multiple experimental curves and subsequently compared each curve to calculate the corresponding error via the RMSE method.