## Validation_ICP-MS
ICP-MS
Reference:
V. Barwick (Ed), Eurachem/CITAC Guide: Guide to Quality in Analytical Chemistry: An Aid to Accreditation (3rd ed. 2016). ISBN 978-0-948926-32-7. Available from www.eurachem.org.
Due to complete combustion of organic material in the plasma, validation of ClickZip tags was performed with Lu(NO3)3, Yb(NO3)3
and Tm(NO3)3 instead of the corresponding HO2CPh{Lu}, HO2CPh{Yb} and HO2CPh{Tm} (with Y(NO3)3 as an internal standard).
limit of detection (LOD) and limit of quantification (LOQ):
Pentaplicate solutions of Lu3+, Yb3+ and Tm3+ (all in form of nitrate) in concentration range of 0.5 and 1.0 ng × L−1 were prepared by dilution of 1 ng × L−1 stock solutions in a matrix blank (2% HNO3) with Y3+ (in form of nitrate) internal standard solution.
raw data | [7 empty] | calculations Lu | [6 empty] | calculations Yb | [6 empty] | calculations Tm
c, ng × L−1 | 175Lu, cps | 174Yb, cps | 169Tm, cps | 89Y, cps |  |  |  | Lu/Y | (ci − cav)2 | [5 empty] | Yb/Y | (ci − cav)2 | [5 empty] | Tm/Y | (ci − cav)2
0 | 0 | 3.1155 | 0.8 | 37543.679 |  |  |  | 0 | 0.25 | [5 empty] | 8.2985e-05 | 0.25 | [5 empty] | 2.1309e-05 | 0.25
0 | 0 | 0.71085 | 0 | 39209.8061 |  |  |  | 0 | 0.25 | [5 empty] | 1.8129e-05 | 0.25 | [5 empty] | 0 | 0.25
0 | 0.8 | 3.8733 | 0.8 | 38950.7299 | LOD and LOQ Lu: |  |  | 2.0539e-05 | 0.25 | [5 empty] | 9.9441e-05 | 0.25 | [5 empty] | 2.0539e-05 | 0.25
0 | 0 | -0.10792 | 0 | 38883.2921 | LOD = | 0.14881 | ng × L−1 | 0 | 0.25 |  | parameters Lu |  |  |  | -2.7754e-06 | 0.25 |  |  | parameters Yb |  |  | 0 | 0.25 | parameters Tm
0 | 0 | 3.0733 | 0 | 38202.3054 | LOQ = | 0.44643 | ng × L−1 | 0 | 0.25 |  | b0 (intercept) |  |  | 3.2998e-05 | 8.0448e-05 | 0.25 |  | b0 (intercept) |  | 6.3398e-05 |  | 0 | 0.25 | b0 (intercept) |  | 1.1681e-05
0.5 | 28 | 13.5015 | 32.8 | 38990.7963 |  |  |  | 0.00071812 | 0 |  | b1 (slope) |  |  | 0.0012981 | 0.00034627 | 0 |  | b1 (slope) |  | 0.00040437 |  | 0.00084123 | 0 | b1 (slope) |  | 0.0014671
0.5 | 24.8 | 9.4921 | 32 | 39012.4673 | LOD and LOQ Yb: |  |  | 0.00063569 | 0 |  | syx (residual std. deviation) |  |  | 0.00010098 | 0.00024331 | 0 |  | syx (residual std. deviation) |  | 6.0354e-05 |  | 0.00082025 | 0 | syx (residual std. deviation) |  | 9.2187e-05
0.5 | 33.6 | 11.8968 | 24.8 | 39032.4624 | LOD = | 0.2855 | ng × L−1 | 0.00086082 | 0 |  | n (number of points) |  |  | 15 | 0.00030479 | 0 |  | n (number of points) |  | 15 |  | 0.00063537 | 0 | n (number of points) |  | 15
0.5 | 28.8 | 9.4921 | 32.8 | 39027.7243 | LOQ = | 0.8565 | ng × L−1 | 0.00073794 | 0 |  | cav (concentration average) |  |  | 0.5 | 0.00024321 | 0 |  | cav (concentration average) |  | 0.5 |  | 0.00084043 | 0 | cav (concentration average) |  | 0.5
0.5 | 28.8 | 10.3343 | 24 | 38582.458 |  |  |  | 0.00074645 | 0 |  | t1−α;(n−2) (critical t-test value) |  |  | 1.7709 | 0.00026785 | 0 |  | t1−α;(n−2) (critical t-test value) |  | 1.7709 |  | 0.00062204 | 0 | t1−α;(n−2) (critical t-test value) |  | 1.7709
1 | 40.8001 | 16.7578 | 52.8001 | 37418.5579 | LOD and LOQ Tm: |  |  | 0.0010904 | 0.25 |  | Σ (ci − cav)2 |  |  | 2.5 | 0.00044785 | 0.25 |  | Σ (ci − cav)2 |  | 2.5 |  | 0.0014111 | 0.25 | Σ (ci − cav)2 |  | 2.5
1 | 48.8001 | 23.1437 | 58.4001 | 38521.4983 | LOD = | 0.12019 | ng × L−1 | 0.0012668 | 0.25 | [5 empty] | 0.0006008 | 0.25 | [5 empty] | 0.001516 | 0.25
1 | 48.0001 | 16.7625 | 52.0001 | 37306.2546 | LOQ = | 0.36058 | ng × L−1 | 0.0012866 | 0.25 | [5 empty] | 0.00044932 | 0.25 | [5 empty] | 0.0013939 | 0.25
1 | 56.8001 | 14.339 | 54.4001 | 39260.2988 |  |  |  | 0.0014468 | 0.25 | [5 empty] | 0.00036523 | 0.25 | [5 empty] | 0.0013856 | 0.25
1 | 54.4001 | 16.7343 | 64.0001 | 38304.088 |  |  |  | 0.0014202 | 0.25 | [5 empty] | 0.00043688 | 0.25 | [5 empty] | 0.0016708 | 0.25
repeatability (precision):
5; 10; 50; 100; and 500 ng × L−1 pentaplicate solutions of Lu3+, Yb3+ and Tm3+ (all in form of nitrate) were prepared by stepwise dilution of 1 ng × L−1 stock solutions in a matrix blank (2% HNO3) with Y3+ (in form of nitrate) internal standard solution.
raw data | [7 empty] | calculations Lu |  |  |  | calculations Yb |  |  | calculations Tm
c, ng × L−1 | 175Lu, cps | 174Yb, cps | 169Tm, cps | 89Y, cps |  |  |  | Lu/Y | c, ng × L−1 | std. dev. |  | Yb/Y | c, ng × L−1 | std. dev. | Tm/Y | c, ng × L−1 | std. dev.
5 | 271.2026 | 89.544 | 284.0028 | 40528.6199 |  |  |  | 0.0066916 | 5 | 0.00049814 |  | 0.0022094 | 5 | 0.000178 | 0.0070075 | 5 | 0.00036282
5 | 226.4018 | 79.9298 | 260.8024 | 40489.3015 | repeatability Lu: |  |  | 0.0055916 |  |  |  | 0.0019741 |  |  | 0.0064413
5 | 283.2028 | 89.5768 | 289.603 | 41832.4799 | c, ng × L−1 | rel. std. dev. | repeatable: | 0.0067699 |  |  |  | 0.0021413 |  |  | 0.0069229
5 | 248.8022 | 75.1204 | 264.0025 | 41325.3986 | 5 | 7.889% | yes | 0.0060206 |  |  |  | 0.0018178 |  |  | 0.0063884
5 | 266.4025 | 91.9112 | 252.8023 | 40991.596 | 10 | 3.162% | yes | 0.006499 |  |  |  | 0.0022422 |  |  | 0.0061672
10 | 481.6082 | 162.3352 | 527.2098 | 38188.5918 | 50 | 2.091% | yes | 0.012611 | 10 | 0.00041469 |  | 0.0042509 | 10 | 0.00028213 | 0.013805 | 10 | 0.00034731
10 | 528.0099 | 178.3589 | 555.2093 | 38653.0416 | 100 | 2.245% | yes | 0.01366 |  |  |  | 0.0046144 |  |  | 0.014364
10 | 502.4088 | 182.3308 | 520.8095 | 38659.4861 | 500 | 1.622% | yes | 0.012996 |  |  |  | 0.0047163 |  |  | 0.013472
10 | 510.4091 | 191.945 | 524.8097 | 38104.3918 |  |  |  | 0.013395 |  |  |  | 0.0050373 |  |  | 0.013773
10 | 501.6089 | 178.3683 | 527.2098 | 38870.4596 | repeatability Yb |  |  | 0.012905 |  |  |  | 0.0045888 |  |  | 0.013563
50 | 2567.4311 | 865.5841 | 2695.4546 | 38475.7538 | c, ng × L−1 | rel. std. dev. | repeatable: | 0.066729 | 50 | 0.0013719 |  | 0.022497 | 50 | 0.00050125 | 0.070056 | 50 | 0.0014366
50 | 2585.8344 | 845.5735 | 2665.0487 | 38556.7769 | 5 | 8.57% | yes | 0.067066 |  |  |  | 0.021931 |  |  | 0.06912
50 | 2487.417 | 877.6039 | 2669.0509 | 38964.3147 | 10 | 6.078% | yes | 0.063838 |  |  |  | 0.022523 |  |  | 0.0685
50 | 2449.01 | 830.3679 | 2663.4492 | 37895.8279 | 50 | 2.275% | yes | 0.064625 |  |  |  | 0.021912 |  |  | 0.070283
50 | 2523.4229 | 816.8001 | 2557.0295 | 38331.3696 | 100 | 2.141% | yes | 0.065832 |  |  |  | 0.021309 |  |  | 0.066709
100 | 5068.9006 | 1736.0681 | 5237.7606 | 37638.3323 | 500 | 1.917% | yes | 0.13467 | 100 | 0.0029894 |  | 0.046125 | 100 | 0.0009676 | 0.13916 | 100 | 0.0023495
100 | 5124.9229 | 1691.2585 | 5309.7871 | 38136.5272 |  |  |  | 0.13438 |  |  |  | 0.044347 |  |  | 0.13923
100 | 4943.2562 | 1663.2593 | 5115.3169 | 37434.6138 | repeatability Tm |  |  | 0.13205 |  |  |  | 0.044431 |  |  | 0.13665
100 | 4908.045 | 1771.2724 | 5257.8047 | 38205.4218 | c, ng × L−1 | rel. std. dev. | repeatable: | 0.12846 |  |  |  | 0.046362 |  |  | 0.13762
100 | 5109.7153 | 1679.233 | 5361.8087 | 37542.069 | 5 | 5.509% | yes | 0.13611 |  |  |  | 0.044729 |  |  | 0.14282
500 | 25228.8827 | 8557.414 | 26204.9197 | 37526.0345 | 10 | 2.518% | yes | 0.6723 | 500 | 0.010619 |  | 0.22804 | 500 | 0.0042377 | 0.69831 | 500 | 0.0063763
500 | 24590.7918 | 8297.5691 | 25566.8996 | 37449.8806 | 50 | 2.084% | yes | 0.65663 |  |  |  | 0.22156 |  |  | 0.6827
500 | 24046.6336 | 8070.219 | 25397.7692 | 37201.9843 | 100 | 1.689% | yes | 0.64638 |  |  |  | 0.21693 |  |  | 0.6827
500 | 24264.5969 | 8231.8972 | 25809.7109 | 37511.6388 | 500 | 0.9268% | yes | 0.64686 |  |  |  | 0.21945 |  |  | 0.68805
500 | 24374.4073 | 8195.9103 | 25723.9618 | 37378.5282 |  |  |  | 0.6521 |  |  |  | 0.21927 |  |  | 0.6882
linearity (0.5‒500 ng × L−1 range):
Duplicate solutions of Lu3+, Yb3+ and Tm3+ (all in form of nitrate) in concentration range of 0.5‒500 ng × L−1 were prepared by stepwise dilution of 1 g × L−1 stock solutions in a matrix blank (2% HNO3) with Y3+ (in form of nitrate) internal standard solution.
raw data | [7 empty] | calculations Lu | [10 empty] | calculations Yb | [8 empty] | calculations Tm
c, ng × L−1 | 175Lu, cps | 174Yb, cps | 169Tm, cps | 89Y, cps |  |  |  | c, ng × L−1 | Lu/Y | yi | Yreg,i | I × (yi − yreg,i)2 |  | yij − yi |  |  |  |  | c, ng × L−1 | Yb/Y | yi | Yreg,i | I × (yi − yreg,i)2 | yij − yi |  |  |  | c, ng × L−1 | Tm/Y | yi | Yreg,i | I × (yi − yreg,i)2 | yij − yi
0.5 | 28 | 13.5015 | 32.8 | 38990.7963 |  |  |  | 0.5 | 0.00071812 | 0.00067691 | 0.00082087 | 4.1453e-08 |  | 4.1212e-05 |  |  |  |  | 0.5 | 0.00034627 | 0.00029479 | 0.0002063 | 1.5662e-08 | 5.1482e-05 |  |  |  | 0.5 | 0.00084123 | 0.00083074 | 0.00099138 | 5.161e-08 | 1.0487e-05
0.5 | 24.8 | 9.4921 | 32 | 39012.4673 |  |  |  |  | 0.00063569 |  |  |  |  | -4.1212e-05 | parameters Lu |  |  |  |  | 0.00024331 |  |  |  | -5.1482e-05 | parameters Yb |  |  |  | 0.00082025 |  |  |  | -1.0487e-05 | parameters Tm
1 | 40.8001 | 16.7578 | 52.8001 | 37418.5579 | linearity Lu: |  |  | 1 | 0.0010904 | 0.0011786 | 0.0014855 | 1.8836e-07 |  | -8.8229e-05 | b0 (intercept) |  |  | 0.00015626 | 1 | 0.00044785 | 0.00052432 | 0.00043115 | 1.7363e-08 | -7.6477e-05 | b0 (intercept) |  | -1.8553e-05 | 1 | 0.0014111 | 0.0014636 | 0.0016818 | 9.5232e-08 | -5.2486e-05 | b0 (intercept) | 0.00030099
1 | 48.8001 | 23.1437 | 58.4001 | 38521.4983 | F = | 0.065909 | linear: |  | 0.0012668 |  |  |  |  | 8.8229e-05 | b1 (slope) |  |  | 0.0013292 |  | 0.0006008 |  |  |  | 7.6477e-05 | b1 (slope) |  | 0.0004497 |  | 0.001516 |  |  |  | 5.2486e-05 | b1 (slope) | 0.0013808
5 | 271.2026 | 89.544 | 284.0028 | 40528.6199 | Fcrit = | 3.9715 | yes | 5 | 0.0066916 | 0.0061416 | 0.0068024 | 8.7325e-07 |  | 0.00054999 | n (number of points) |  |  | 14 | 5 | 0.0022094 | 0.0020917 | 0.0022299 | 3.8198e-08 | 0.00011765 | n (number of points) |  | 14 | 5 | 0.0070075 | 0.0067244 | 0.0072049 | 4.6174e-07 | 0.0002831 | n (number of points) | 14
5 | 226.4018 | 79.9298 | 260.8024 | 40489.3015 |  |  |  |  | 0.0055916 |  |  |  |  | -0.00054999 | I (number of repetitions) |  |  | 2 |  | 0.0019741 |  |  |  | -0.00011765 | I (number of repetitions) |  | 2 |  | 0.0064413 |  |  |  | -0.0002831 | I (number of repetitions) | 2
10 | 481.6082 | 162.3352 | 527.2098 | 38188.5918 | linearity Yb: |  |  | 10 | 0.012611 | 0.013136 | 0.013449 | 1.9568e-07 |  | -0.00052447 | N (number of calibration points) |  |  | 7 | 10 | 0.0042509 | 0.0044326 | 0.0044785 | 4.2012e-09 | -0.00018174 | N (number of calibration points) |  | 7 | 10 | 0.013805 | 0.014085 | 0.014109 | 1.1559e-09 | -0.00027925 | N (number of calibration points) | 7
10 | 528.0099 | 178.3589 | 555.2093 | 38653.0416 | F = | 0.022529 | linear: |  | 0.01366 |  |  |  |  | 0.00052447 | Σ I × (yi − Yreg,i)2 |  |  | 5.8409e-06 |  | 0.0046144 |  |  |  | 0.00018174 | Σ I × (yi − Yreg,i)2 |  | 3.6709e-07 |  | 0.014364 |  |  |  | 0.00027925 | Σ I × (yi − Yreg,i)2 | 2.136e-06
50 | 2567.4311 | 865.5841 | 2695.4546 | 38475.7538 | Fcrit = | 3.9715 | yes | 50 | 0.066729 | 0.066897 | 0.066618 | 1.5597e-07 |  | -0.00016854 | Σ Σ (yij − yi)2 |  |  | 0.00012407 | 50 | 0.022497 | 0.022214 | 0.022466 | 1.2775e-07 | 0.00028313 | Σ Σ (yij − yi)2 |  | 2.2812e-05 | 50 | 0.070056 | 0.069588 | 0.06934 | 1.2341e-07 | 0.00046791 | Σ Σ (yij − yi)2 | 0.0001227
50 | 2585.8344 | 845.5735 | 2665.0487 | 38556.7769 |  |  |  |  | 0.067066 |  |  |  |  | 0.00016854 | F1−α;(N−2,n−N) (critical f-test value) |  |  | 3.9715 |  | 0.021931 |  |  |  | -0.00028313 | F1−α;(N−2,n−N) (critical f-test value) |  | 3.9715 |  | 0.06912 |  |  |  | -0.00046791 | F1−α;(N−2,n−N) (critical f-test value) | 3.9715
100 | 5068.9006 | 1736.0681 | 5237.7606 | 37638.3323 | linearity Tm: |  |  | 100 | 0.13467 | 0.13453 | 0.13308 | 4.2012e-06 |  | 0.00014516 |  |  |  |  | 100 | 0.046125 | 0.045236 | 0.044951 | 1.6216e-07 | 0.00088876 |  |  |  | 100 | 0.13916 | 0.1392 | 0.13838 | 1.3363e-06 | -3.5368e-05
100 | 5124.9229 | 1691.2585 | 5309.7871 | 38136.5272 | F = | 0.024372 | linear: |  | 0.13438 |  |  |  |  | -0.00014516 | [5 empty] | 0.044347 |  |  |  | -0.00088876 |  |  |  |  | 0.13923 |  |  |  | 3.5368e-05
500 | 25228.8827 | 8557.414 | 26204.9197 | 37526.0345 | Fcrit = | 3.9715 | yes | 500 | 0.6723 | 0.66447 | 0.66477 | 1.8504e-07 |  | 0.0078357 |  |  |  |  | 500 | 0.22804 | 0.2248 | 0.22483 | 1.7585e-09 | 0.0032374 |  |  |  | 500 | 0.69831 | 0.6905 | 0.69069 | 6.6569e-08 | 0.0078083
500 | 24590.7918 | 8297.5691 | 25566.8996 | 37449.8806 |  |  |  |  | 0.65663 |  |  |  |  | -0.0078357 | [5 empty] | 0.22156 |  |  |  | -0.0032374 |  |  |  |  | 0.6827 |  |  |  | -0.0078083
accuracy:
Accuracy was determined in presence of "Senonorm™ Trace elements in whole blood" reference material (with certified concentration of 1.2 ng × L−1 Lu3+, 8.6 ng × L−1 Yb3+ and 2.3 ng × L−1 Tm3+.
Blood samples were first diluted 20× (resulting Lu, Yb and Tm content lower then LOQ) followed by adition of Lu(NO3)3, Yb(NO3)3 and Tm(NO3)3 analytes to reach 250 ng × L−1 final concentration of the analytes.
raw data | [7 empty] | parameters Lu |  |  |  |  | parameters Yb |  |  |  | parameters Tm
c, ng × L−1 | Lu(NO3)3, ng × L−1 | Yb(NO3)3, ng × L−1 | Tm(NO3)3, ng × L−1 |  |  |  |  | cav (concentration average) |  |  | 253.0084 |  | cav (concentration average) |  | 249.3843 |  | cav (concentration average) |  | 256.1716
250 | 246.652 | 239.5245 | 250.8472 |  |  |  |  | n (number of points) |  |  | 5 |  | n (number of points) |  | 5 |  | n (number of points) |  | 5
| 248.6564 | 252.6717 | 249.1273 |  | accuracy Lu: |  |  | sd (std. deviation) |  |  | 6.1423 |  | sd (std. deviation) |  | 6.3306 |  | sd (std. deviation) |  | 5.7514
| 250.971 | 251.9567 | 258.9697 |  | t = | 0.37802 | acurate: | R (recovery) |  |  | 101.2% |  | R (recovery) |  | 99.75% |  | R (recovery) |  | 102.5%
| 257.622 | 247.0446 | 261.5556 |  | tcrit = | 2.7764 | yes | u(R) (recovery uncertainty) |  |  | 3.183% |  | u(R) (recovery uncertainty) |  | 3.224% |  | u(R) (recovery uncertainty) |  | 3.081%
| 261.1405 | 255.7242 | 260.3581 |  |  |  |  | t1−α;(n−1) (critical t-test value) |  |  | 2.7764 |  | t1−α;(n−1) (critical t-test value) |  | 2.7764 |  | t1−α;(n−1) (critical t-test value) |  | 2.7764
accuracy Yb:
[5 empty] | t = | 0.076392 | acurate:
[5 empty] | tcrit = | 2.7764 | yes
accuracy Tm:
[5 empty] | t = | 0.80125 | acurate:
[5 empty] | tcrit = | 2.7764 | yes

## Validation_LC-MS-MS
LC-MS/MS
Reference:
V. Barwick (Ed), Eurachem/CITAC Guide: Guide to Quality in Analytical Chemistry: An Aid to Accreditation (3rd ed. 2016). ISBN 978-0-948926-32-7. Available from www.eurachem.org.
Validation was performed with HO2CPh{Lu} and HO2CPh{Tm} (toghether) in presence of mouse liver lysate (matrix blank) with HO2CPh{Y} as an internal standard. Concentration of stock solution of HO2CPh{Lu}, HO2CPh{Tm}
and HO2CPh{Y} was determined by ICP-OES. Stock solution of matrix blank was obtained from 100 mg mouse liver that was processed (see Materials and Methods) into a final volume of 400 µL in MOPS/NaOH buffer (pH 7.0).
limit of detection (LOD) and limit of quantification (LOQ):
1, 2 and 5 nM triplicate solutions of HO2CPh{Lu} and HO2CPh{Tm} were prepared by combining 2, 4 and 10 µL of 10 nM stock solution mixture with 4 µL of matrix blanc, 2 µL of 1 µM HO2CPh{Y} internal standard solution and 12, 10 and 4 µL of 10 mM NH4OAc diluent solution.
raw data (integration of peak area) | [6 empty] | calculations Lu | [6 empty] | calculations Tm
c, nM | HO2CPh{Lu} | HO2CPh{Tm} | HO2CPh{Y} |  |  |  | Lu/Y | (ci − cav)2 | [5 empty] | Tm/Y | (ci − cav)2
1 | 3100 | 21400 | 2180000 |  |  |  | 0.001422 | 2.7778 |  | parameters Lu |  |  |  | 0.0098165 | 2.7778 | parameters Tm
1 | 4130 | 20200 | 2230000 |  |  |  | 0.001852 | 2.7778 |  | b0 (intercept) |  |  | -0.0002751 | 0.0090583 | 2.7778 | b0 (intercept) |  | -0.00018783
1 | 3170 | 17400 | 2200000 | LOD and LOQ Lu: |  |  | 0.0014409 | 2.7778 |  | b1 (slope) |  |  | 0.0017059 | 0.0079091 | 2.7778 | b1 (slope) |  | 0.0091257
2 | 6840 | 39000 | 2090000 | LOD = | 0.5661 | nM | 0.0032727 | 0.44444 |  | syx (residual std. deviation) |  |  | 0.00043317 | 0.01866 | 0.44444 | syx (residual std. deviation) |  | 0.0010586
2 | 4960 | 39000 | 2140000 | LOQ = | 1.6983 | nM | 0.0023178 | 0.44444 |  | n (number of points) |  |  | 9 | 0.018224 | 0.44444 | n (number of points) |  | 9
2 | 6870 | 36600 | 2110000 |  |  |  | 0.0032559 | 0.44444 |  | cav (concentration average) |  |  | 2.6667 | 0.017346 | 0.44444 | cav (concentration average) |  | 2.6667
5 | 16800 | 92200 | 2070000 | LOD and LOQ Tm: |  |  | 0.0081159 | 5.4444 |  | t1−α;(n−2) (critical t-test value) |  |  | 1.8946 | 0.044541 | 5.4444 | t1−α;(n−2) (critical t-test value) |  | 1.8946
5 | 17300 | 87200 | 1960000 | LOD = | 0.2586 | nM | 0.0088265 | 5.4444 |  | Σ (ci − cav)2 |  |  | 26 | 0.04449 | 5.4444 | Σ (ci − cav)2 |  | 26
5 | 16400 | 97400 | 2060000 | LOQ = | 0.77581 | nM | 0.0079612 | 5.4444 | [5 empty] | 0.047282 | 5.4444
repeatability (precision):
A 50 nM solution of HO2CPh{Lu} and HO2CPh{Tm} was prepared by mixing 3 µL of 500 nM stock solution mixture with 27 µL of a matrix blank stock solution. 10 nM pentaplicate solutions of HO2CPh{Lu}
and HO2CPh{Tm} were prepared by combining 4 µL of 50 nM stock solution mixture in matrix blanc with 2 µL of 1 µM HO2CPh{Y} internal standard solution and 14 µL of 10 mM NH4OAc diluent solution.
raw data | [6 empty] | calculation Lu
c, nM | HO2CPh{Lu}, nM | HO2CPh{Tm}, nM |  |  |  |  | c, nM | std. dev.
50 | 53.4 | 52.9 |  | c, nM | rel. std. dev. | repeatable: | 50 | 1.7065
50 | 49.7 | 47.4 |  | 50 | 3.328% | yes
50 | 52.7 | 47.1 |  |  |  |  | calculation Tm
50 | 49.7 | 48 |  | c, nM | rel. std. dev. | repeatable: | c, nM | std. dev.
50 | 50.9 | 53 |  | 50 | 6.044% | yes | 50 | 3.0028
linearity (2‒1000 nM range for HO2CPh{Lu} and 1‒1000 nM range for HO2CPh{Tm}):
1‒1000 nM duplicate solutions of HO2CPh{Lu} and HO2CPh{Tm} were prepared by combining 2‒10 µL of stock solution mixture of higher concentration with 4 µL of matrix blank, 2 µL of 1 µM HO2CPh{Y} internal standard solution and 4‒12 µL of 10 mM NH4OAc diluent solution (to reach final volume 20 µL). 1 nM HO2CPh{Lu} was excluded from linearity test due to peak area below LOQ.
raw data (integration of peak area) | [16 empty] | calculations Tm
c, nM | HO2CPh{Lu} | HO2CPh{Tm} | HO2CPh{Y} | [13 empty] | c, nM | Lu/Y | yi | Yreg,i | I × (yi − yreg,i)2 | yij − yi
1 | below LOQ | 27900 | 3360000 |  |  |  | calculations Lu | [9 empty] | 1 | 0.0083036 | 0.0088685 | 0.017102 | 0.00013559 | -0.00056488
1 | below LOQ | 28300 | 3000000 |  |  |  | c, nM | Lu/Y | yi | Yreg,i | I × (yi − yreg,i)2 |  | yij − yi |  |  |  |  | 0.0094333 |  |  |  | 0.00056488
2 | 7010 | 49600 | 2870000 |  |  |  | 2 | 0.0024425 | 0.0025849 | 0.0046239 | 8.3148e-06 |  | -0.00014238 |  |  |  | 2 | 0.017282 | 0.016753 | 0.024966 | 0.0001349 | 0.00052923
2 | 7800 | 46400 | 2860000 | linearity Lu: |  |  |  | 0.0027273 |  |  |  |  | 0.00014238 |  |  |  |  | 0.016224 |  |  |  | -0.00052923
5 | 17200 | 112000 | 2850000 | F = | 0.070922 | linear: | 5 | 0.0060351 | 0.0059577 | 0.0083901 | 1.1833e-05 |  | 7.7403e-05 | parameters Lu |  |  | 5 | 0.039298 | 0.039896 | 0.048556 | 0.00015 | -0.00059736 | parameters Tm
5 | 16700 | 115000 | 2840000 | Fcrit = | 3.9715 | yes |  | 0.0058803 |  |  |  |  | -7.7403e-05 | b0 (intercept) |  | 0.0021131 |  | 0.040493 |  |  |  | 0.00059736 | b0 (intercept) | 0.0092389
10 | 37700 | 230000 | 2910000 |  |  |  | 10 | 0.012955 | 0.013583 | 0.014667 | 2.3507e-06 |  | -0.0006276 | b1 (slope) |  | 0.0012554 | 10 | 0.079038 | 0.081098 | 0.087873 | 9.1804e-05 | -0.00206 | b1 (slope) | 0.0078634
10 | 40500 | 237000 | 2850000 |  |  |  |  | 0.014211 |  |  |  |  | 0.0006276 | n (number of points) |  | 14 |  | 0.083158 |  |  |  | 0.00206 | n (number of points) | 16
50 | 189000 | 1140000 | 2800000 | linearity Tm: |  |  | 50 | 0.0675 | 0.06855 | 0.064883 | 2.6894e-05 |  | -0.00105 | I (number of repetitions) |  | 2 | 50 | 0.40714 | 0.41957 | 0.40241 | 0.00058908 | -0.012429 | I (number of repetitions) | 2
50 | 174000 | 1080000 | 2500000 | F = | 0.040797 | linear: |  | 0.0696 |  |  |  |  | 0.00105 | N (number of calibration points) |  | 7 |  | 0.432 |  |  |  | 0.012429 | N (number of calibration points) | 8
100 | 394000 | 2430000 | 2910000 | Fcrit = | 3.5806 | yes | 100 | 0.1354 | 0.13376 | 0.12765 | 7.4655e-05 |  | 0.0016326 | Σ I × (yi − Yreg,i)2 |  | 0.00031758 | 100 | 0.83505 | 0.83088 | 0.79558 | 0.0024927 | 0.0041684 | Σ I × (yi − Yreg,i)2 | 0.011206
100 | 366000 | 2290000 | 2770000 |  |  |  |  | 0.13213 |  |  |  |  | -0.0016326 | Σ Σ (yij − yi)2 |  | 0.0062691 |  | 0.82671 |  |  |  | -0.0041684 | Σ Σ (yij − yi)2 | 0.36624
500 | 1700000 | 10600000 | 2770000 |  |  |  | 500 | 0.61372 | 0.62073 | 0.62981 | 0.00016507 |  | -0.0070094 | F1−α;(N−2,n−N) (critical f-test value) |  | 3.9715 | 500 | 3.8267 | 3.8842 | 3.9409 | 0.0064483 | -0.057446 | F1−α;(N−2,n−N) (critical f-test value) | 3.5806
500 | 1720000 | 10800000 | 2740000 |  |  |  |  | 0.62774 |  |  |  |  | 0.0070094 |  |  |  |  | 3.9416 |  |  |  | 0.057446
1000 | 3340000 | 20700000 | 2770000 |  |  |  | 1000 | 1.2058 | 1.2613 | 1.2575 | 2.8471e-05 |  | -0.055509 |  |  |  | 1000 | 7.4729 | 7.8968 | 7.8726 | 0.0011638 | -0.42384
1000 | 3450000 | 21800000 | 2620000 |  |  |  |  | 1.3168 |  |  |  |  | 0.055509 |  |  |  |  | 8.3206 |  |  |  | 0.42384
accuracy:
10 nM pentaplicate solutions of HO2CPh{Lu} and HO2CPh{Tm} were prepared by combining 2 µL of 100 nM stock solution mixture with 4 µL of matrix blank, 2 µL of 1 µM HO2CPh{Y} internal standard solution and 12 µL of 10 mM NH4OAc diluent solution.
raw data | [6 empty] | parameters Lu |  |  |  |  | parameters Tm
c, nM | HO2CPh{Lu}, nM | HO2CPh{Tm}, nM |  |  |  |  | cav (concentration average) |  |  | 10.156 |  | cav (concentration average) |  | 10.18
10 | 9.08 | 10.1 |  |  |  |  | n (number of points) |  |  | 5 |  | n (number of points) |  | 5
| 10.8 | 10.6 |  | accuracy Lu: |  |  | sd (std. deviation) |  |  | 0.71671 |  | sd (std. deviation) |  | 0.27559
| 10.4 | 10.3 |  | t = | 0.20941 | acurate: | R (recovery) |  |  | 101.6% |  | R (recovery) |  | 101.8%
| 10.7 | 9.93 |  | tcrit = | 2.7764 | yes | u(R) (recovery uncertainty) |  |  | 7.449% |  | u(R) (recovery uncertainty) |  | 3.426%
| 9.8 | 9.97 |  |  |  |  | t1−α;(n−1) (critical t-test value) |  |  | 2.7764 |  | t1−α;(n−1) (critical t-test value) |  | 2.7764
accuracy Tm:
|  |  |  | t = | 0.52533 | acurate:
|  |  |  | tcrit = | 2.7764 | yes