# Supplemental material - Caregiver’s response to cardiorespiratory events in preterm infants in the NICU – a quantitative overview

### Authors

Sophie J.E. Cramer1, Henriette A. van Zanten1, Hylke H. Salverda1, Stuart B. Hooper23, Janneke Dekker1, Arjan B. te Pas1

## Affiliations

- Willem-Alexander Children’s Hospital, department of pediatrics, division of Neonatology, Leiden University Medical Center, Leiden, the Netherlands

- The Ritchie Centre, Hudson Institute of Medical Research, Melbourne, VIC, Australia,

- Department of Obstetrics and Gynecology, Monash University, Melbourne, VIC, Australia,

## Corresponding author

Sophie J.E. Cramer

Postzone J6-S, Room J7-73
Willem-Alexander Peadiatric Hospital, Leiden University Medical Center

Albinusdreef 2, 2333 ZA, Leiden, the Netherlands

s.j.e.cramer@lumc.nl

## Cardiorespiratory alarm settings

Alarms that are activated for cardiorespiratory monitoring by default are (1) low heart frequency (HF<100 bpm) and bradycardia (HF<80 bpm) alarms originating from the electrocardiogram (ECG) signal via the Philips monitor, (2) desaturation alarms (SpO2<90%) originating from the ventilator’s built-in pulse oximeter (Masimo SET, Masimo, USA), (3) alarms indicating an (automated) increase in FiO2 above a set limit (FiO2>0.6) originating from the ventilator and (4) alarms indicating a steep automated increase in FiO2 (0.05 over 30 sec) originating from the ventilator. Apnoea alarms based on impedance from the ECG leads are disabled but apnoea alarms originating from the ventilator’s pressure breath detection module (no breaths detected for 15 sec) are occasionally activated on discretion of the caregiver. Only a selection of these alarms is transferred to the handhelds of the caregivers, including apnoea, bradycardia, desaturation and alarms related to the automated FiO2 titration.

#### Overview cardiorespiratory alarms settings

| Alarm type | Detection method | Device | Default alarm limit | Status | Transferred to handheld caregivers |
| --- | --- | --- | --- | --- | --- |
| Apnoea | Impedance via ECG leads | Philips patient monitor | - | Disabled | No |
| Apnoea | Breath detection via a pressure line | SLE 6000 | No breath detected for 15 seconds | Activated on discretion of the caregiver | Yes |
| Low heart rate | ECG | Philips patient monitor | <100 bpm | Activated by default | No |
| Bradycardia | ECG | Philips patient monitor | <80 bpm | Activated by default | Yes |
| Desaturation | SpO2 via pulse oximeter | Masimo SET module build in SLE6000 ventilator | <90% | Yes | Yes |
| High FiO2 requirement | Oxygen sensor | SLE6000 | FiO2 > 0.6 | Yes | Yes |
| Steep automated FiO2 increase | Oxygen sensor | SLE6000 | 0.05 increase over 30 seconds | Yes | Yes |