Supplemental data

Evaluation of radionuclidic impurities

Each stock solution was checked for the presence of gamma-ray emitting radionuclide impurities by gamma-ray spectrometry analysis using a high-purity germanium detector (Canberra HPGe GR1018; Mirion Technologies, Georgia, USA), calibrated conform the procedures described by the International Atomic Energy Agency (IAEA technical reports series no 454, 2006). From each stock solution, three aliquots were prepared by pipetting 20 µL of solution into a 1.5-mL Eppendorf cup. Each aliquot was measured twice at a fixed distance of 18.5 cm from the detector cap in order to minimize system dead-time effect. The first measurement, having a measurement time of 5 min, was performed at the day of sample preparation to verify the absence of short-lived radionuclide impurities. The second measurement, having a measurement time of 30 min, was performed after a minimum of 10 half-lives of the main radionuclide (6.5 for 131I, 8 for 177Lu) to check for short and long-lived radionuclide impurities, respectively.

Each measurement was corrected for radioactive decay, dead-time, background and true coincidence effects. For each stock solution, the radionuclide impurities were determined as the average values of the three aliquots measured.

Determination of reference activity

Table S1: Corrections applied to activity measurements performed with the reference chambers (Fidelis and ISOCAL-III).

| Correction | Fidelis | ISOCAL-III |
| --- | --- | --- |
| Linearity | < +0.4% (dependent on measured current) | none |
| Radionuclide impurities | 177Lu: -0.5% (177mLu) | 177Lu: -0.5% (177mLu) |
| Relative response against NPL master chamber | gamma emitters: ≤ -1.1% | gamma emitters: ≤ -0.6% |
|  | 90Y: -14.5% |  |

Table S2: Uncertainty assessment of radionuclide activity measurements using the reference chambers (Fidelis and ISOCAL-III).

| Source of uncertainty | Standard uncertainty (k=1) (%) |
| --- | --- |
| NPL calibration factor | 99mTc: 0.90, 111In: 0.75, 123I: 0.90, 124I: 0.50, 131I: 0.40, 177Lu: 0.50, 90Y: 2.17 |
| Energy response correction | gamma emitters: 0.27, 90Y: 2.0 |
| Day-to-day system stability | Fidelis: ≤ 0.1, ISOCAL: ≤ 0.2 |
| Linearity (correction) | Fidelis (linearity correction): 0.03, ISOCAL: 0.1 |
| Counting statistics of ionization current | Fidelis: ≤ 0.05, ISOCAL: ≤ 0.7 (90Y excluded) |
| Geometry effects for 124I measurements in Schott 1+ vial (this study) vs P6 vial (NPL calibration factor) | 124I: 0.6 |

Radionuclide calibrator measurement settings

Table S3: Radionuclide calibrator specifications.

| Manufacturer | Model (Chamber) | University hospital | Radionuclide calibrator | Year of purchase |
| --- | --- | --- | --- | --- |
| Capintec | CRC-25R | A | 4 | 2010 |
| Capintec | CRC-15R | D | 2 | 1995 |
| Capintec | CRC-25PET | E | 2 | 2018 |
| Capintec | CRC-15PET | J | 2 | 1992 |
| Nuvia | Isomed 2010 (640000-17) | A | 1 | 2015 |
| Nuvia | Isomed 2010 (640000-17) | A | 2 | 2015 |
| Nuvia | Isomed 2010 (640000-17) | A | 3 | 2015 |
| Nuvia | Isomed 2010 (638605-01)* | D | 1 | 2017 |
| Nuvia | Isomed 2010 (640000-17) | E | 1 | 2018 |
| Nuvia | Isomed 2010 (638600-17) | F | 1 | 1999 |
| Nuvia | Isomed 2010 (640000) | F | 2 | 2010 |
| Nuvia | Isomed 2010 (638600-17) | F | 3 | 2012 |
| Nuvia | Isomed 2010 (640000) | F | 4 | 2011 |
| Nuvia | Isomed 2010 (unknown) | F | 5 | 2012 |
| Nuvia | Isomed 2010 (unknown) | F | 6 | 2012 |
| PTW Freiburg | Curiementor 4 (TN33003-1,8) | D | 3 | 2006 |
| Comecer | VIK-202 | C | 1 | 2010 |
| Comecer | VIK-202 | C | 2 | 2010 |
| Comecer | VIK-202 | C | 3 | 2004 |
| Comecer | VIK-202 | E | 3 | 1994 |
| Comecer | VIK-202 | E | 4 | 2002 |
| Comecer | VIK-202 | G | 1 | 2007 |
| Comecer | VIK-202 | G | 2 | 2007 |
| Comecer | VDC-404 | G | 3 | 1987 |
| Comecer | VIK-202 | H | 1 | 2000 |
| Comecer | VIK-202 | H | 2 | 2006 |
| Comecer | VIK-202 | H | 3 | 2011 |
| Comecer | VIK-202 | H | 4 | 2017 |
| Comecer | VIK-202 | H | 5 | 2017 |
| Comecer | VIK-202 | J | 1 | 2016 |
| Comecer | VIK-202 | J | 3 | 2011 |
| Comecer | VIK-202 | J | 4 | 2015 |

* A specific ionization chamber especially designed for the institution.

Radionuclide calibrator measurement settings

Table S4: Radionuclide calibrator measurement settings for 99mTc.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.128 | 0.128 | 0.129 | 0.126 |
| A2 | 0.128 | 0.128 | 0.129 | 0.126 |
| A3 | 0.128 | 0.128 | 0.129 | 0.126 |
| A4 | 80 | 80 | 80 | 80 |
| C1 | 236 | 236 | 236 | 236 |
| C2 | 236 | 236 | 236 | 236 |
| C3 | 236 | 236 | 236 | 236 |
| D1 | 0.128 | 0.128 | 0.129 | 0.126 |
| D2 | 80 | 80 | 80 | 80 |
| D3 | 1.256 | 1.256 | 1.256 | 1.256 |
| E1 | 0.128 | 0.128 | 0.129 | 0.126 |
| E2 | 90 | 90 | 90 | 90 |
| E3 | 236 | 236 | 236 | 236 |
| E4 | 236 | 236 | 236 | 236 |
| F1 | 0.128 | 0.128 | 0.129 | 0.126 |
| F2 | 0.128 | 0.128 | 0.129 | 0.126 |
| F3 | 0.128 | 0.128 | 0.129 | 0.126 |
| F4 | 0.128 | 0.128 | 0.129 | 0.126 |
| F5 | 0.128 | 0.128 | 0.129 | 0.126 |
| F6 | 0.128 | 0.128 | 0.129 | 0.126 |
| G1 | 236 | 236 | 236 | 236 |
| G2 | 236 | 236 | 236 | 236 |
| G3 | 236 | 236 | 236 | 236 |
| H1 | 236 | 236 | 236 | 236 |
| H2 | 236 | 236 | 236 | 236 |
| H3 | 236 | 236 | 236 | 236 |
| H4 | 236 | 236 | 236 | 236 |
| H5 | 236 | 236 | 236 | 236 |
| J1 | 236 | 236 | 236 | 236 |
| J2 | Not measured | Not measured | Not measured | Not measured |
| J3 | Not measured | Not measured | Not measured | Not measured |
| J4 | Not measured | Not measured | Not measured | Not measured |

Table S5: Radionuclide calibrator measurement settings for 111In.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.069 | 0.067 | 0.08 | 0.078 |
| A2 | 0.069 | 0.067 | 0.08 | 0.078 |
| A3 | 0.069 | 0.067 | 0.08 | 0.078 |
| A4 | 303 | 303 | 303 | 303 |
| C1 | 645 (81%) | 645 (81%) | 645 (100%) | 645 (100%) |
| C2 | 645 (81%) | 645 (81%) | 645 (100%) | 645 (100%) |
| C3 | 645 (81%) | 645 (81%) | 645 (100%) | 645 (100%) |
| D1 | 0.061 | 0.061 | 0.076 | 0.076 |
| D2 | 303 | 303 | 303 | 303 |
| D3 | 0.55 | 0.55 | 0.55 | 0.55 |
| E1 | 0.061 | 0.06 | 0.076 | 0.076 |
| E2 | 1† | 1† | 1† | 1† |
| E3 | 676 | 676 | 676 | 676 |
| E4 | 676 | 676 | 676 | 676 |
| F1 | 0.061 | 0.06 | 0.076 | 0.076 |
| F2 | 0.061 | 0.06 | 0.076 | 0.076 |
| F3 | 0.061 | 0.06 | 0.076 | 0.076 |
| F4 | 0.061 | 0.06 | 0.076 | 0.076 |
| F5 | 0.061 | 0.06 | 0.076 | 0.076 |
| F6 | 0.061 | 0.06 | 0.076 | 0.076 |
| G1 | 676 | 676 | 676 | 676 |
| G2 | 676 | 676 | 676 | 676 |
| G3 | 676 | 676 | 676 | 676 |
| H1 | 676 (95%) | 676 (95%) | 676 (110%) | 676 (110%) |
| H2 | 676 (95%) | 676 (95%) | 676 (110%) | 676 (110%) |
| H3 | 676 (95%) | 676 (95%) | 676 (110%) | 676 (110%) |
| H4 | 711 | 711 | 711 | 711 |
| H5 | 711 | 711 | 711 | 711 |
| J1 | Not measured | Not measured | Not measured | Not measured |
| J2 | Not measured | Not measured | Not measured | Not measured |
| J3 | Not measured | Not measured | Not measured | Not measured |
| J4 | Not measured | Not measured | Not measured | Not measured |

† Calibration factor was not present in the communication module. In this case, it appears that the system uses a calibration factor of 1 (no error warning was displayed). This measurement was excluded from the detailed analysis.

Table S6: Radionuclide calibrator measurement settings for 123I.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.066 | 0.067 | 0.095 | 0.095 |
| A2 | 0.066 | 0.067 | 0.095 | 0.095 |
| A3 | 0.066 | 0.067 | 0.095 | 0.095 |
| A4 | 277 | 277 | 277 | 277 |
| C1 | 618 (81%) | 618 (81%) | 618 (100%) | 618 (100%) |
| C2 | 618 (81%) | 618 (81%) | 618 (100%) | 618 (100%) |
| C3 | 618 (81%) | 618 (81%) | 618 (100%) | 618 (100%) |
| D1 | 0.066 | 0.067 | 0.095 | 0.095 |
| D2 | 277 | 277 | 277 | 277 |
| D3 | 0.93 | 0.93 | 0.93 | 0.93 |
| E1 | 0.066 | 0.067 | 0.095 | 0.095 |
| E2 | 1† | 1† | 1† | 1† |
| E3 | 676 (85%) | 676 (85%) | 676 (133%) | 676 (133%) |
| E4 | 676 (85%) | 676 (85%) | 676 (133%) | 676 (133%) |
| F1 | 0.066 | 0.067 | 0.095 | 0.095 |
| F2 | 0.066 | 0.067 | 0.095 | 0.095 |
| F3 | 0.066 | 0.067 | 0.095 | 0.095 |
| F4 | 0.066 | 0.067 | 0.095 | 0.095 |
| F5 | 0.066 | 0.067 | 0.095 | 0.095 |
| F6 | 0.066 | 0.067 | 0.095 | 0.095 |
| G1 | 618 | 618 | 618 | 618 |
| G2 | 618 | 618 | 618 | 618 |
| G3 | 618 | 618 | 618 | 618 |
| H1 | 618 (100%) | 618 (100%) | 618 (130%) | 618 (130%) |
| H2 | 618 (100%) | 618 (100%) | 618 (130%) | 618 (130%) |
| H3 | 618 (100%) | 618 (100%) | 618 (130%) | 618 (130%) |
| H4 | 618 | 618 | 618 | 618 |
| H5 | 618 | 618 | 618 | 618 |
| J1 | 618 | 618 | 618 | 618 |
| J2 | Not measured | Not measured | Not measured | Not measured |
| J3 | Not measured | Not measured | Not measured | Not measured |
| J4 | Not measured | Not measured | Not measured | Not measured |

† Calibration factor was not present in the communication module. In this case, it appears that the system uses a calibration factor of 1 (no error warning was displayed). This measurement was excluded from the detailed analysis.

Table S7: Radionuclide calibrator measurement settings for 124I.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.099 | 0.101 | 0.129 | 0.128 |
| A2 | 0.099 | 0.101 | 0.129 | 0.128 |
| A3 | 0.099 | 0.101 | 0.129 | 0.128 |
| A4 | 570 | 570 | 570 | 570 |
| B1 | 813 | 813 | 813 | 813 |
| C1 | 813 | 813 | 813 | 813 |
| C2 | 813 | 813 | 813 | 813 |
| C3 | 813 | 813 | 813 | 813 |
| D1 | 0.099 | 0.101 | 0.1145 | 0.128 |
| D2 | 570 | 570 | 570 | 570 |
| D3 | Not available | Not available | Not available | Not available |
| E1 | 0.099 | 0.101 | 0.129 | 0.128 |
| E2 | 1† | 1† | 1† | 1† |
| E3 | Not available | Not available | Not available | Not available |
| E4 | Not available | Not available | Not available | Not available |
| F1 | 0.099 | 0.101 | 0.129 | 0.128 |
| F2 | 0.099 | 0.101 | 0.129 | 0.128 |
| F3 | 0.099 | 0.101 | 0.129 | 0.128 |
| F4 | 0.099 | 0.101 | 0.129 | 0.128 |
| F5 | 0.099 | 0.101 | 0.129 | 0.128 |
| F6 | 0.099 | 0.101 | 0.129 | 0.128 |
| G1 | 813 | 813 | 813 | 813 |
| G2 | 813 | 813 | 813 | 813 |
| G3 | Not available | Not available | Not available | Not available |
| H1 | 813 | 813 | 813 | 813 |
| H2 | 813 | 813 | 813 | 813 |
| H3 | 813 | 813 | 813 | 813 |
| H4 | 813 | 813 | 813 | 813 |
| H5 | 813 | 813 | 813 | 813 |
| J1 | 813 | 813 | 813 | 813 |
| J2 | Not available | Not available | Not available | Not available |
| J3 | 813 | 813 | 813 | 813 |
| J4 | 813 | 813 | 813 | 813 |

† Calibration factor was not present in the communication module. In this case, it appears that the system uses a calibration factor of 1 (no error warning was displayed). This measurement was excluded from the detailed analysis.

Table S8: Radionuclide calibrator measurement settings for 131I.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.26 | 0.26 | 0.265 | 0.265 |
| A2 | 0.26 | 0.26 | 0.265 | 0.265 |
| A3 | 0.26 | 0.26 | 0.265 | 0.265 |
| A4 | 151 | 151 | 151 | 151 |
| B1 | 480 | 480 | 480 | 480 |
| C1 | 480 | 480 | 480 | 480 |
| C2 | 480 | 480 | 480 | 480 |
| C3 | 480 | 480 | 480 | 480 |
| D1 | 0.2453 | 0.2453 | 0.25 | 0.25 |
| D2 | 151 | 151 | 151 | 151 |
| D3 | 0.93 | 0.93 | 0.93 | 0.93 |
| E1 | 0.26 | 0.26 | 0.265 | 0.265 |
| E2 | 165‡ | 165‡ | 165‡ | 165‡ |
| E3 | 480 | 480 | 480 | 480 |
| E4 | 480 | 480 | 480 | 480 |
| F1 | 0.26 | 0.26 | 0.265 | 0.265 |
| F2 | 0.26 | 0.26 | 0.265 | 0.265 |
| F3 | 0.26 | 0.26 | 0.265 | 0.265 |
| F4 | 0.26 | 0.26 | 0.265 | 0.265 |
| F5 | 0.26 | 0.26 | 0.265 | 0.265 |
| F6 | 0.26 | 0.26 | 0.265 | 0.265 |
| G1 | 480 | 480 | 480 | 480 |
| G2 | 480 | 480 | 480 | 480 |
| G3 | 447 | 447 | 447 | 447 |
| H1 | 480 | 480 | 480 | 480 |
| H2 | 480 | 480 | 480 | 480 |
| H3 | 480 | 480 | 480 | 480 |
| H4 | 480 | 480 | 480 | 480 |
| H5 | 480 | 480 | 480 | 480 |
| J1 | 480 | 480 | 480 | 480 |
| J2 | 151 | 151 | 151 | 151 |
| J3 | 480 | 480 | 480 | 480 |
| J4 | 480 | 480 | 480 | 480 |

‡ System was not commissioned to measure this radionuclide, although a calibration factor was present in the communication module (no error warning was displayed). This measurement was excluded from the detailed analysis.

Table S9: Radionuclide calibrator measurement settings for 177Lu.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.44 | 0.44 | 0.44 | 0.44 |
| A2 | 0.44 | 0.44 | 0.44 | 0.44 |
| A3 | 0.44 | 0.44 | 0.44 | 0.44 |
| A4 | 150x10 | 150x10 | 150x10 | 150x10 |
| C1 | 751x10 | 751x10 | 751x10 | 751x10 |
| C2 | 751x10 | 751x10 | 751x10 | 751x10 |
| C3 | 751x10 | 751x10 | 751x10 | 751x10 |
| D1 | 0.44 | 0.44 | 0.44 | 0.44 |
| D2 | 450x10 | 450x10 | 450x10 | 450x10 |
| D3 | 4.88 | 4.88 | 4.88 | 4.88 |
| E1 | 0.44 | 0.44 | 0.44 | 0.44 |
| E2 | 1† | 1† | 1† | 1† |
| E3 | 751x10 | 751x10 | 751x10 | 751x10 |
| E4 | 751x10 | 751x10 | 751x10 | 751x10 |
| F1 | 0.44 | 0.44 | 0.44 | 0.44 |
| F2 | 0.44 | 0.44 | 0.436 | 0.436 |
| F3 | 0.44 | 0.44 | 0.44 | 0.44 |
| F4 | 0.44 | 0.44 | 0.44 | 0.44 |
| F5 | 0.44 | 0.44 | 0.44 | 0.44 |
| F6 | 0.44 | 0.44 | 0.44 | 0.44 |
| G1 | 751x10 | 751x10 | 751x10 | 751x10 |
| G2 | 751x10 | 751x10 | 751x10 | 751x10 |
| G3 | 751x10 | 751x10 | 751x10 | 751x10 |
| H1 | 751x10 | 751x10 | 751x10 | 751x10 |
| H2 | 751x10 | 751x10 | 751x10 | 751x10 |
| H3 | 751x10 | 751x10 | 751x10 | 751x10 |
| H4 | 751x10 | 751x10 | 751x10 | 751x10 |
| H5 | 751x10 | 751x10 | 751x10 | 751x10 |
| J1 | 751x10 | 751x10 | 751x10 | 751x10 |
| J2 | Not measured | Not measured | Not measured | Not measured |
| J3 | 751x10 | 751x10 | 751x10 | 751x10 |
| J4 | 751x10 | 751x10 | 751x10 | 751x10 |

† Calibration factor was not present in the communication module. In this case, it appears that the system uses a calibration factor of 1 (no error warning was displayed). This measurement was excluded from the detailed analysis.

Table S10: Radionuclide calibrator measurement settings for 90Y.

|  | Calibration factor/dial setting (sample geometry correction factor) |  |  |  |
| --- | --- | --- | --- | --- |
| Radionuclide calibrator | 1 mL syringe | 3 mL syringe | 1 mL vial | 10 mL vial |
| A1 | 0.81 | 0.67 | 1.8 | 2.07 |
| A2 | 0.81 | 0.67 | 1.8 | 2.07 |
| A3 | 0.81 | 0.67 | 1.8 | 2.07 |
| A4 | 48x10 | 48x10 | 48x10 | 48x10 |
| C1 | 902x100 | 902x100 | 902x100 | 902x100 |
| C2 | 902x100 | 902x100 | 902x100 | 902x100 |
| C3 | 902x100 | 902x100 | 902x100 | 902x100 |
| D1 | 0.81 | 0.67 | 1.8 | 2.07 |
| D2 | 48x10 | 48x10 | 48x10 | 48x10 |
| D3 | 28.95 | 28.95 | 28.95 | 28.95 |
| E1 | 0.81 | 0.67 | 1.8 | 2.07 |
| E2 | 1† | 1† | 1† | 1† |
| E3 | 902x100 | 902x100 | 902x100 | 902x100 |
| E4 | 902x100 | 902x100 | 902x100 | 902x100 |
| F1 | 0.81§ | 0.67§ | 1.8§ | 2.07§ |
| F2 | 2.13§ | 2.13§ | 3.35§ | 3.05§ |
| F3 | 0.81§ | 0.67§ | 1.8§ | 2.07§ |
| F4 | 0.81§ | 0.67§ | 1.8§ | 2.07§ |
| F5 | 0.81§ | 0.67§ | 1.8§ | 2.07§ |
| F6 | 0.81§ | 0.67§ | 1.8§ | 2.07§ |
| G1 | 902x100 | 902x100 | 902x100 | 902x100 |
| G2 | 902x100 | 902x100 | 902x100 | 902x100 |
| G3 | 902x100 | 902x100 | 902x100 | 902x100 |
| H1 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |
| H2 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |
| H3 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |
| H4 | 902x100 | 902x100 | 902x100 | 902x100 |
| H5 | 902x100 | 902x100 | 902x100 | 902x100 |
| J1 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |
| J2 | Not measured | Not measured | Not measured | Not measured |
| J3 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |
| J4 | 902x100(100%) | 902x100(100%) | 902x100(92.4%) | 902x100(92.4%) |

† Calibration factor was not present in the communication module. In this case, it appears that the system uses a calibration factor of 1 (no error warning was displayed). This measurement was excluded from the detailed analysis.

§ Measurements performed with aluminum dipper.