Table 3.  Results under different lighting conditions. 
Lighting 	Observer 1	Observer 2	Comments
Conditions	1      2       3	1      2      3

WHPM (Lot #C31103100)
DL		2       2	     2	2      2       2	Took longer to fully develop but still clear after 5
IL		2       2	     2	2      2       2	min dwell. Some water may have gotten to flow IL Off		U      U	     U	U     U      U 	area from dip step. A slight pink tone was noted.
IL Indirect	2       2	     2	2      2       2
Flashlight	2       2	     2	2      2       2
WHPM (Lot #F3290919)
DL		2       2	     2	2      2       2	Still readable after 10 plus min beyond 5 
IL		2       2	     2	2      2       2	min post-dwell.
IL Off		U      U	     U	U     U      U 	
IL Indirect	2       2	     2	2      2       2
Flashlight	2       2	     2	2      2       2
WHPM (Lot #D3130455)
DL		2       2	     2	2      2       2	Sample 2 had much slower development and much
IL		2       2	     2	2      2       2	more pronounced pink hue than Sample 1. Sample 2
IL Off		U      U	     U	U     U      U 	was also very difficult to make out the test line 
IL Indirect	2       2	     2	2      2       2	IL indirect conditions.
Flashlight	2       2	     2	2      2       2
Henso 
DL		2       2	     2	2      2       2	Started quickly with strong clear C and T lines but  
IL		2       2	     2	2      2       2	the T line began to visibly fade at minute 4 of the 
IL Off		U      U	     U	U     U      U 	dwell period. Inconsistent test line (some spots
IL Indirect	2       2	     2	2      2       2	colored but other blank regions) but still readable 
Flashlight	2       2	     2	2      2       2	during initial observations and 5 min post initial
						dwell period.
Hangzhou SingClean
DL		2       2	     2	2      2       2	Inconsistent and slow development during dwell
IL		2       2	     2	2      2       2	time. At 5 min exactly, Sample 1 did not have 2
IL Off		U      U	     U     U    U	     U 	visible lines because of pink background hue but
IL Indirect	1C     2	     1C	2      2       2	then became visible later.
Flashlight	2       2	     2	2      2       2	
China Med
DL		2       2	     2	2      2       2	Pink hue throughout dwell but did rapidly show 
IL		2       2	     2	2      2       2	control and test lines and the background cleaned
IL Off		U      U	     U	U     U      U 	up by the end of the 5-min dwell.
IL Indirect	2       2	     2	2      2       2
Flashlight	2       2	     2	2      2       2
Tianjin Mingao (Lot #20231114)
DL		2       2	     2	2      2       2	Test line was very faint.
IL		2       2	     2	2      2       2	
IL Off		U      U	     U     U    U	     U 	
IL Indirect	1C    1C    2    2      1C     1C	
Flashlight	2       2	     2	2      2       2	
Tianjin Mingao (Lot #20231103)
DL		2       2	     2	2      2       2	Test line was extremely faint. Under daylight and
IL		2       2	     2	2      2       2	room light conditions could not make out the 2 lines
IL Off		U      U	     U    U     U	     U 	from more than 2 feet away from the sample.
IL Indirect	U      1C   1C	1C   U       U	
Flashlight	2       2	     2	2      2       2	

2= two lines.  U=unreadable result. 1C=strip furthest from we test area visible. 

