Table 3.  Results under different lighting conditions.

Lighting 	Observer 1	Observer 2	Comments

Conditions	1      2       3	1      2      3

WHPM (Lot #C31103100)

DL		2       2	     2	2      2       2	Took longer to fully develop but still clear after 5

IL		2       2	     2	2      2       2	min dwell. Some water may have gotten to flow IL Off		U      U	     U	U     U      U 	area from dip step. A slight pink tone was noted.

IL Indirect	2       2	     2	2      2       2

Flashlight	2       2	     2	2      2       2

WHPM (Lot #F3290919)

DL		2       2	     2	2      2       2	Still readable after 10 plus min beyond 5

IL		2       2	     2	2      2       2	min post-dwell.

IL Off		U      U	     U	U     U      U

IL Indirect	2       2	     2	2      2       2

Flashlight	2       2	     2	2      2       2

WHPM (Lot #D3130455)

DL		2       2	     2	2      2       2	Sample 2 had much slower development and much

IL		2       2	     2	2      2       2	more pronounced pink hue than Sample 1. Sample 2

IL Off		U      U	     U	U     U      U 	was also very difficult to make out the test line

IL Indirect	2       2	     2	2      2       2	IL indirect conditions.

Flashlight	2       2	     2	2      2       2

Henso

DL		2       2	     2	2      2       2	Started quickly with strong clear C and T lines but

IL		2       2	     2	2      2       2	the T line began to visibly fade at minute 4 of the

IL Off		U      U	     U	U     U      U 	dwell period. Inconsistent test line (some spots

IL Indirect	2       2	     2	2      2       2	colored but other blank regions) but still readable

Flashlight	2       2	     2	2      2       2	during initial observations and 5 min post initial

dwell period.

Hangzhou SingClean

DL		2       2	     2	2      2       2	Inconsistent and slow development during dwell

IL		2       2	     2	2      2       2	time. At 5 min exactly, Sample 1 did not have 2

IL Off		U      U	     U     U    U	     U 	visible lines because of pink background hue but

IL Indirect	1C     2	     1C	2      2       2	then became visible later.

Flashlight	2       2	     2	2      2       2

China Med

DL		2       2	     2	2      2       2	Pink hue throughout dwell but did rapidly show

IL		2       2	     2	2      2       2	control and test lines and the background cleaned

IL Off		U      U	     U	U     U      U 	up by the end of the 5-min dwell.

IL Indirect	2       2	     2	2      2       2

Flashlight	2       2	     2	2      2       2

Tianjin Mingao (Lot #20231114)

DL		2       2	     2	2      2       2	Test line was very faint.

IL		2       2	     2	2      2       2

IL Off		U      U	     U     U    U	     U

IL Indirect	1C    1C    2    2      1C     1C

Flashlight	2       2	     2	2      2       2

Tianjin Mingao (Lot #20231103)

DL		2       2	     2	2      2       2	Test line was extremely faint. Under daylight and

IL		2       2	     2	2      2       2	room light conditions could not make out the 2 lines

IL Off		U      U	     U    U     U	     U 	from more than 2 feet away from the sample.

IL Indirect	U      1C   1C	1C   U       U

Flashlight	2       2	     2	2      2       2

2= two lines.  U=unreadable result. 1C=strip furthest from we test area visible.