Supplementary appendix.

Droplex EGFR Mutation Test v2 PCR master mixture and patient DNA samples from targeted washing fluid were combined in an 8-strip PCR tube (Axygen, Tewksbury, MA, USA) to a final reaction volume of 20 μL, and then the mixture was loaded into each well of DG8 Cartridge (Bio-Rad). The droplets were generated using QX200 Droplet Generator (Bio-Rad) following the manufacturer’s instructions. The generated uniform nanoliter-sized droplets were transferred to a 96-well plate (Eppendorf, Hamburg, Germany or Bio-Rad) and the plate was placed into a thermal cycler (Veriti 96-Well Fast Thermal Cycler, Thermo Fisher Scientific, Waltham, MA, USA) for PCR amplification. The thermocycling conditions were as follows: one cycle of 30 minutes at 37°C, followed by 10 minutes at 95°C; 40 cycles of 30 seconds at 94°C; 1 minute at 58°C, and a final step at 98°C for 10 minutes, with the ramping rate settled at 2°C/sec. After thermal cycling, the droplet reader connected to a computer running data analysis QuantaSoft Software (v1.6.6.0320; Bio-Rad) and run data analysis was carried out. Thresholds were set manually based results from positive control containing wild-type genomic DNA and standard positive DNA. The number of positive and negative droplets were distinguished by threshold, and the given numbers were used for calculating the concentration target in terms of copies/µL.

The kit is designed to detect total 107 EGFR mutations related with EGFR-tyrosine kinase inhibitor through 6 reactions per specimen. Each reaction consists of specific primers and fluorescent labeled probes for EGFR mutations; exon 18 (L718X, G719X, G724S), exon 19 deletions (59 deletions), exon 20 (S768I, T790M, C797X, 33 insertions), and exon 21 (L858R, L861Q). To confirm the validity of the test, the positive control and internal control are measured in every test. Detailed procedure of Droplex EGFR Mutation Test v2 was performed according to the manufacturer's instructions.