APPENDIX A: Cyclotron types

There is a wide range of cyclotrons, depending on the energy and intensity of the output beam, and a lot of options may be added to support multiple beams, beams with higher intensity, different target technology (gas, liquid or solid targets), etc.. The smallest cyclotrons may also benefit of a self-shielding option that enables installation in almost any conventional building, as they do not need extra radiation protection measures. Cyclotrons work per production “runs”, each providing a specific number of doses of one radiopharmaceutical. Larger cyclotrons produce larger quantities faster.

| Type of cyclotron | Main Beam Energy (MeV) | Acquisition price estimation (million €) without options | Average 18FDG Doses per run |
| --- | --- | --- | --- |
| Baby cyclotron (BC) – out of scope | < 9 | Alone: 1, but comes always with a PET scanner, so combined: 1.5 – 2.5 | 4 |
| Small cyclotron sub-group “A” (SCa) – out of scope | 9-11 | 0.8 – 1.6 | < 20 |
| Small cyclotron sub-group “B” (SCb) – within scope | 12-19 | 1.2 – 2 | < 100 |
| Medium Cyclotron (MC) – out of scope | 20-40 | 15 | > 100 |
| Large Cyclotron (LC) – out of scope | > 40 | >50 | > 100 |

Cyclotron types and corresponding main features.

EXAMPLES OF OPTIONS:

- Self-shielding (only available for BC and SCa): ≈ 1 million €.

- Solid or gaseous targets for producing other than 18F isotopes: ≈ 0.5 million € per target.

- Multi-beam: starts at 0.2 million € for the smaller cyclotrons and can go up to several million € for MC and LC.

BC, MC and LC:

A “baby cyclotron” (BC) is a low energy cyclotron intended for very small productions, sufficient only for in-hospital use. BC is an appealing and not too expensive solution that by-passes the main disadvantage of off-hospital production, which is that radiopharmaceuticals have a short half-life and thus need to complete a “speed race” to reach the hospital. The sole use of BC is for PET radiopharmaceutical production and not γ-camera, while its production rate is small and sufficient only for a few doses per production cycle, thus sufficient for only one PET scanner located in the same hospital, and only for a small number of patients per day. This small number of radio-pharmaceutical doses produced in each run can be a relative advantage when different radiopharmaceuticals are needed in the same day, but its small production makes it out of scope for this article (too expensive to install one BC per NM department), although they can prove quite versatile for research centres and University Hospitals. Similarly, medium (MC) and large cyclotrons (LC) have production rates, along with acquisition prices and running costs, that are too big for places with rather limited population sizes, such as islands, so again out of scope for this article.

SCa, SCb and HC:

SCa models cannot be used in a hybrid way (HC), since their beam energy does not correspond to the optimal beam energy to produce the main γ-camera radionuclide (99mTc), which is between 11 MeV and 14 MeV (please note that the maximal real-life energy is usually about 0.5 MeV lower than the nominal one, so no SCa can have a real-life beam of 11 MeV). This 11 MeV to 14 MeV energy range offers a plateau in 99mTc cross section production and minimizes contaminations due to target impurities. SCb PET radiopharmaceutical production is much higher than SCa, and several magnitudes higher than BC. SCb are typically used to provide PET radiopharmaceuticals for several NM departments in an area up to 4 hours of delivery range. Their production goes up to 100 doses per run, but given the 110-minutes half-life of 18F, after an equal transportation time there will be 50 doses left to be used. If we round-up the 110-minutes into two hours and we consider that each PET examination lasts about half an hour, 4 patients can be scanned per half-life, after which the equivalent of (50 doses – 4 doses used) / 2 will be left. By continuing in the same logic, each run producing 100 doses can lead to 16 patients scanned at PET scanners adjacent to the production site, 12 patients for PET scanners at distances around 2 hours away and 8 patients at PET scanners at distances around 4 hours away.

HC are SCb with at least one extra solid target and their nominal energy has been slightly decreased, so their price (see “example of options”) would be 1.7 – 2.5 million €. Since almost all commercially available SCb have an output energy that is higher than 14 MeV, the latter needs to be lowered to at most 14 MeV to be used for HC, without lowering its beam intensity. Examples of SCb that have been validated as HC include (this is not an exclusive list, others may exist as well): “PetTrace” 16 MeV of GE, “Cube” 18 MeV of IBA and “TR19” 19 MeV of ACSI.

Not all cyclotrons can produce all isotopes, nor are all isotopes deliverable at distances more than a few minutes. However, the most common PET isotope (18F) can indeed be produced by SCb and, thus, by HC, as can 99mTc be produced by HC and both can be delivered at distances going up to a few hours.