Appendix 1: Calculation of conversion factor in PCXMC

Siemens ArtisQ Ceiling DynaCT
Parameters in PCXMC:
Avagerage tube voltage: 90 kV
Anode angle: 11.5 degrees
Tube-to-detector distance: 105 cm
Field of view: 30x38 cm (diameter 48 cm)
Area: abdomen, kidney
	
	DAP (mGycm2)
	Effective dose (mSv)*
	

	0 degrees
	250
	0.023383
	

	90 degrees
	250
	0.050701
	

	180 degrees
	250
	0.016885
	

	270 degrees
	250
	0.082371
	

	Total
	1000
	0.17334
	


*ICRP103

This results in a conversion factor of 0.17 mSv Gy-1 cm-2

