{"text_format":"3","workbook":"xlsx","n_sheets":1,"sheets":["Deidentified Dataset N95 & Duo"],"read":"whole","date_system":"1900"}

{"sheet":"Deidentified Dataset N95 & Duo","index":0,"n_rows":43,"n_cols":42,"rows_kept":43,"merged":[[0,0,0,1],[0,2,0,10],[0,11,0,24],[0,25,0,39],[0,40,0,41]]}
["Study Identification",null,"Demographic Information",{"$nulls":8},"3M N95",{"$nulls":13},"Duo ",{"$nulls":14},"Feedback"]
["Study ID","Date of Test (yyyy/mm/dd)","Age (Years)","Sex (M/F)","Height (cm)","Weight (kg)","BMI (kg/m^2)","Face Width (mm)","Face Length (mm)","NIOSH Panel","Naso-Menton Distance (mm)","3M Model","Negative Pressure Test","Run #1\nNormal Breathing","Run #2 Deep Breathing","Run #3 Head Side-to-Side","Run #4 Head Up-and-Down","Run #5 Talking out loud","Run #6 Bending over","Run #7 Normal Breathing","Overall Harmonic Fit Factor (Run 1-7)","Stationary Harmonic Fit Factor (Run 1,2,7)","Dynamic Harmonic Fit Factor (Run 3-6)","Pass/Fail (1,0)","Retest Y/N (1,0)","Run #1\nNormal Breathing","Run #2 Deep Breathing","Run #3 Head Side-to-Side","Run #4 Head Up-and-Down","Run #5 Talking out loud","Run #6 Bending over","Run #7 Normal Breathing","Overall Harmonic Fit Factor (Run 1-7)","Stationary Harmonic Fit Factor (Run 1,2,7)","Dynamic Harmonic Fit Factor (Run 3-6)","Pass/Fail (1,0)","Retest Y/N (1,0)","Duo Model","Duo Size","Filter Type","Comfort (0-5)","Ease of Breathing (0-5)"]
[1,"2020-05-07",35,"F",157.5,47.7,{"v":19.22902494331066,"d":"19.23"},144,110,4,85,"1870+","Pass",8.13,5.19,7.92,9.27,21.53,6.9,9.73,{"v":8.2941707636759,"d":"8.29"},{"v":7.169240283518522,"d":"7.17"},{"v":9.40044334179973,"d":"9.40"},0,0,{"v":11455,"d":"11455.00"},{"v":13817,"d":"13817.00"},{"v":7836,"d":"7836.00"},{"v":24748,"d":"24748.00"},{"v":2814,"d":"2814.00"},{"v":16099,"d":"16099.00"},{"v":15947,"d":"15947.00"},{"v":8665,"d":"8665.00"},{"v":13490.39429577652,"d":"13490.39"},{"v":6831.707947583242,"d":"6831.71"},1,0,"Duo Nos","M","Mini x 2, Big Red, H100",3.5,4]
[2,"2020-05-07",32,"F",167,64,{"v":22.948115744558788,"d":"22.95"},117,114,"NA",95,"1870+","Pass",2368.05,346.23,1054.18,581.28,640.39,784.89,367.45,{"v":606.7667256073393,"d":"606.77"},{"v":497.3474342858078,"d":"497.35"},{"v":726.6704850591582,"d":"726.67"},1,0,38374,15294,59543,26391,3503,99999999,29577,15042,{"v":23951.23825857918,"d":"23951.24"},{"v":11758.96859790235,"d":"11758.97"},1,0,"Duo Nos","M","Big Red, Mini x 2, H100",5,5]
[3,"2020-05-07",34,"M",180,72.7,{"v":22.438271604938272,"d":"22.44"},129,129,6,105,1860,"Pass",487.2,429.76,715.84,721.4,235.53,510.38,542.23,{"v":460.1665711834287,"d":"460.17"},{"v":482.03138700517405,"d":"482.03"},{"v":445.02684642413897,"d":"445.03"},1,0,{"v":17813,"d":"17813.00"},{"v":23094,"d":"23094.00"},{"v":38783,"d":"38783.00"},{"v":99999999,"d":"99999999.00"},{"v":10386,"d":"10386.00"},{"v":5446,"d":"5446.00"},{"v":132447,"d":"132447.00"},{"v":16962,"d":"16962.00"},{"v":28039.936825021206,"d":"28039.94"},{"v":13084.768187673648,"d":"13084.77"},1,0,"Duo Nos","M","Big Red, Mini x 2, H100",5,5]
[4,"2020-05-08",32,"M",182,90,{"v":27.17063156623596,"d":"27.17"},148,135,10,120,1860,"Pass",98.81,77.56,27.87,53.15,66.83,30.52,40.53,46.63214356700647,{"v":62.91054535566235,"d":"62.91"},{"v":39.05323610881529,"d":"39.05"},0,0,{"v":3904,"d":"3904.00"},{"v":13045,"d":"13045.00"},{"v":4631,"d":"4631.00"},{"v":4559,"d":"4559.00"},{"v":979,"d":"979.00"},{"v":16126,"d":"16126.00"},{"v":4503,"d":"4503.00"},{"v":3375.7325572879968,"d":"3375.73"},{"v":5406.579879938885,"d":"5406.58"},{"v":2633.7543094807584,"d":"2633.75"},1,0,"Duo 3","M","Yellow T3, Mini x 2",5,5]
[5,"2020-05-08",31,"F",163,58,{"v":21.829952199932254,"d":"21.83"},121,115,3,100,"1870+","Pass",1877,3064,1380,1335,574,810,2292,{"v":1218.2672885589307,"d":"1218.27"},{"v":2315.8226581322215,"d":"2315.82"},{"v":898.7896497723945,"d":"898.79"},1,0,{"v":28214,"d":"28214.00"},{"v":10971,"d":"10971.00"},{"v":9756,"d":"9756.00"},{"v":20018,"d":"20018.00"},{"v":2287,"d":"2287.00"},{"v":99999999,"d":"99999999.00"},{"v":20762,"d":"20762.00"},{"v":9158,"d":"9158.00"},{"v":17166.624832634338,"d":"17166.62"},{"v":6782.8787447982995,"d":"6782.88"},1,0,"Duo Nos","M","Big red, Mini x2, H100 ",4,3.5]
[6,"2020-05-09",29,"F",164,54,{"v":20.077334919690664,"d":"20.08"},115,114,"NA",90,"1870+","Pass",1395,409,665,696,611,372,526,{"v":567.7970527470449,"d":"567.80"},{"v":592.5370740506984,"d":"592.54"},{"v":550.5566342278904,"d":"550.56"},1,0,25715.66,15599.62,28520.74,{"v":17360,"d":"17360.00"},{"v":1692,"d":"1692.00"},{"v":32082,"d":"32082.00"},{"v":60344,"d":"60344.00"},{"v":8390,"d":"8390.00"},{"v":25091.46003244057,"d":"25091.46"},{"v":5595.559410448978,"d":"5595.56"},1,0,"Duo Nos","M","Big red, Mini x 2, H100",4,4]
[7,"2020-05-10",33,"M",178,74,{"v":23.355636914530994,"d":"23.36"},132,130,6,105,8210,"Pass",777.56,369.77,490.08,234.05,367.66,443.08,692.2,{"v":418.5342535209326,"d":"418.53"},{"v":551.9649861699867,"d":"551.96"},{"v":354.29873845266087,"d":"354.30"},1,0,{"v":9999999,"d":"9999999.00"},{"v":99999999,"d":"99999999.00"},{"v":24424,"d":"24424.00"},{"v":19822,"d":"19822.00"},{"v":2748,"d":"2748.00"},{"v":9921,"d":"9921.00"},{"v":57569,"d":"57569.00"},{"v":12207,"d":"12207.00"},{"v":171620.1995408019,"d":"171620.20"},{"v":7193.084912016046,"d":"7193.08"},1,0,"Duo Int","M","Big red, Mini x 2, H100",5,4]
[8,"2020-05-11",46,"M",178,75,{"v":23.671253629592222,"d":"23.67"},133,126,6,105,"1870+","Pass",161.69,134.29,184.13,84.57,233.69,138.67,270.25,{"v":151.91786302347145,"d":"151.92"},{"v":173.09500399688028,"d":"173.10"},{"v":139.14975619494243,"d":"139.15"},1,0,{"v":24091,"d":"24091.00"},{"v":7167,"d":"7167.00"},{"v":27507,"d":"27507.00"},{"v":8965,"d":"8965.00"},{"v":5159,"d":"5159.00"},{"v":2746,"d":"2746.00"},{"v":315,"d":"315.00"},{"v":1724,"d":"1724.00"},{"v":894.0170112711678,"d":"894.02"},{"v":5666.514136127905,"d":"5666.51"},1,1,"Duo Int","M","Red Big, Mini x 2, H100",4,3.5]
[9,"2020-05-11",29,"F",167,58,{"v":20.796729893506402,"d":"20.80"},116,110,"NA",97.5,"1870+","Pass",48.58,74.1,41.21,40.9,25.78,39.85,52.25,42.22,{"v":56.371231246393535,"d":"56.37"},{"v":35.52407490813866,"d":"35.52"},0,1,8540.27,8631.06,14614.33,10058.66,{"v":4450,"d":"4450.00"},183.97,139.93,{"v":530,"d":"530.00"},{"v":406.53803690732934,"d":"406.54"},{"v":686.3146222276497,"d":"686.31"},1,0,"Duo Nos","M","Mini 2, Big Red",3,3]
[10,"2020-05-13",37,"M",178,72.7,{"v":22.945335184951396,"d":"22.95"},126,115,3,95,8210,"Pass",432.84,321.21,238.19,155.21,183.96,120.35,180.16,197.96,{"v":273.3691360809486,"d":"273.37"},{"v":164.0266865166562,"d":"164.03"},1,0,{"v":11869,"d":"11869.00"},{"v":8666,"d":"8666.00"},{"v":3896,"d":"3896.00"},{"v":3330,"d":"3330.00"},{"v":641,"d":"641.00"},{"v":7430,"d":"7430.00"},{"v":16987,"d":"16987.00"},{"v":2787,"d":"2787.00"},{"v":11604.736427257983,"d":"11604.74"},{"v":1776.4941887910913,"d":"1776.49"},1,0,"Duo Nos","M","Mini 2, Big Red",5,5]
[11,"2020-05-13",34,"F",168,54,{"v":19.13265306122449,"d":"19.13"},112,120,"NA",100,"1860S","Pass",231.8,58.85,242.73,126.41,73,28,179,{"v":79,"d":"79.00"},{"v":111.55315300745661,"d":"111.55"},{"v":65.10047265786825,"d":"65.10"},0,0,{"v":35284,"d":"35284.00"},{"v":8194,"d":"8194.00"},{"v":7749,"d":"7749.00"},{"v":14400,"d":"14400.00"},{"v":2085,"d":"2085.00"},{"v":30413,"d":"30413.00"},{"v":11235,"d":"11235.00"},{"v":7366,"d":"7366.00"},{"v":12531.875819325433,"d":"12531.88"},{"v":5625.9555046697205,"d":"5625.96"},1,0,"Duo Nos","M","Mini x 2,  Big red, H100",3,5]
[12,"2020-05-14",40,"M",169,53,{"v":18.556773222226116,"d":"18.56"},132,118,3,92,"1870+","Pass",980.47,441.23,750.38,576.82,350.82,643.04,829.86,585.06,{"v":667.9531101931634,"d":"667.95"},{"v":535.3405724364087,"d":"535.34"},1,0,{"v":47858,"d":"47858.00"},{"v":29915,"d":"29915.00"},{"v":33186,"d":"33186.00"},{"v":23058,"d":"23058.00"},{"v":3676,"d":"3676.00"},{"v":99999999,"d":"99999999.00"},{"v":70143,"d":"70143.00"},{"v":16903,"d":"16903.00"},{"v":43744.669465727384,"d":"43744.67"},{"v":11575.852488802982,"d":"11575.85"},1,1,"Duo 1","M","Red T3, Mini x 2",4,5]
[13,"2020-05-14",36,"M",178,83,{"v":26.19618735008206,"d":"26.20"},135,124,7,115,8210,"Pass",269.68,120.25,149.17,131.27,97.77,179.59,365.92,155.99,{"v":203.29411575408767,"d":"203.29"},{"v":132.81106401008765,"d":"132.81"},1,0,{"v":27219,"d":"27219.00"},{"v":7256,"d":"7256.00"},{"v":17555,"d":"17555.00"},{"v":5047,"d":"5047.00"},{"v":1786,"d":"1786.00"},{"v":4553,"d":"4553.00"},{"v":5295,"d":"5295.00"},{"v":5006.936075477258,"d":"5006.94"},{"v":8255.060812200383,"d":"8255.06"},{"v":3866.0522873162477,"d":"3866.05"},1,0,"Duo Int","M","Big red, Mini x 2, H100",4,4]
[14,"2020-05-14",36,"M",190,110,{"v":30.470914127423825,"d":"30.47"},139,117,4,95,1860,"Pass",174.85,93.82,187.36,127.94,112.95,134.34,220.89,139.01,{"v":143.506089383444,"d":"143.51"},{"v":135.81993990504182,"d":"135.82"},1,0,{"v":7643,"d":"7643.00"},{"v":6397,"d":"6397.00"},{"v":4570,"d":"4570.00"},{"v":8170,"d":"8170.00"},{"v":2927,"d":"2927.00"},{"v":7242,"d":"7242.00"},{"v":7311,"d":"7311.00"},{"v":5623,"d":"5623.00"},{"v":7076.437253336541,"d":"7076.44"},{"v":4872.418515500533,"d":"4872.42"},1,0,"Duo Int","M","Red T3, Mini x 2 (new), H100",4,5]
[15,"2020-05-26",33,"M",186,85,{"v":24.56931437160365,"d":"24.57"},135,127,7,102,"1870+","Pass",2028.9,929.43,354.47,641.96,498.59,643.03,1503.23,688.05,{"v":1342.858456987065,"d":"1342.86"},{"v":503.80143302082297,"d":"503.80"},1,1,{"v":1548,"d":"1548.00"},{"v":2574,"d":"2574.00"},{"v":3329,"d":"3329.00"},{"v":3781,"d":"3781.00"},{"v":1381,"d":"1381.00"},{"v":2808,"d":"2808.00"},{"v":3558,"d":"3558.00"},{"v":2364,"d":"2364.00"},{"v":2280.4115790631818,"d":"2280.41"},{"v":2431.449651223297,"d":"2431.45"},1,0,"Duo 3","M","Red T3, Mini/AirGuard",4,5]
[16,"2020-05-18",30,"M",174,80,{"v":26.42356982428326,"d":"26.42"},147,129,10,100,"1860","Pass",157.4,99.7,95.7,112.5,88.4,71.1,130.7,{"v":101.8,"d":"101.80"},{"v":124.8208828955843,"d":"124.82"},{"v":89.45519096559624,"d":"89.46"},0,0,{"v":10905,"d":"10905.00"},{"v":27432,"d":"27432.00"},{"v":6921,"d":"6921.00"},{"v":36019,"d":"36019.00"},{"v":1385,"d":"1385.00"},{"v":16667,"d":"16667.00"},{"v":7975,"d":"7975.00"},{"v":5796,"d":"5796.00"},{"v":11832.140699614183,"d":"11832.14"},{"v":4191.679745005748,"d":"4191.68"},1,1,"Duo 1","M","Red T3, Mini x 2",5,5]
[17,"2020-05-21",36,"M",183,76,{"v":22.694018931589476,"d":"22.69"},143,130,9,120,"1870+","Pass",167.5,96.1,134,160.5,128.9,109.7,162.5,{"v":131.8,"d":"131.80"},{"v":133.15684032800655,"d":"133.16"},{"v":130.8603418108837,"d":"130.86"},1,0,{"v":12636,"d":"12636.00"},{"v":28318,"d":"28318.00"},{"v":19663,"d":"19663.00"},{"v":19085,"d":"19085.00"},{"v":2016,"d":"2016.00"},{"v":99999999,"d":"99999999.00"},{"v":99999999,"d":"99999999.00"},{"v":9809,"d":"9809.00"},{"v":26209.52677230049,"d":"26209.53"},{"v":6674.499681350066,"d":"6674.50"},1,0,"Duo Nos","M","Red T3, Mini x2",4,3.5]
[18,"2020-05-21",43,"F",161,52,{"v":20.060954438486167,"d":"20.06"},124,111,3,100,"1860S","Pass",156,87.9,142.5,150.7,138.5,105,332,{"v":136.7,"d":"136.70"},{"v":144.23861303322943,"d":"144.24"},{"v":131.59076432158605,"d":"131.59"},1,0,{"v":5238,"d":"5238.00"},{"v":7818,"d":"7818.00"},{"v":49709,"d":"49709.00"},{"v":11694,"d":"11694.00"},{"v":874,"d":"874.00"},{"v":4895,"d":"4895.00"},{"v":20067,"d":"20067.00"},{"v":3841,"d":"3841.00"},{"v":8137.677068302161,"d":"8137.68"},{"v":2750.8691731757744,"d":"2750.87"},1,0,"Duo Int","M","Big Red, Mini x2, H100",4,4]
[19,"2020-05-21",36,"M",185,79.5,{"v":23.22863403944485,"d":"23.23"},121,119,6,105,"8210","Pass",108.4,111.8,148.25,128.06,164.1,114.2,199.2,{"v":133.1,"d":"133.10"},{"v":129.36768816797112,"d":"129.37"},{"v":136.03306915027832,"d":"136.03"},1,1,{"v":16660,"d":"16660.00"},{"v":12925,"d":"12925.00"},{"v":31978,"d":"31978.00"},{"v":61640,"d":"61640.00"},{"v":1997,"d":"1997.00"},{"v":6968,"d":"6968.00"},{"v":20172,"d":"20172.00"},{"v":7966,"d":"7966.00"},{"v":16045.602445611674,"d":"16045.60"},{"v":5782.360112189326,"d":"5782.36"},1,0,"Duo Nos","M","Big Red, Mini x 2, H100",5,3]
[20,"2020-05-21",35,"M",179,79,{"v":24.65590961580475,"d":"24.66"},145,122,7,108,"1860","Pass",52.94,99.5,111,103,81,79,112.5,{"v":86,"d":"86.00"},{"v":79.30539304270013,"d":"79.31"},{"v":91.48784774652934,"d":"91.49"},0,0,{"v":4650.7,"d":"4650.70"},4282.65,6598.77,8988.16,1148.02,4327,5083.62,{"v":3482,"d":"3482.00"},{"v":4649.4911438242025,"d":"4649.49"},{"v":2930.461246017785,"d":"2930.46"},1,0,"Duo Nos","M","Big Red, Mini x 2, H100",4.5,4.5]
[21,"2020-05-23",29,"F",178,68,{"v":21.461936624163616,"d":"21.46"},144,115,5,92,"1870+","Pass",425,240,254,143,210,136,273,210,{"v":294.61990160292015,"d":"294.62"},{"v":173.57449167421439,"d":"173.57"},1,0,{"v":17303,"d":"17303.00"},{"v":17097,"d":"17097.00"},{"v":32909,"d":"32909.00"},{"v":9161,"d":"9161.00"},{"v":2995,"d":"2995.00"},{"v":19536,"d":"19536.00"},{"v":44425,"d":"44425.00"},{"v":10552,"d":"10552.00"},{"v":21614.90926729242,"d":"21614.91"},{"v":7624.528847185578,"d":"7624.53"},1,0,"Duo Nos","M","Mini x2, Red Big, H100",4,4]
[22,"2020-05-23",28,"F",157,47,{"v":19.067710657633167,"d":"19.07"},148,108,5,101,"1870+","Pass",67,68,70,40,84,53,77,62,{"v":70.39234833790383,"d":"70.39"},{"v":57.09522282782944,"d":"57.10"},0,0,21588,16341,9171,14346,6578,20588,23123,13209,{"v":19898.546658425505,"d":"19898.55"},{"v":10544.655737588771,"d":"10544.66"},1,0,"Duo Nos","M","Mini x 2, Red Big, H100",4.5,4.5]
[23,"2020-05-23",29,"F",153,52,{"v":22.21367849972233,"d":"22.21"},135,127,7,101,"1860S","Pass",147,103,151,122.5,103,76,139,114,{"v":126.5518229740835,"d":"126.55"},{"v":106.23488711182783,"d":"106.23"},1,0,9132,6967,18448,4863,1463,2437,4387,3816,{"v":6237.200171946214,"d":"6237.20"},{"v":2954.8461491899607,"d":"2954.85"},1,0,"Duo Nos","M","Mini x 2, Red T3 (old), H100",4,5]
[24,"2020-05-29",28,"F",152,44,{"v":19.04432132963989,"d":"19.04"},136,112,4,106,"1870+","Pass",410,162,223,207,87,134,142,160,{"v":191.64287068895453,"d":"191.64"},{"v":141.4819271668165,"d":"141.48"},1,0,2411,2951,1807,2547,1895,1946,3075,2282,{"v":2780.772026289358,"d":"2780.77"},{"v":2012.476045019409,"d":"2012.48"},1,0,"Duo 1","M","Mini x2, Red T3 (old)",5,5]
[25,"2020-05-29",31,"F",156,46,{"v":18.902038132807363,"d":"18.90"},116,120,"NA",112,"1860S","Pass",14,21,12,12,19,13.5,28,16,{"v":19.384615384615383,"d":"19.38"},{"v":13.634551495016611,"d":"13.63"},0,0,{"v":1171,"d":"1171.00"},{"v":1745,"d":"1745.00"},{"v":1175,"d":"1175.00"},{"v":1427,"d":"1427.00"},{"v":1047,"d":"1047.00"},{"v":291,"d":"291.00"},{"v":567,"d":"567.00"},{"v":767,"d":"767.00"},{"v":940.2309896728683,"d":"940.23"},{"v":673.018771256781,"d":"673.02"},1,0,"Duo 4","M","Mini x2, Red T3 (new)",4,5]
[26,"2020-05-29",49,"M",174,77.5,{"v":25.597833267274407,"d":"25.60"},144,120,7,117,"1860","Pass",90,91,99.4,80.2,93.8,54.5,121,85.5,{"v":98.79930876341763,"d":"98.80"},{"v":77.61143849230433,"d":"77.61"},0,0,{"v":3780,"d":"3780.00"},{"v":4556,"d":"4556.00"},99999999,41863,3611,46796,2806,{"v":6028,"d":"6028.00"},{"v":3569.6427582288957,"d":"3569.64"},{"v":12414.711395376306,"d":"12414.71"},1,0,"Duo 1","M","Mini x2, Red T3 (new)",4,5]
[27,"2020-05-29",26,"F",178,73,{"v":23.04002019946976,"d":"23.04"},135,128,7,105,"8210","pass",181,327,557,277,63,134,145,{"v":159,"d":"159.00"},{"v":193.8044893749953,"d":"193.80"},{"v":139.17331042534664,"d":"139.17"},1,0,{"v":957,"d":"957.00"},{"v":1711,"d":"1711.00"},{"v":1583,"d":"1583.00"},{"v":1956,"d":"1956.00"},{"v":1856,"d":"1856.00"},{"v":3116,"d":"3116.00"},{"v":1971,"d":"1971.00"},{"v":1691,"d":"1691.00"},{"v":1404.006471961143,"d":"1404.01"},{"v":1997.3268292618945,"d":"1997.33"},1,0,"Duo 3","M","Mini x2, Red T3 (new)",5,5]
[28,"2020-05-29",34,"M",180,92,{"v":28.39506172839506,"d":"28.40"},160,140,"NA",120,"8210","Pass",76,67,71,98,42,59,77,66,{"v":73.04551946842203,"d":"73.05"},{"v":61.493746099113714,"d":"61.49"},0,0,{"v":2130,"d":"2130.00"},{"v":2403,"d":"2403.00"},2562,1551,1095,1386,1867,{"v":1711,"d":"1711.00"},{"v":2110.8198858137994,"d":"2110.82"},{"v":1498.2350124759378,"d":"1498.24"},1,0,"Duo 4","M","Mini x2, Red T3 (new)",3,4]
[29,"2020-05-30",43,"F",157,53,{"v":21.50188648626719,"d":"21.50"},125,117,3,102,"1860S","Pass",71.7,85.2,101,77.2,34.5,64.3,94.5,67.5,{"v":82.72187807635596,"d":"82.72"},{"v":59.35425706273986,"d":"59.35"},0,0,{"v":2582,"d":"2582.00"},{"v":1608,"d":"1608.00"},2536,1530,918,1771,2154,{"v":1847.9228180290884,"d":"1847.92"},{"v":2036.0520065240376,"d":"2036.05"},{"v":1480.4430356225364,"d":"1480.44"},1,0,"Duo 3","M","Mini x2, Red T3 (new)",2.5,3]
[30,"2020-05-30",36,"F",150,48,{"v":21.333333333333332,"d":"21.33"},136,114,4,101,"1870+","Pass",80.7,93.6,93.6,121,138,68.9,69.2,89.5,{"v":79.94414256786106,"d":"79.94"},{"v":98.25984660228656,"d":"98.26"},0,0,{"v":2649,"d":"2649.00"},{"v":5382,"d":"5382.00"},2931,3785,2681,3037,3334,{"v":3224,"d":"3224.00"},{"v":3475.257081450425,"d":"3475.26"},{"v":3058.9254505716003,"d":"3058.93"},1,0,"Duo 5","M","Mini x2, Red T3 (new)",2,3]
[31,"2020-05-30",42,"F",157,51,{"v":20.690494543389182,"d":"20.69"},125,108,1,98,"1870+","Pass",180,125,182,148,45.8,157,166,117,{"v":153.2202926295638,"d":"153.22"},{"v":98.87590422219394,"d":"98.88"},1,0,{"v":11917,"d":"11917.00"},{"v":36483,"d":"36483.00"},{"v":24272,"d":"24272.00"},{"v":13755,"d":"13755.00"},{"v":4166,"d":"4166.00"},{"v":16084,"d":"16084.00"},{"v":11984,"d":"11984.00"},{"v":11459,"d":"11459.00"},{"v":15402.912776678171,"d":"15402.91"},{"v":9612.783451968478,"d":"9612.78"},1,1,"Duo 1","M","Mini x2, Red T3 (new)",4,5]
[32,"2020-06-01",59,"F",157,57,{"v":23.124670372023203,"d":"23.12"},131,116,3,101,"1870+","Pass",787,196,478,495,121,343,240,271,{"v":284.6473862809223,"d":"284.65"},{"v":261.571824135026,"d":"261.57"},1,0,{"v":3143,"d":"3143.00"},{"v":4105,"d":"4105.00"},2011,1186,1590,1398,1918,{"v":1857.8367837590388,"d":"1857.84"},{"v":2769.7016633425583,"d":"2769.70"},{"v":1489.938947446469,"d":"1489.94"},1,0,"Duo 3","M","Mini x 2, Red T3 (old)",5,5]
[33,"2020-06-01",39,"M",183,113,{"v":33.74242288512645,"d":"33.74"},148,131,8,110,"1870+","Pass",418,356,233,174,67,181,213,173,{"v":303.14761373417923,"d":"303.15"},{"v":131.193966686691,"d":"131.19"},1,0,{"v":1666,"d":"1666.00"},{"v":1807,"d":"1807.00"},2521,6199,2395,4776,2125,{"v":2492,"d":"2492.00"},{"v":1847.0270495786237,"d":"1847.03"},{"v":3375.8091732128632,"d":"3375.81"},1,0,"Duo 3","M","Mini x2, Red T3 (old)",4,3]
[34,"2020-06-02",42,"M",178,79,{"v":24.933720489837143,"d":"24.93"},138,125,7,109,"1860","Pass",68,55.6,65.8,66,35.7,46.9,94.2,57,{"v":69.27253772063133,"d":"69.27"},{"v":50.19938591002988,"d":"50.20"},0,0,{"v":1330,"d":"1330.00"},{"v":1219,"d":"1219.00"},5353,2735,1991,2295,5777,{"v":2163,"d":"2163.00"},{"v":1718.8780959112407,"d":"1718.88"},{"v":2683.785883428837,"d":"2683.79"},1,1,"Duo 5","M","Mini x2, Red T3 (old) ",3,3]
[35,"2020-06-02",30,"M",180,66,{"v":20.37037037037037,"d":"20.37"},136,126,6,115,"1870+","Pass",439,247,445,341,274,321,374,334,{"v":333.32281087241813,"d":"333.32"},{"v":334.8781775780759,"d":"334.88"},1,0,{"v":3018,"d":"3018.00"},{"v":4146,"d":"4146.00"},10045,6932,3445,5690,10119,{"v":5068.061152106048,"d":"5068.06"},{"v":4468.504261295058,"d":"4468.50"},{"v":5635.1265669977665,"d":"5635.13"},1,1,"Duo 1","M","Mini x2, Red T3 (old)",5,5]
[36,"2020-06-03",47,"M",175,73,{"v":23.836734693877553,"d":"23.84"},137,127,7,117,1860,"Pass",73.3,115,87.1,100,29.9,74.5,140,71.6,{"v":101.76020004886252,"d":"101.76"},{"v":58.52343235064334,"d":"58.52"},0,0,{"v":18147,"d":"18147.00"},{"v":62981,"d":"62981.00"},14522,33101,3160,11669,6850,{"v":9746,"d":"9746.00"},{"v":13826.877076031786,"d":"13826.88"},{"v":7980.456682476957,"d":"7980.46"},1,0,"Duo 3","M","Mini x2, Red T3 (old)",4,5]
[37,"2020-06-03",28,"F",173,59,{"v":19.713321527615356,"d":"19.71"},125,118,3,109,"9105S","Pass",152.2,52.3,147,95.5,60.4,43.3,133,77.7,{"v":90.33543865277758,"d":"90.34"},{"v":70.26801375908904,"d":"70.27"},0,0,{"v":2727,"d":"2727.00"},{"v":5404,"d":"5404.00"},1266,2157,1737,2295,2933,{"v":2217,"d":"2217.00"},{"v":3360.5941576518567,"d":"3360.59"},{"v":1766.0579321993607,"d":"1766.06"},1,0,"Duo 1","M","Mini x2, Red T3 (old)",2,3]
[38,"2020-06-03",39,"M",157,56,{"v":22.7189744005842,"d":"22.72"},151,113,5,105,"1860S","Pass",196,116,133,157,134,123,175,143,{"v":154.3446777697321,"d":"154.34"},{"v":135.68064645305034,"d":"135.68"},1,0,{"v":21877,"d":"21877.00"},{"v":32055,"d":"32055.00"},4827,8615,1361,12891,7601,{"v":5208,"d":"5208.00"},{"v":14390.689268866958,"d":"14390.69"},{"v":3522.453937574913,"d":"3522.45"},1,0,"Duo 5","M","Mini x2, Red T3 (old)",2,4]
[39,"2020-06-19",46,"M",182,93,{"v":28.076319285110493,"d":"28.08"},135,139,9,110,1860,"Pass",73.9,82.9,70,56.54,110.31,54.4,82.28,72.04,{"v":79.474098306826,"d":"79.47"},{"v":67.31737657491676,"d":"67.32"},0,1,{"v":1301,"d":"1301.00"},{"v":1736,"d":"1736.00"},1227.76,955.29,914,1006,1185,{"v":1140,"d":"1140.00"},{"v":1370.765446195151,"d":"1370.77"},{"v":1012.8065175438869,"d":"1012.81"},1,0,"Duo Nos","M","Mini x2, Big Red, H100",3.5,4]
[40,"2020-06-19",33,"F",160,54,{"v":21.093749999999996,"d":"21.09"},130,110,1,100,"1870+","Pass",866.79,657.21,535.16,817.69,319.27,356.03,1232.77,559.15,{"v":860.4748706247195,"d":"860.47"},{"v":442.8475029860304,"d":"442.85"},1,0,3144.98,11538.98,8183.54,7580.26,1814.09,5023.35,7316.48,4528.59,{"v":5542.1357337629015,"d":"5542.14"},{"v":3982.3620729636245,"d":"3982.36"},1,0,"Duo Nos","M","Mini x2, Big Red, H100",2,3]
[41,"2020-06-19",44,"F",163,52,19.57168128269788,110,95,"NA",90,"1860S","Pass",42.8,26.42,56.73,50.49,36.28,36.16,54.4,40.64,{"v":37.68988518780544,"d":"37.69"},{"v":43.172522970483804,"d":"43.17"},0,1,1964.6,1954.84,3393.56,2935.41,2775.4,6299.7,4367.34,2911.11,{"v":2400.8977217051315,"d":"2400.90"},{"v":3465.033922487567,"d":"3465.03"},1,0,"Duo Nos","M","Mini x2, Big Red, H100",4,4]