Supplementary Table 1. %CVs for RNA dilutions and concentration assays across 4 operators

|  | Operator-1 | Operator-2 | Operator-3 | Operator-4 | All operators |
| --- | --- | --- | --- | --- | --- |
| Replicate 1 | 2.51 | 2.6 | 2.56 | 2.68 | NA |
| Replicate 2 | 2.7 | 2.57 | 2.40 | 2.77 | NA |
| Replicate 3 | 2.52 | 2.54 | 2.40 | 2.75 | NA |
| Replicate 4 | 2.55 | 2.53 | 2.30 | NA | NA |
| Replicate 5 | 2.5 | 2.58 | 2.30 | NA | NA |
| mean | 2.56 | 2.56 | 2.39 | 2.73 | 2.54 |
| ST DEV | 0.08 | 0.03 | 0.11 | 0.05 | 0.13 |
| %CV | 3.23 | 1.12 | 4.45 | 1.73 | 5.28 |

Five-fold dilutions of Qubit RNA HS Standard #2 were prepared by each of four different operators. Replicates of 2 µl aliquots were assayed according to manufacturer’s instructions. Measurements are shown in ng/µl.

Supplementary Table 2 Descriptive statistics for percent DNA contamination in 65 RNA samples

| Mean | 12.2 |
| --- | --- |
| Median | 11.8 |
| Standard Deviation | 5.2 |
| Minimum | 2.4 |
| Maximum | 25.9 |
| Count | 65 |

Supplementary Table 3 OncoPrism Score significance testing for intermediate precision

| Factor tested | p value (linear mixed effects models) |
| --- | --- |
| Operator | 0.59 |
| Reagent lot | 0.11 |
| Day (Batch-1 vs Batch-2) | 0.22 |
| Day (Batch-3 vs Batch-4) | 0.27 |
| Day (Batch-5 vs Batch-6) | 0.81 |
| Day (Batch-7 vs Batch-8) | 0.43 |
| Day (overall) | 0.48 |
| Machine (alternate sequencer) | 0.24 |

Supplementary Table 4 Fusion-positive clinical samples

| Sample ID | Cancer Indication | Expected Fusion | Tumor Cellularity |
| --- | --- | --- | --- |
| UWD23 | Lung | EML4::ALK | 60% |
| 6MV24 | NSCLC | EML4::ALK | 40% |
| 73Z6P | NSCLC | EML4::ALK | 30% |
| COUCS | NSCLC | EML4::ALK | 30% |
| O2USE | NSCLC | EML4::ALK | 30% |
| ST6HK | NSCLC | EML4::ALK | 60% |
| U6DI7 | NSCLC | EML4::ALK | 40% |
| WM11J | NSCLC | EML4::ALK | 70% |
| B1NVL | Colorectal | TMP3::NTRK1 | 40% |
| RNMWE | Colon | EML4::NTRK3 | 80% |
| 1KS6A | NHL (Anaplastic Large Cell Lymphoma) | NPM1::ALK | 90% |
| JRL2T | NHL (Diffuse Large B-cell Lymphoma) | CLTC::ALK | 70% |
| PB05R | Head and Neck (salivary gland) | ETV6:NTRK3 | 25% |
| KXFRY | Head and Neck (thyroid) | ETV6:NTRK3 | 30% |

Supplementary Table 5 Fusion detection - repeatability (intra-run)

| Sample ID | Expected Fusion Result | Observed Fusion Result | Concordant? |
| --- | --- | --- | --- |
| RNMWE_Rep1 | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE_Rep2 | EML4::NTRK3 | EML4::NTRK3 | Yes |
| UWD23_Rep1 | EML4::ALK | EML4::ALK | Yes |
| UWD23_Rep2 | EML4::ALK | EML4::ALK | Yes |
| MR5VU_Rep1 | not detected | not detected | Yes |
| MR5VU_Rep2 | not detected | not detected | Yes |

Supplementary Table 6 Fusion detection - intermediate precision (inter-operator, -lot, -run)

| Sample ID | Library Batch | Operator | Reagent Lot | Expected Fusion | Observed Fusion | Concordant? |
| --- | --- | --- | --- | --- | --- | --- |
| B1NVL | Batch 7 | 1 | Lot B | TPM3::NTRK1 | TPM3::NTRK1 | Yes |
| B1NVL | Batch 9 | 3 | Lot C | TPM3::NTRK1 | TPM3::NTRK1 | Yes |
| MR5VU | Batch 4 | 1 | Lot A | not detected | not detected | Yes |
| MR5VU | Batch 5 | 2 | Lot B | not detected | not detected | Yes |
| MR5VU Rep 1 | Batch 6 | 1 | Lot B | not detected | not detected | Yes |
| MR5VU Rep 2 | Batch 6 | 1 | Lot B | not detected | not detected | Yes |
| PB05R | Batch 7 | 1 | Lot C | ETV6::NTRK3; NTRK3::ETV6 | ETV6::NTRK3; NTRK3::ETV6 | Yes |
| PB05R | Batch 8 | 3 | Lot C | NTRK3::ETV6 | NTRK3::ETV6 | Yes |
| RNMWE | Batch 4 | 1 | Lot B | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE | Batch 5 | 2 | Lot B | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE Rep 1 | Batch 6 | 1 | Lot B | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE Rep 2 | Batch 6 | 1 | Lot B | EML4::NTRK3 | EML4::NTRK3 | Yes |
| UWD23 | Batch 4 | 1 | Lot B | EML4::ALK | EML4::ALK | Yes |
| UWD23 | Batch 10 | 2 | Lot C | EML4::ALK | EML4::ALK | Yes |
| UWD23 | Batch 9 | 3 | Lot C | EML4::ALK | EML4::ALK | Yes |
| UWD23 | Batch 5 | 2 | Lot B | EML4::ALK | EML4::ALK | Yes |
| UWD23 Rep 1 | Batch 6 | 1 | Lot B | EML4::ALK | EML4::ALK | Yes |
| UWD23 Rep 2 | Batch 6 | 1 | Lot B | EML4::ALK | EML4::ALK | Yes |

Supplementary Table 7 Fusion detection - intermediate precision (machine)

| Sample ID | Fusions detected on Sequencer-1 | Fusions detected on Sequencer-2 | Concordant between sequencers? |
| --- | --- | --- | --- |
| RNMWE | EML4::NTRK3 | EML4::NTRK3 | Yes |
| UWD23 | EML4::ALK | EML4::ALK | Yes |
| 6MV24 | EML4::ALK | EML4::ALK | Yes |
| 1KS6A | NPM1::ALK | NPM1::ALK | Yes |
| RYUGZ | not detected | not detected | Yes |
| PLBGS | not detected | not detected | Yes |
| 174DS | not detected | not detected | Yes |
| TX79T | not detected | not detected | Yes |
| E0DLZ | not detected | not detected | Yes |
| YUKWB | not detected | not detected | Yes |
| Fusion-positive reference RNA | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | Yes |
| Fusion-negative reference RNA | not detected | not detected | Yes |

Supplementary Table 8 Fusion detection at varying RNA input amounts

| Sample ID | RNA input | Expected Result | Observed Result |
| --- | --- | --- | --- |
| B1NVL | 20 ng | TPM3::NTRK1 | TPM3::NTRK1 |
|  | 35 ng | TPM3::NTRK1 | TPM3::NTRK1 |
|  | 50 ng | TPM3::NTRK1 | TPM3::NTRK1 |
|  | 65 ng | TPM3::NTRK1 | TPM3::NTRK1 |
|  | 80 ng | TPM3::NTRK1 | TPM3::NTRK1 |
| PB05R | 20 ng | ETV6::NTRK3 | ETV6::NTRK3 |
|  | 35 ng | ETV6::NTRK3 | ETV6::NTRK3 |
|  | 50 ng | ETV6::NTRK3 | ETV6::NTRK3 |
|  | 65 ng | ETV6::NTRK3 | ETV6::NTRK3 |
|  | 80 ng | ETV6::NTRK3 | ETV6::NTRK3 |
| UWD23 | 20 ng | EML4::ALK | EML4::ALK |
|  | 35 ng | EML4::ALK | EML4::ALK |
|  | 50 ng | EML4::ALK | EML4::ALK |
|  | 65 ng | EML4::ALK | EML4::ALK |
|  | 80 ng | EML4::ALK | EML4::ALK |

Supplementary Table 9 Fusion gene detection in fusion-negative/fusion-positive RNA mixes

| Fusion-positive RNA ID | %fusion- positive RNA in mix | Expected Result | Observed Result | Supporting Reads | Concordance |
| --- | --- | --- | --- | --- | --- |
| Positive-reference RNA | 100% | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | 1430 1242 995 169 | 4/4 |
|  | 25% | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | 404 375 274 69 | 4/4 |
|  | 10% | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | 197 154 142 27 | 4/4 |
|  | 5% | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | 94 81 75 21 | 4/4 |
|  | 1% | TPM3::NTRK1 EML4::ALK QKI::NTRK2 ETV6::NTRK3 | TPM3::NTRK1 EML4::ALK QKI::NTRK2 | 23 16 12 | 3/4 |
| PB05R | 100% | ETV6::NTRK3 | ETV6::NTRK3 | 293 14 | 1/1 |
|  | 25% | ETV6::NTRK3 | ETV6::NTRK3 | 65 | 1/1 |
|  | 10% | ETV6::NTRK3 | ETV6::NTRK3 | 29 | 1/1 |
|  | 5% | ETV6::NTRK3 | ETV6::NTRK3 | 11 | 1/1 |
|  | 1% | ETV6::NTRK3 | not detected | N/A | 0/1 |
| UWD23 | 100% | EML4::ALK | EML4::ALK | 71 | 1/1 |
|  | 25% | EML4::ALK | EML4::ALK | 21 | 1/1 |
|  | 10% | EML4::ALK | EML4::ALK | 9 | 1/1 |
|  | 5% | EML4::ALK | not detected | N/A | 0/1 |
|  | 1% | EML4::ALK | not detected | N/A | 0/1 |

Supplementary Table 10 Extrapolated fusion gene copy number in fusion-positive/fusion-negative reference RNA mixes

| Fusion gene | RNA mix with minimum copy number detected | *Copy number in 50 ng | *Copy number in RNA mix |
| --- | --- | --- | --- |
| TPM3::NTRK1 | 1% fusion-positive reference RNA | 14,540 | 145.4 |
| QKI::NTRK2 | 1% fusion-positive reference RNA | 31,254 | 312.54 |
| EML4::ALK | 1% fusion-positive reference RNA | 5,775 | 57.75 |
| ETV6::NTRK3 | 5% fusion-positive reference RNA | 2,242 | 112.1 |

*Copy number input extrapolated from manufacturer’s ddPCR measurements.

Supplementary Table 11 RNA sample properties for study of gDNA as an interferent

| Sample ID | Baseline %gDNA | Expected fusion |
| --- | --- | --- |
| RNMWE | 13.7% | EML4::NTRK3 |
| UWD23 | 4.5% | EML4::ALK |
| 226BJ | 6.7% | none detected |

Supplementary Table 12 Fusions detected for study of gDNA as an interferent

| Sample ID | Expected Fusion Result | Observed Fusion Result | Concordant? |
| --- | --- | --- | --- |
| RNMWE_0%_gDNA_spike-in | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE_5%_gDNA_spike-in | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE_10%_gDNA_spike-in | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE_20%_gDNA_spike-in | EML4::NTRK3 | EML4::NTRK3 | Yes |
| RNMWE_30%_gDNA_spike-in | EML4::NTRK3 | EML4::NTRK3 | Yes |
| UWD23_0%_gDNA_spike-in | EML4::ALK | EML4::ALK | Yes |
| UWD23_5%_gDNA_spike-in | EML4::ALK | EML4::ALK | Yes |
| UWD23_10%_gDNA_spike-in | EML4::ALK | EML4::ALK | Yes |
| UWD23_20%_gDNA_spike-in | EML4::ALK | EML4::ALK | Yes |
| UWD23_30%_gDNA_spike-in | EML4::ALK | EML4::ALK | Yes |
| 226BJ_0%_gDNA_spike-in | not detected | not detected | Yes |
| 226BJ_5%_gDNA_spike-in | not detected | not detected | Yes |
| 226BJ_10%_gDNA_spike-in | not detected | not detected | Yes |
| 226BJ_20%_gDNA_spike-in | not detected | not detected | Yes |
| 226BJ_30%_gDNA_spike-in | not detected | not detected | Yes |

Supplementary Table 13 Validation of OncoPrism-HNSCC fusion detection by orthogonal methods

| SAMPLE ID | Cancer type | Expected fusion | Orthogonal testing method | Concordant? (Y/N) | Comments |
| --- | --- | --- | --- | --- | --- |
| UWD23 | Lung | EML4::ALK | TSO500, ALK IHC, GatewaySeq | Y | GatewaySeq detected no fusions, but sample was positive for EML4::ALK fusion by TSO500 and positive for ALK fusion by IHC, and is therefore designated as concordant |
| 6MV24 | NSCLC | EML4::ALK | GatewaySeq, ALK IHC | Y |  |
| 73Z6P | NSCLC | EML4::ALK | GatewaySeq, ALK IHC | Y |  |
| COUCS | NSCLC | EML4::ALK | TSO500, GatewaySeq, ALK IHC | Y | GatewaySeq detected no fusions and sample failed ALK IHC QC, but was positive for EML4::ALK fusion by TSO500, and is therefore designated as concordant |
| O2USE | NSCLC | EML4::ALK | TSO500, GatewaySeq, ALK IHC | Y | GatewaySeq detected no fusions, but sample was positive for EML4::ALK fusion by TSO500 and positive for ALK fusion by IHC, so is therefore designated as concordant |
| ST6HK | NSCLC | EML4::ALK | GatewaySeq, ALK IHC | Y |  |
| U6DI7 | NSCLC | EML4::ALK | TSO500, GatewaySeq, ALK IHC | Y | GatewaySeq detected no fusions and sample failed ALK IHC QC, but was positive for EML4::ALK fusion by TSO500, and is therefore designated as concordant |
| WM11J | NSCLC | EML4::ALK | GatewaySeq, ALK IHC | Y |  |
| B1NVL | Colorectal | TMP3::NTRK1 | TSO500, Pan-Trk IHC | Y |  |
| RNMWE | Colon | EML4::NTRK3 | TSO500, GatewaySeq, Pan-Trk IHC | Y | Pan-Trk IHC was equivocal, but EML4::NTRK3 fusion was confirmed by TSO500 and GatewaySeq |
| 1KS6A | NHL | NPM1::ALK | ALK IHC, GatewaySeq | Y |  |
| JRL2T | NHL | CLTC::ALK | ALK IHC, GatewaySeq | Y |  |
| PB05R | Head and Neck (SGC) | ETV6:NTRK3 | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| KXFRY | Head and Neck (Thyroid) | ETV6:NTRK3 | GatewaySeq | Y |  |
| Y462C | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq, STAR-Fusion | Y | Pan-Trk IHC was positive; NGS-based GatewaySeq was negative for ALK, NTRK1, ETV6::NTRK3 fusions; sample RNA was negative by STAR-Fusion; probable Pan-Trk IHC false-positive |
| LQQUH | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, Star-Fusion | Y | Pan-Trk IHC was positive, but sample was QNS for additional orthogonal DNA testing; sample RNA was negative by STAR-Fusion; probable Pan-Trk IHC false-positive |
| 07H1V | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 1AMJ7 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 226BJ | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 3HD42 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 5AX1V | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 5ZW04 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| B0U4W | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| I96PP | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| J16T3 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| KD4VA | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| KTL7F | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| MR5VU | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| RYUGZ | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| ULR4R | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| V7MWZ | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| VS8KQ | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| W69KR | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| WDCLX | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC, GatewaySeq | Y |  |
| 0OUOY | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 174DS | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 3Y1GY | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 4K9GH | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 51E2C | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 5GHP9 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 6EDBL | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 79Y53 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| 7Z3F1 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| A15X0 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| DCL56 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| DUJQC | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| E0DLZ | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| H2AW8 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| HF89S | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| IDUDB | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| IPX5J | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| J0BUV | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| L851R | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| ODL4U | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| OPB1Q | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| PLBGS | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| PLZTE | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| PZY1K | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| QFPJW | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| QP4DC | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| QR103 | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| R26TL | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| SUU9G | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| TX79T | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| UXTBE | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| WN52W | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |
| YUKWB | Head and Neck (HNSCC) | not detected | ALK IHC, Pan-Trk IHC | Y |  |

Supplementary Table 14 OncoPrism-HNSCC fusion-detection compared to orthogonal methods

| Accuracy (95% CI) | Negative percent agreement (95% CI) | Positive percent agreement (95% CI) |
| --- | --- | --- |
| 100% (94.6–100%) | 100% (93.3–100%) | 100% (76.8–100%) |