Table S2: Summary of analytical conditions of DI-LC-MS/MS and SBSE-GC-MS/MS methods

- DI-LC-MS/MS: direct injection liquid chromatography coupled with tandem mass spectrometry

| DI-LC-MS/MS | Operating conditions |
| --- | --- |
| Column | XSelect HSS T3, 150 x 2.1 mm, 2.5 µm (Waters) with pre-column XSelect HSS T3, 5 x 2.1 mm, 2.5 µm (Waters) |
| Oven temperature | 30°C |
| Sample temperature | 10°C |
| Solvent flush | Water / mix 25% water, 25% MeOH, 25% AcN, 25% Isopropanol |
| Injection volume | 100 μL |
| Solvents | Water + 0.1% formic acid (A) AcN + 0.1% formic acid (B) |
| Gradient | 90% A during 0.5 min, then 100% B at 8 min (0% A) until 10 min, 90% A at 10.5 min until 13 min Total run time 13 min |
| Flow | 300 μL/min |
| Ionisation parameters | Electospray ionisation: switch ESI+/ESI- Ion spray voltage: +5500V/-4500V Source temperature: 300°C Curtain gas flow: 35 au Ion source gas 1: 40 au Ion source gas 2: 50 au |

Table A: MS/MS fragmentation parameters and retention times of the direct injection LC-MS/MS method for compounds analysed in MRM ESI+, first run

| Transition | RT (min) | Q1 | Q3 | DP (V) | EP (V) | CE (V) | CXP (V) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 13C-DEDIA-1 | 2.8 | 149.0 | 105.8 | 16 | 10 | 27 | 14 |
| 15N-DPC-1 | 2.6 | 148.0 | 67.0 | 81 | 10 | 49 | 10 |
| 2,6 dichlorobenzamide-1 | 5.1 | 190.1 | 172.9 | 50 | 10 | 25 | 13 |
| 2,6 dichlorobenzamide-2 | 5.1 | 190.1 | 145.0 | 50 | 10 | 40 | 13 |
| 6 chloronicotinic acid-1 | 5.2 | 157.9 | 122.1 | 81 | 10 | 25 | 8 |
| 6 chloronicotinic acid-2 | 5.2 | 157.9 | 78.0 | 81 | 10 | 33 | 8 |
| Acetamiprid-1 | 5.5 | 222.9 | 126.0 | 51 | 10 | 29 | 14 |
| Acetamiprid-2 | 5.5 | 222.9 | 98.9 | 51 | 10 | 53 | 10 |
| Acetochlor-1 | 8.4 | 270.1 | 223.9 | 61 | 10 | 13 | 16 |
| Acetochlor-2 | 8.4 | 270.1 | 148.1 | 61 | 10 | 27 | 10 |
| Atrazine-1 | 6.9 | 216.1 | 174.0 | 86 | 4.5 | 23 | 4 |
| Atrazine-2 | 6.9 | 216.1 | 104.1 | 86 | 4.5 | 39 | 4 |
| Bentazon-methyl-1 | 7.8 | 254.9 | 213.0 | 61 | 10 | 17 | 14 |
| Bentazon-methyl-2 | 7.8 | 254.9 | 149.1 | 61 | 10 | 29 | 14 |
| Benzamide-1 | 4.5 | 122.0 | 77.0 | 95 | 10 | 33 | 13 |
| Benzamide-2 | 4.5 | 122.0 | 79.0 | 95 | 10 | 19 | 13 |
| Boscalid-1 | 7.9 | 343.0 | 307.0 | 111 | 10 | 29 | 20 |
| Boscalid-2 | 7.9 | 343.0 | 140.0 | 111 | 10 | 29 | 12 |
| Carbofuran-3-OH -1 | 5.2 | 237.9 | 181.1 | 41 | 10 | 15 | 10 |
| Carbofuran-3-OH -2 | 5.2 | 237.9 | 163.0 | 41 | 10 | 21 | 16 |
| CGA 37735-1 | 5.4 | 194.0 | 136.1 | 61 | 10 | 21 | 12 |
| CGA 37735-2 | 5.4 | 194.0 | 148.0 | 61 | 10 | 21 | 12 |
| CGA 50267-1 | 6.4 | 208.0 | 135.0 | 61 | 10 | 22 | 13 |
| CGA 50267-2 | 6.4 | 208.0 | 120.0 | 61 | 10 | 41 | 13 |
| Chloridazon-1 | 5.3 | 222.0 | 104.0 | 91 | 10 | 31 | 8 |
| Chloridazon-2 | 5.3 | 222.0 | 92.0 | 91 | 10 | 33 | 10 |
| Chlorotoluron-1 | 6.7 | 213.1 | 72.0 | 74 | 6 | 33 | 4 |
| Chlorotoluron-2 | 6.7 | 213.1 | 46.2 | 74 | 6 | 27 | 4 |
| Chlorpyrifos-Et-1 | 9.9 | 350.3 | 96.9 | 151 | 10 | 37 | 12 |
| Chlorpyrifos-Et-2 | 9.9 | 350.3 | 198.0 | 151 | 10 | 25 | 14 |
| Chlorthiamid-1 | 6.4 | 205.9 | 118.9 | 76 | 10 | 55 | 6.5 |
| Chlorthiamid-2 | 6.4 | 205.9 | 154.0 | 76 | 10 | 45 | 7 |
| Clothianidine-1 | 5.2 | 250.0 | 169.0 | 71 | 10 | 19 | 14 |
| Clothianidine-2 | 5.2 | 250.0 | 131.9 | 71 | 10 | 23 | 12 |
| D10-Simazine-1 | 6.2 | 212.0 | 137.1 | 66 | 10 | 29 | 10 |
| D3-chlothianidin-1 | 5.2 | 252.9 | 131.9 | 116 | 10 | 23 | 10 |
| D3-MeDPC-1 | 3.7 | 162.9 | 88.0 | 96 | 10 | 41 | 8 |
| D3-thiamethoxam-1 | 4.9 | 295.0 | 214.1 | 81 | 10 | 19 | 18 |
| D4-Boscalid-1 | 7.9 | 347.0 | 311.1 | 116 | 10 | 31 | 22 |
| D4-imidacloprid-1 | 5.4 | 260.0 | 213.1 | 61 | 10 | 23 | 14 |
| D5-Atrazine-1 | 6.9 | 221.1 | 179.1 | 121 | 10 | 25 | 16 |
| D5-DIA-1 | 4.7 | 179.0 | 100.9 | 101 | 10 | 27 | 10 |
| D6-chlorpyrifos Me-1 | 9.1 | 327.9 | 130.9 | 76 | 10 | 31 | 10 |
| D6-DEA-1 | 5.3 | 194.1 | 147.1 | 66 | 10 | 27 | 10 |
| D6-Dimethoate-1 | 5.4 | 235.9 | 205.0 | 61 | 10 | 15 | 22 |
| D6-Isoproturon-1 | 6.9 | 213.1 | 78.1 | 86 | 10 | 25 | 10 |
| D6-Metolachlor-1 | 8.5 | 290.1 | 258.2 | 171 | 10 | 21 | 20 |
| D6-Omethoate-1 | 3.7 | 219.9 | 189.0 | 46 | 10 | 15 | 10 |
| D6-Tebuconazole-1 | 8.0 | 314.1 | 72.1 | 101 | 10 | 53 | 10 |
| DEA-1 | 5.3 | 188.1 | 145.9 | 100 | 10 | 25 | 15 |
| DEA-2 | 5.3 | 188.1 | 104.0 | 100 | 10 | 33 | 6.5 |
| DEA-OH 1 | 2.6 | 170.0 | 128.0 | 60 | 10 | 31 | 13 |
| DEA-OH 2 | 2.6 | 170.0 | 86.0 | 60 | 10 | 23 | 13 |
| DEDIA-1 | 2.8 | 146.0 | 78.9 | 106 | 10 | 25 | 10 |
| DEDIA-2 | 2.8 | 148.1 | 81.0 | 106 | 10 | 25 | 10 |
| Desmethyl isoproturon-1 | 6.5 | 193.0 | 136.0 | 40 | 10 | 22 | 13 |
| Desmethyl isoproturon-2 | 6.5 | 193.0 | 94.0 | 40 | 10 | 32 | 13 |
| DIA-1 | 4.7 | 174.0 | 103.9 | 106 | 10 | 31 | 12 |
| DIA-2 | 4.7 | 174.0 | 68.0 | 106 | 10 | 37 | 8 |
| DIA-OH-1 | 2.1 | 156.1 | 69.0 | 61 | 10 | 35 | 32 |
| DIA-OH-2 | 2.1 | 156.1 | 86.1 | 61 | 10 | 21 | 8 |
| Didesmethyl isoproturon-1 | 6.1 | 179.0 | 137.1 | 106 | 10 | 19 | 10 |
| Didesmethyl isoproturon-2 | 6.1 | 179.0 | 94.1 | 106 | 10 | 31 | 8 |
| Dimethachlor-1 | 7.3 | 256.0 | 224.2 | 64 | 10.5 | 19 | 6.5 |
| Dimethachlor-2 | 7.3 | 256.0 | 148.0 | 64 | 10 | 33 | 6 |
| Dimethenamid-1 | 7.9 | 275.9 | 244.1 | 16 | 10 | 19 | 16 |
| Dimethenamid-2 | 7.9 | 275.9 | 168.0 | 16 | 10 | 35 | 20 |
| Dimethoate-1 | 5.5 | 230.0 | 198.9 | 51 | 10 | 13 | 14 |
| Dimethoate-2 | 5.5 | 230.0 | 124.9 | 51 | 10 | 29 | 16 |
| DPC-1 | 2.2 | 146.0 | 117.0 | 290 | 10 | 30 | 8 |
| Fenthion oxon sulfone-1 | 5.6 | 295.0 | 217.1 | 101 | 10 | 27 | 14 |
| Fenthion oxon sulfone-2 | 5.6 | 295.0 | 78.0 | 101 | 10 | 53 | 10 |
| Fenthion oxon sulfoxide-1 | 4.9 | 279.0 | 264.0 | 76 | 10 | 25 | 20 |
| Fenthion oxon sulfoxide-2 | 4.9 | 279.0 | 104.0 | 76 | 10 | 37 | 10 |
| Fenthion oxon-1 | 7.0 | 262.9 | 231.0 | 116 | 10 | 23 | 18 |
| Fenthion oxon-2 | 7.0 | 262.9 | 216.0 | 116 | 10 | 33 | 14 |
| Fenthion sulfone-1 | 7.2 | 310.9 | 125.0 | 101 | 10 | 27 | 10 |
| Fenthion sulfone-2 | 7.2 | 310.9 | 139.0 | 101 | 10 | 37 | 10 |
| Fenthion sulfoxide-1 | 6.4 | 295.0 | 280.0 | 111 | 10 | 25 | 20 |
| Fenthion sulfoxide-2 | 6.4 | 295.0 | 109.0 | 96 | 10 | 41 | 14 |
| Fenthion-1 | 8.8 | 278.9 | 246.9 | 71 | 10 | 17 | 18 |
| Fenthion-2 | 8.8 | 278.9 | 169.1 | 71 | 10 | 23 | 12 |
| Flufenacet-1 | 8.3 | 364.0 | 152.0 | 91 | 10 | 27 | 14 |
| Flufenacet-2 | 8.3 | 364.0 | 194.1 | 91 | 10 | 15 | 16 |
| Imidacloprid olefin-1 | 4.8 | 254.0 | 205.1 | 56 | 10 | 21 | 18 |
| Imidacloprid olefin-2 | 4.8 | 254.0 | 171.0 | 56 | 10 | 27 | 18 |
| Imidacloprid urea-1 | 4.8 | 211.9 | 127.9 | 101 | 10 | 27 | 10 |
| Imidacloprid urea-2 | 4.8 | 211.9 | 125.9 | 101 | 10 | 35 | 14 |
| Imidacloprid-1 | 5.4 | 255.9 | 175.1 | 66 | 10 | 29 | 14 |
| Imidacloprid-2 | 5.4 | 255.9 | 209.1 | 66 | 10 | 23 | 14 |
| Isoproturon-1 | 6.9 | 207.2 | 72.1 | 86 | 4 | 29 | 4 |
| Isoproturon-2 | 6.9 | 207.2 | 165.2 | 86 | 10.5 | 19 | 6.5 |
| MeDPC-1 | 3.7 | 160.0 | 87.9 | 61 | 10 | 39 | 10 |
| MeDPC-2 | 3.7 | 160.0 | 130.0 | 61 | 10 | 33 | 12 |
| Metazachlor-1 | 7.3 | 278.1 | 210.1 | 58 | 3.5 | 15 | 4 |
| Metazachlor-2 | 7.3 | 278.1 | 134.2 | 58 | 3.5 | 27 | 4 |
| Metazachlor OXA-1 | 5.1 | 274.0 | 134.3 | 71 | 10 | 27 | 8 |
| Metazachlor OXA-2 | 5.1 | 274.0 | 162.0 | 71 | 10 | 11 | 10 |
| Metsulfuron-methyl-1 | 6.4 | 382.0 | 167.0 | 86 | 10 | 23 | 14 |
| Metsulfuron-methyl-2 | 6.4 | 382.0 | 141.0 | 86 | 10 | 21 | 10 |
| MMTA-1 | 2.3 | 140.9 | 57.0 | 66 | 10 | 25 | 10 |
| MMTA-2 | 2.3 | 140.9 | 85.0 | 66 | 10 | 23 | 10 |
| Omethoate-1 | 3.7 | 213.9 | 183.0 | 66 | 10 | 15 | 16 |
| Omethoate-2 | 3.7 | 213.9 | 108.9 | 66 | 10 | 37 | 14 |
| Prosulfocarb-1 | 9.5 | 252.3 | 91.1 | 71 | 10 | 35 | 14 |
| Prosulfocarb-2 | 9.5 | 252.3 | 128.1 | 71 | 10 | 17 | 14 |
| Metolachlor-1 | 8.5 | 284.1 | 252.1 | 101 | 10 | 21 | 18 |
| Metolachlor-2 | 8.5 | 284.1 | 176.2 | 101 | 10 | 35 | 14 |
| TCP-1 | 7.1 | 197.9 | 107.0 | 56 | 10 | 43 | 8 |
| TCP-2 | 7.1 | 197.9 | 134.0 | 56 | 10 | 31 | 10 |
| Tebuconazole-1 | 8.0 | 308.1 | 70.1 | 61 | 10 | 39 | 4 |
| Tebuconazole-2 | 8.0 | 308.1 | 125.0 | 61 | 10 | 51 | 4 |
| Terbumeton desethyl-1 | 4.8 | 198.1 | 142.1 | 36 | 10 | 23 | 10 |
| Terbumeton desethyl-2 | 4.8 | 198.1 | 86.0 | 36 | 10 | 33 | 10 |
| Terbumeton-1 | 5.7 | 226.0 | 170.2 | 56 | 11.5 | 23 | 6.5 |
| Terbumeton-2 | 5.7 | 226.0 | 114.0 | 56 | 10 | 35 | 6 |
| Terbuthylazine-desethyl-OH-1 | 3.8 | 184.0 | 128.1 | 81 | 10 | 21 | 10 |
| Terbuthylazine-desethyl-OH-2 | 3.8 | 184.0 | 86.1 | 81 | 10 | 35 | 10 |
| Terbuthylazine desethyl-1 | 6.3 | 202.0 | 146.0 | 46 | 10 | 23 | 16 |
| Terbuthylazine desethyl-2 | 6.3 | 202.0 | 103.9 | 46 | 10 | 37 | 12 |
| Terbuthylazine hydroxy-1 | 4.7 | 212.1 | 156.1 | 51 | 10 | 21 | 12 |
| Terbuthylazine hydroxy-2 | 4.7 | 212.1 | 69.0 | 51 | 10 | 49 | 8 |
| Terbuthylazine-1 | 7.7 | 230.1 | 174.1 | 81 | 11.5 | 23 | 6.5 |
| Terbuthylazine-2 | 7.7 | 230.1 | 104.1 | 86 | 7.5 | 41 | 4 |
| Terbutryn-1 | 7.1 | 242.0 | 186.1 | 41 | 10 | 27 | 14 |
| Terbutryn-2 | 7.1 | 242.0 | 68.0 | 41 | 10 | 61 | 8 |
| Thiamethoxam-1 | 4.9 | 292.0 | 211.1 | 66 | 10 | 17 | 16 |
| Thiamethoxam-2 | 4.9 | 292.0 | 180.9 | 66 | 10 | 31 | 22 |
| TZMU-1 | 4.5 | 205.9 | 131.9 | 61 | 10 | 23 | 16 |
| TZMU-2 | 4.5 | 205.9 | 174.9 | 61 | 10 | 25 | 18 |

RT: retention time (min) – Q1: parent ion – Q3: fragment ion – DP: declustering potential – EP: entrance potential – CE: collision energy – CXP: cell exit potential. Transition of quantification: -1; transition of confirmation: -2.

Table B: MS/MS fragmentation parameters and retention times of the direct injection LC-MS/MS method for compounds analysed in MRM ESI-, first run

| ID | RT (min) | Q1 | Q3 | DP (V) | EP (V) | CE (V) | CXP (V) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 13C-15N-Fipronil - 1 | 8.5 | 441.0 | 336.0 | -30 | -10 | -24 | -17 |
| 13C6-Saccharin - 1 | 4.1 | 187.9 | 42.0 | -130 | -10 | -54 | -15 |
| Acetochlor OXA-1 | 5.6 | 264.0 | 146.2 | -30 | -10 | -14 | -15 |
| Acetochlor OXA-2 | 5.6 | 264.0 | 144.3 | -30 | -10 | -34 | -15 |
| Alachlor OXA-1 | 5.6 | 264.0 | 159.9 | -15 | -10 | -16 | -15 |
| Alachlor OXA-2 | 5.6 | 264.0 | 158.1 | -15 | -10 | -36 | -13 |
| Alachlor-acetochlor ESA-1 | 5.4 | 314.0 | 80.0 | -135 | -10 | -52 | -15 |
| Alachlor-acetochlor ESA-2 | 5.4 | 314.0 | 120.9 | -135 | -10 | -30 | -13 |
| Bentazon-1 | 6.8 | 239.0 | 132.0 | -100 | -10 | -38 | -4 |
| Bentazon-2 | 6.8 | 239.0 | 175.0 | -100 | -10 | -28 | -6 |
| CGA 357704-1 | 5.5 | 278.0 | 206.0 | -45 | -10 | -18 | -15 |
| CGA 357704-2 | 5.5 | 278.0 | 162.0 | -45 | -10 | -26 | -15 |
| CGA 368208-1 | 4.7 | 256.0 | 79.8 | -145 | -10 | -58 | -11 |
| CGA 368208-2 | 4.7 | 256.0 | 120.8 | -145 | -10 | -28 | -11 |
| CGA 50720-1 | 5.4 | 206.0 | 134.1 | -60 | -10 | -14 | -11 |
| D3-2,4 D -1 | 6.9 | 221.9 | 163.8 | -25 | -10 | -20 | -17 |
| D3-alachlor ESA-1 | 5.4 | 316.9 | 79.8 | -55 | -10 | -66 | -11 |
| D3-alachlor OA-1 | 5.6 | 266.9 | 160.0 | -45 | -10 | -16 | -13 |
| D3-Mecoprop-1 | 7.3 | 216.0 | 144.0 | -60 | -10 | -18 | -11 |
| D5-acetochlor ESA-1 | 5.4 | 319.0 | 79.9 | -135 | -10 | -66 | -11 |
| D5-acetochlor OA-1 | 5.6 | 269.0 | 146.0 | -50 | -10 | -16 | -11 |
| D6-Bentazon-1 | 6.8 | 245.0 | 131.9 | -40 | -10 | -36 | -15 |
| Dimethachlor ESA-1 | 4.8 | 300.0 | 80.0 | -25 | -10 | -66 | -13 |
| Dimethachlor ESA-2 | 4.8 | 300.0 | 120.9 | -25 | -10 | -30 | -15 |
| Dimethachlor OXA-1 | 5.5 | 250.0 | 178.0 | -50 | -10 | -26 | -19 |
| Dimethachlor OXA-2 | 5.5 | 250.0 | 146.0 | -50 | -10 | -22 | -15 |
| Dimethenamid ESA-1 | 5.0 | 319.9 | 80.0 | -50 | -10 | -66 | -9 |
| Dimethenamid ESA-2 | 5.0 | 319.9 | 120.7 | -50 | -10 | -30 | -17 |
| Dimethenamid OXA-1 | 5.9 | 269.9 | 198.0 | -50 | -10 | -16 | -7 |
| Dimethenamid OXA-2 | 5.9 | 269.9 | 166.0 | -50 | -10 | -24 | -11 |
| Fipronil desulfinyl-1 | 8.6 | 386.9 | 350.9 | -110 | -10 | -20 | -23 |
| Fipronil desulfinyl-2 | 8.6 | 386.9 | 281.9 | -110 | -10 | -44 | -19 |
| Fipronil sulfide-1 | 8.9 | 418.8 | 261.9 | -120 | -10 | -40 | -17 |
| Fipronil sulfide-2 | 8.9 | 418.8 | 382.8 | -120 | -10 | -18 | -27 |
| Fipronil sulfone-1 | 8.8 | 450.8 | 414.9 | -65 | -10 | -24 | -21 |
| Fipronil sulfone-2 | 8.8 | 450.8 | 281.9 | -65 | -10 | -38 | -19 |
| Fipronil-1 | 8.5 | 434.9 | 329.9 | -110 | -10 | -22 | -21 |
| Fipronil-2 | 8.5 | 434.9 | 249.9 | -110 | -10 | -38 | -17 |
| Flufenacet ESA-1 | 4.9 | 274.1 | 121.0 | -25 | -10 | -28 | -11 |
| Flufenacet ESA-2 | 4.9 | 274.1 | 79.9 | -25 | -10 | -54 | -29 |
| Flufenacet OXA -1 | 5.1 | 224.0 | 152.0 | -20 | -10 | -15 | -10 |
| Flufenacet OXA -2 | 5.1 | 152.0 | 136.0 | -90 | -8 | -17 | -10 |
| MCPA-1 | 6.9 | 199.0 | 141.0 | -65 | -10 | -20 | -12 |
| MCPA-2 | 6.9 | 201.0 | 143.0 | -65 | -10 | -20 | -12 |
| Metazachlor ESA-1 | 4.8 | 322.0 | 77.0 | -65 | -10 | -52 | -11 |
| Metazachlor ESA-2 | 4.8 | 322.0 | 79.6 | -65 | -10 | -76 | -19 |
| Metolachlor ESA-1 | 5.3 | 328.1 | 79.8 | -135 | -10 | -72 | -11 |
| Metolachlor ESA-2 | 5.3 | 328.1 | 121.0 | -135 | -10 | -30 | -13 |
| Metolachlor OXA-1 | 6.7 | 278.1 | 206.1 | -65 | -10 | -14 | -13 |
| Metolachlor OXA-2 | 6.7 | 278.1 | 174.1 | -65 | -10 | -22 | -15 |
| Saccharin-1 | 4.1 | 181.9 | 41.9 | -150 | -10 | -58 | -19 |
| Saccharin-2 | 4.1 | 181.9 | 105.9 | -150 | -10 | -26 | -13 |
| Metolachlor NOA-1 | 5.1 | 327.9 | 205.9 | -30 | -10 | -22 | -17 |
| Metolachlor NOA-2 | 5.1 | 327.9 | 256.0 | -30 | -10 | -28 | -21 |

RT: retention time (min) – Q1: parent ion – Q3: fragment ion – DP: declustering potential – EP: entrance potential – CE: collision energy – CXP: cell exit potential. Transitions of quantification: -1; transitions of confirmation: -2.

Table C: MS/MS fragmentation parameters and retention times of the direct injection LC-MS/MS method for compounds analysed in MRM ESI+, second run

| ID | RT (min) | Q1 | Q3 | DP (V) | EP (V) | CE (V) | CXP (V) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 13C-DEDIA-1 | 2.6 | 149.0 | 105.8 | 16 | 10 | 27 | 14 |
| Acetamiprid-N-desmethyl-1 | 5.1 | 209.1 | 126.0 | 76 | 10 | 27 | 22 |
| Acetamiprid-N-desmethyl-2 | 5.1 | 211.0 | 128.0 | 76 | 10 | 13 | 22 |
| Alachlor-1 | 8.1 | 270.0 | 238.0 | 20 | 10 | 15 | 12 |
| Alachlor-2 | 8.1 | 270.0 | 162.0 | 20 | 10 | 28 | 12 |
| Atrazine-2-hydroxy-1 | 4.6 | 198.1 | 156.0 | 71 | 10 | 25 | 12 |
| Atrazine-2-hydroxy-2 | 4.6 | 198.1 | 114.0 | 71 | 10 | 31 | 20 |
| Carbofuran-1 | 6.4 | 222.1 | 165.2 | 56 | 9.5 | 17 | 6.5 |
| Carbofuran-2 | 6.4 | 222.1 | 123.0 | 7 | 10 | 29 | 2 |
| Carbofuran-3-keto-1 | 5.9 | 236.0 | 179.0 | 31 | 10 | 17 | 14 |
| Carbofuran-3-keto-2 | 5.9 | 236.0 | 161.1 | 31 | 10 | 23 | 14 |
| Chlorotoluron-desmethyl-1 | 6.3 | 199.0 | 106.0 | 66 | 10 | 55 | 12 |
| Chlorotoluron-desmethyl-2 | 6.3 | 199.0 | 107.0 | 66 | 10 | 33 | 12 |
| Chlorpyrifos-Me-1 | 9.0 | 323.9 | 291.8 | 31 | 10 | 23 | 22 |
| Chlorpyrifos-Me-2 | 9.0 | 323.9 | 125.1 | 31 | 10 | 31 | 8 |
| CPO-1 | 7.8 | 335.9 | 279.8 | 61 | 10 | 25 | 20 |
| CPO-2 | 7.8 | 335.9 | 199.9 | 61 | 10 | 43 | 14 |
| D10-Simazine-1 | 6.3 | 212.0 | 137.1 | 66 | 10 | 29 | 10 |
| D3-chlothianidin-1 | 5.2 | 252.9 | 131.9 | 116 | 10 | 23 | 10 |
| D3-MeDPC-1 | 3.5 | 162.9 | 88.0 | 96 | 10 | 41 | 8 |
| D3-thiamethoxam-1 | 4.9 | 295.0 | 214.1 | 81 | 10 | 19 | 18 |
| D4-Boscalid-1 | 8.0 | 347.0 | 311.1 | 120 | 10 | 47 | 13 |
| D4-imidacloprid-1 | 5.4 | 260.0 | 213.1 | 61 | 10 | 23 | 14 |
| D4-phtalimide-1 | 5.2 | 152.0 | 106.1 | 91 | 10 | 35 | 12 |
| D5-Atrazine-1 | 6.9 | 221.1 | 179.1 | 121 | 10 | 25 | 16 |
| D5-DIA-1 | 4.7 | 179.0 | 100.9 | 101 | 10 | 27 | 10 |
| D6-chlorpyrifos Me-1 | 9.1 | 327.9 | 130.9 | 76 | 10 | 31 | 10 |
| D6-DEA-1 | 5.3 | 194.1 | 147.1 | 66 | 10 | 27 | 10 |
| D6-Dimethoate-1 | 5.5 | 235.9 | 205.0 | 61 | 10 | 15 | 22 |
| D6-DMS-1 | 2.4 | 131.0 | 114.0 | 90 | 10 | 17 | 8 |
| D6-Isoproturon-1 | 6.9 | 213.1 | 78.1 | 86 | 10 | 25 | 10 |
| D6-metalaxyl-1 | 6.9 | 286.0 | 198.5 | 66 | 10 | 41 | 16 |
| D6-Metolachlor-1 | 8.5 | 290.1 | 258.2 | 171 | 10 | 21 | 20 |
| D6-Omethoate-1 | 3.5 | 219.9 | 189.0 | 46 | 10 | 15 | 10 |
| D6-Tebuconazole-1 | 8.0 | 314.1 | 72.1 | 101 | 10 | 53 | 10 |
| D7-DEET-1 | 6.9 | 199.0 | 125.9 | 81 | 10 | 25 | 12 |
| DEET-1 | 6.9 | 192.0 | 119.0 | 41 | 10 | 25 | 8 |
| DEET-2 | 6.9 | 192.0 | 91.1 | 41 | 10 | 39 | 8 |
| DMS-1 | 2.4 | 125.0 | 107.9 | 81 | 10 | 15 | 10 |
| DMS-2 | 2.4 | 125.0 | 44.2 | 81 | 10 | 41 | 8 |
| DMST-1 | 6.9 | 215.0 | 106.0 | 116 | 10 | 21 | 8 |
| DMST-2 | 6.9 | 215.0 | 77.0 | 116 | 10 | 59 | 10 |
| Epoxyconazole-1 | 7.8 | 330.0 | 121.1 | 61 | 10 | 27 | 6 |
| Epoxyconazole-2 | 7.8 | 330.0 | 101.0 | 61 | 10 | 65 | 12 |
| Fipronil Carboxamide-1 | 7.1 | 455.9 | 386.9 | 120 | 10 | 47 | 13 |
| Fipronil Carboxamide-2 | 7.1 | 455.9 | 256.0 | 120 | 10 | 47 | 13 |
| IM-1-4 -1 | 2.5 | 156.9 | 126.0 | 71 | 10 | 21 | 10 |
| IM-1-4 -2 | 2.5 | 156.9 | 90.1 | 71 | 10 | 35 | 8 |
| Imidacloprid desnitro-1 | 3.9 | 210.9 | 126.0 | 161 | 10 | 33 | 10 |
| Imidacloprid desnitro-2 | 3.9 | 210.9 | 99.0 | 161 | 10 | 55 | 12 |
| IN-KF313-1 | 5.4 | 249.0 | 124.1 | 56 | 10 | 39 | 10 |
| IN-KF313-2 | 5.4 | 249.0 | 167.0 | 56 | 10 | 19 | 10 |
| Lenacile-1 | 6.4 | 235.0 | 153.0 | 74 | 10 | 21 | 4 |
| Lenacile-2 | 6.4 | 235.0 | 136.0 | 74 | 10 | 43 | 4 |
| LM2-1 | 2.0 | 214.1 | 126.0 | 91 | 10 | 31 | 10 |
| LM2-2 | 2.0 | 214.1 | 168.0 | 91 | 10 | 23 | 14 |
| LM3-1 | 2.1 | 199.0 | 129.0 | 91 | 10 | 27 | 10 |
| LM3-2 | 2.1 | 199.0 | 138.2 | 91 | 10 | 27 | 10 |
| LM4-1 | 3.6 | 242.1 | 196.1 | 96 | 10 | 25 | 14 |
| LM4-2 | 3.6 | 242.1 | 126.1 | 96 | 10 | 33 | 10 |
| LM5-1 | 4.0 | 185.0 | 129.0 | 86 | 10 | 17 | 16 |
| LM5-2 | 4.0 | 185.0 | 86.0 | 86 | 10 | 33 | 12 |
| LM6-1 | 4.4 | 199.0 | 143.0 | 111 | 10 | 19 | 12 |
| LM6-2 | 4.4 | 199.0 | 57.0 | 111 | 10 | 29 | 14 |
| Metalaxyl CGA 62826-1 | 6.0 | 266.0 | 219.9 | 71 | 10 | 17 | 16 |
| Metalaxyl CGA 62826-2 | 6.0 | 266.0 | 192.1 | 71 | 10 | 23 | 14 |
| Metalaxyl-1 | 6.9 | 280.0 | 220.1 | 66 | 10 | 19 | 20 |
| Metalaxyl-2 | 6.9 | 280.0 | 192.2 | 66 | 10 | 25 | 14 |
| Methabenzthiazuron-1 | 6.6 | 222.0 | 150.0 | 71 | 10 | 45 | 10 |
| Methabenzthiazuron-2 | 6.6 | 222.0 | 165.1 | 71 | 10 | 25 | 8 |
| Methabenzthiazuron-3 | 6.6 | 222.0 | 95.9 | 71 | 10 | 69 | 12 |
| Pethoxamid MET101-1 | 7.0 | 310.0 | 116.0 | 71 | 10 | 29 | 8 |
| Pethoxamid MET101-2 | 7.0 | 310.0 | 162.0 | 71 | 10 | 17 | 12 |
| Pethoxamide-1 | 8.3 | 296.0 | 131.0 | 46 | 10 | 27 | 12 |
| Pethoxamide-2 | 8.3 | 296.0 | 91.0 | 46 | 10 | 53 | 8 |
| Phthalamic acid-1 | 3.5 | 166.0 | 149.0 | 65 | 10 | 15 | 15 |
| Phthalamic acid-2 | 3.5 | 166.0 | 93.0 | 65 | 10 | 35 | 12 |
| Phthalamic acid-3 | 3.5 | 166.0 | 65.0 | 65 | 10 | 42 | 12 |
| Phthalic acid-1 | 4.5 | 167.0 | 149.0 | 86 | 10 | 13 | 14 |
| Phthalic acid-2 | 4.5 | 167.0 | 64.9 | 86 | 10 | 39 | 10 |
| Phthalic acid-3 | 4.5 | 167.0 | 121.0 | 86 | 10 | 27 | 14 |
| Phthalic acid-4 | 4.5 | 167.0 | 92.9 | 86 | 10 | 33 | 10 |
| Phthalimide-1 | 5.5 | 147.9 | 102.0 | 90 | 10 | 33 | 14 |
| Phthalimide-2 | 5.2 | 147.9 | 76.0 | 90 | 10 | 60 | 8 |
| Pinoxaden M3-1 | 6.8 | 333.1 | 314.9 | 76 | 10 | 17 | 12 |
| Pinoxaden M3-2 | 6.8 | 314.9 | 155.0 | 156 | 10 | 65 | 12 |
| Pinoxaden-1 | 8.3 | 401.2 | 317.1 | 66 | 10 | 29 | 22 |
| Pinoxaden-2 | 8.3 | 401.2 | 57.0 | 66 | 10 | 55 | 8 |
| Sulfoxaflor-1 | 5.9 | 278.0 | 174.0 | 96 | 10 | 15 | 20 |
| Sulfoxaflor-2 | 5.9 | 278.0 | 154.0 | 96 | 10 | 39 | 10 |
| Thiacloprid-1 | 5.8 | 253.0 | 125.9 | 121 | 10 | 29 | 18 |
| Thiacloprid-2 | 5.8 | 253.0 | 90.0 | 121 | 10 | 55 | 16 |
| THPI-1 | 4.7 | 151.9 | 78.9 | 120 | 10 | 47 | 13 |
| THPI-2 | 4.7 | 151.9 | 81.0 | 120 | 10 | 47 | 13 |
| X11519540-1 | 6.0 | 254.0 | 175.1 | 66 | 10 | 25 | 14 |
| X11519540-2 | 6.0 | 254.0 | 174.0 | 66 | 10 | 39 | 12 |
| X11579457-1 | 4.9 | 253.0 | 174.0 | 41 | 10 | 11 | 18 |
| X11579457-2 | 4.9 | 253.0 | 154.1 | 41 | 10 | 35 | 10 |
| X11719474-1 | 5.1 | 296.0 | 174.0 | 81 | 10 | 17 | 14 |
| X11719474-2 | 5.1 | 296.0 | 61.8 | 81 | 10 | 45 | 8 |

RT: retention time (min) – Q1: parent ion – Q3: fragment ion – DP: declustering potential – EP: entrance potential – CE: collision energy – CXP: cell exit potential. Transitions of quantification: -1; transitions of confirmation: -2.

Table D: MS/MS fragmentation parameters and retention times of the direct injection LC-MS/MS method for compounds analysed in MRM ESI-, second run

| ID | RT (min) | Q1 | Q3 | DP (V) | EP (V) | CE (V) | CXP (V) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2,4-D-1 | 6.8 | 220.9 | 163.0 | -65 | -10 | -20 | -14 |
| 2,4-D-2 | 6.8 | 218.9 | 161.0 | -65 | -10 | -20 | -14 |
| Boscalid M510F01-1 | 6.6 | 359.0 | 246.0 | -160 | -10 | -26 | -17 |
| Boscalid M510F01-2 | 6.6 | 359.0 | 243.9 | -160 | -10 | -28 | -19 |
| Chlorothalonil R417888-1 | 4.6 | 326.9 | 219.8 | -35 | -10 | -34 | -15 |
| Chlorothalonil R417888-2 | 4.6 | 326.9 | 283.8 | -35 | -10 | -26 | -21 |
| D3-2,4 D -1 | 6.8 | 221.9 | 163.8 | -25 | -10 | -20 | -17 |
| D3-alachlor ESA-1 | 5.5 | 316.9 | 79.8 | -55 | -10 | -66 | -11 |
| D3-alachlor OA-1 | 5.7 | 266.9 | 160.0 | -45 | -10 | -16 | -13 |
| D3-Mecoprop-1 | 7.4 | 216.0 | 144.0 | -60 | -10 | -18 | -11 |
| D5-acetochlor ESA-1 | 5.5 | 319.0 | 79.9 | -135 | -10 | -66 | -11 |
| D5-acetochlor OA-1 | 5.7 | 269.0 | 146.0 | -50 | -10 | -16 | -11 |
| D6-Bentazon-1 | 6.8 | 245.0 | 131.9 | -40 | -10 | -36 | -15 |
| Dimethachlor CGA 369873-1 | 4.4 | 242.0 | 80.0 | -140 | -10 | -54 | -9 |
| Dimethachlor CGA 369873-2 | 4.4 | 242.0 | 121.0 | -140 | -10 | -26 | -9 |
| Meptyldinocap-1 | 10.2 | 295.0 | 193.0 | -185 | -10 | -38 | -13 |
| Meptyldinocap-2 | 10.2 | 295.0 | 209.0 | -185 | -10 | -42 | -13 |
| Fipronil 13C 15N- 1 | 8.5 | 441.0 | 336.0 | -200 | -10 | -40 | -16 |
| IN-KE121-1 | 4.9 | 247.0 | 151.0 | -180 | -10 | -30 | -11 |
| IN-KE121-2 | 4.9 | 247.0 | 107.0 | -180 | -10 | -42 | -11 |
| Metalaxyl CGA 108906-1 | 5.2 | 294.0 | 203.9 | -30 | -10 | -16 | -15 |
| Metalaxyl CGA 108906-2 | 5.2 | 294.0 | 88.9 | -30 | -10 | -30 | -11 |
| Pethoxamid MET100-1 | 4.9 | 326.0 | 120.8 | -60 | -10 | -28 | -11 |
| Pethoxamid MET100-2 | 4.9 | 326.0 | 203.9 | -60 | -10 | -22 | -17 |
| Pethoxamid MET42-1 | 5.6 | 340.1 | 79.9 | -155 | -10 | -68 | -9 |
| Pethoxamid MET42-2 | 5.6 | 340.1 | 120.9 | -155 | -10 | -32 | -9 |
| Pinoxaden M2-1 | 6.8 | 315.1 | 187.1 | -205 | -10 | -34 | -11 |
| Pinoxaden M2-2 | 6.8 | 315.1 | 172.0 | -205 | -10 | -54 | -11 |
| Chlorothalonil R471811-1 | 2.1 | 344.8 | 301.7 | -65 | -10 | -30 | -19 |
| Chlorothalonil R471811-2 | 2.1 | 346.8 | 303.7 | -150 | -10 | -30 | -19 |
| Chlorothalonil R471811-3 | 2.1 | 344.8 | 158.7 | -65 | -10 | -50 | -17 |
| Chlorothalonil R611965-1 | 4.5 | 265.8 | 222.0 | -75 | -10 | -12 | -17 |
| Chlorothalonil R611965-2 | 4.5 | 265.8 | 42.0 | -75 | -10 | -32 | -5 |
| Saccharin 13C6 - 1 | 3.8 | 187.9 | 42.0 | -130 | -10 | -54 | -15 |
| Chlorothalonil SDS3701-1 | 6.9 | 244.8 | 181.7 | -110 | -10 | -40 | -19 |
| Chlorothalonil SDS3701-2 | 6.9 | 244.8 | 174.9 | -110 | -10 | -36 | -15 |
| Sedaxan-1 | 8.0 | 330.1 | 131.1 | -155 | -10 | -28 | -7 |
| Sedaxan-2 | 8.0 | 330.1 | 90.9 | -155 | -10 | -48 | -11 |
| Sedaxane M1-1 | 4.7 | 174.9 | 90.9 | -45 | -10 | -26 | -11 |
| Sedaxane M2-1 | 3.9 | 160.9 | 141.0 | -30 | -10 | -14 | -9 |
| Sedaxane M2-2 | 3.9 | 160.9 | 66.0 | -30 | -10 | -28 | -9 |
| Chlorothalonil SYN507900-1 | 4.7 | 264.9 | 221.8 | -45 | -10 | -28 | -17 |
| Chlorothalonil SYN507900-2 | 4.7 | 262.8 | 219.8 | -180 | -10 | -28 | -15 |

RT: retention time (min) – Q1: parent ion – Q3: fragment ion – DP: declustering potential – EP: entrance potential – CE: collision energy – CXP: cell exit potential. Transitions of quantification: -1; transitions of confirmation: -2.

- SBSE-GC-MS/MS: stir bar sorptive extraction gas chromatography coupled with tandem mass spectrometry

| SBSE-GC-MS/MS | Operating conditions |
| --- | --- |
| SBSE | Stirring of a PDMS (Polydimethylsiloxane) bar (20mm x 1mm, Gerstel®) in 20 mL of sample, at 600 rpm during 2h00 at room temperature |
| GC-MS/MS analysis | TDU/CIS: - TDU mode: Splitless (30°C during 0.1 min, to 280°C at 60°C/min, hold time 8 min), - CIS mode: Solvent vent (=Splitless) -35°C (cryogenic cooling) for 0.05 min, 12°C/s until 300°C, hold time 2 min Column: VF-5ms 30 m x 0.25 mm, 0.25 µm (Agilent®) Duration of the chromatographic run: 34 min Column pressure (Helium): 71.3 kPa Oven temperature gradient: 65°C for min, 150°C with 25°C/min, 200°C with 3°C/min, 280°C with 8°C/min stay 2 min Electronic ionisation: EI, 70 eV Temperatures: Source: 250°C, transfer line: 280°C |

Table E: MS/MS fragmentation parameters and retention times of the SBSE-GC-MS/MS method

| ID | RT | Q | CE (V) | q1 | CE (V) | q2 | CE (V) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2-Chlorophenol | 6.74 | 128 - 65 | 20 | 128 - 100 | 10 | 130 - 102 | 10 |
| 3-Chlorophenol | 6.97 | 128 - 65 | 20 | 128 - 100 | 10 |  |  |
| 4-Chlorophenol | 7.05 | 128 - 65 | 20 | 128 - 100 | 10 |  |  |
| 4-Chlorophenol-13C | 7.05 | 137 - 70 | 10 |  |  |  |  |
| 2,3-Dichlorophenol | 8.67 | 162 - 63 | 25 | 162 - 126 | 10 |  |  |
| 2,4-Dichlorophenol | 8.28 | 162 - 63 | 25 | 164 - 63 | 10 | 164 - 63 | 25 |
| 2,4-Chlorophenol-D3 | 8.26 | 165 - 66 | 25 |  |  |  |  |
| 2,5-Dichlorophenol | 8.28 | 162 - 63 | 25 | 164 - 63 | 10 | 164 - 63 | 25 |
| 2,6-Dichlorophenol | 8.02 | 162 - 63 | 25 | 162 - 64 | 10 | 164 - 63 | 25 |
| 3,4-Dichlorophenol | 9.05 | 162 - 99 | 10 | 162 - 63 | 25 |  |  |
| 3,5-Dichlorophenol | 8.43 | 162 - 99 | 10 | 162 - 63 | 25 |  |  |
| 2,3-Dichloranisole | 8.03 | 176 - 133 | 20 | 178 - 135 | 20 |  |  |
| 2,4-Dichloroanisole | 7.64 | 176 - 161 | 10 | 176 - 133 | 20 |  |  |
| 2,5-Dichloroanisole | 7.55 | 176 - 133 | 20 | 176 - 161 | 10 |  |  |
| 2,6-Dichloroanisole | 6.82 | 176 - 133 | 20 | 176 - 161 | 10 |  |  |
| 3,4-Dichloroanisole | 7.71 | 176 - 133 | 20 | 176 - 146 | 20 | 178 - 163 | 10 |
| 3,5-Dichoroanisole | 7.26 | 176 - 146 | 20 | 178 - 148 | 20 | 178 - 148 | 15 |
| Dichlobenil | 8.33 | 171 - 100 | 30 | 171 - 136 | 15 | 173 - 100 | 25 |
| Endosulfan I | 25.46 | 195 - 159 | 5 | 239 - 204 | 15 |  |  |
| Endosulfan I-D4 | 25.38 | 199 - 129 | 15 |  |  |  |  |
| Endosulfan II | 27.55 | 195 - 159 | 5 | 239 - 204 | 15 |  |  |
| Endosulfan-ether | 18.70 | 241 - 206 | 15 | 272 - 237 | 15 |  |  |
| Endosulfan-lacton | 28.74 | 321 - 263 | 15 | 277 - 242 | 15 |  |  |
| Endosulfan-sulfate | 23.94 | 272 - 237 | 10 | 387 - 289 | 5 |  |  |

RT: retention time (min) – Q: MS/MS quantitative transition – q: MS/MS qualitative transitions – CE: collision energy