- Supplementary Information

Supplementary Table S1: Typical measured characteristics of tuned waveforms conditions prior to pulse shaping. The untuned characteristics are used to determine the voltage envelope for shaping the current waveform. The resonant frequencies, rampup time constants in milliseconds and number of cycles for each resonant frequency, in addition to the series resistance at resonance are presented. Resonant frequency and rampup time constants are obtained by fitting I(t)=I0(1-exp(-t/τ))sin⁡(2πfrest+ϕ) over I0, τ, fres, and ϕ. Rseries is obtained by additionally fitting Vt=V0sin⁡(2πfrest+ϕ) to the applied voltage waveforms, and computing Rseries=V0/I0 assuming real impedance at resonance.

Supplementary Figure S2: (a) Sites reported by subjects during PNS experiments mapped to generic head model. Black dots represent a site reported by any subject across all frequencies at any point during a titration. (b) Distribution of stimulation sites across subjects for each frequency condition. In some cases, subjects did not report stimulation, and the corresponding subject marker is absent from the head model.

Supplementary Figure S3: (a) Bland-Altman plot of PNS threshold pairs measured in 5 subjects at 1.8 kHz. (b) Bland-Altman plot of PNS threshold pairs measured in 5 subjects at 25.3 kHz.

Supplementary Figure S4: Error function estimation of the mean PNS threshold across subjects for each frequency. Each point presents the threshold measured for a single subject. In the top-left most pane, each subject data point is labeled with subject number to demonstrate construction of the CDF for each frequency (ordering changes across panes based on the field threshold peak for each frequency). For some subjects and frequencies, the PNS threshold could not be reached within the amplifier limits. We constructed CDFs assuming normally distributed thresholds across subjects and fitted the corresponding mean and standard deviation to impute missing data points. Note that for 88.1 kHz, only 3 out of 8 subjects reported stimulation, resulting in a relatively large standard deviation about the estimated mean (mu=7.09 mT, sigma=1.53 mT).

Pulse Shaping

- LCR Second Order System Analysis

In the resonant coil case, the series LCR circuit is modeled as a second order system. The dynamics of the LCR circuit can be expressed as:

vt=Lditdt+1C-∞tit'dt'+Ri(t)

Taking the derivative with respect to t, and dividing through by L:

1Ldvtdt=d2itdt2+RL ditdt+1LCi(t)

Taking the Laplace transform of the above equation, assuming all zero initial conditions yields:

1LsVs=s2Is+RLsIs+1LCI(s)

Rearranging the above, and substituting α=R/2L, and ω0=1/LC, the circuit’s voltage to current impulse response can be expressed as

Hnats=IsV(s)=1Lss2+2αs+ω02

where the subscript “nat” refers to the natural impulse response to a step input into the circuit.

To determine the response to a sinusoidal step voltage input such as vsrct=V0sinωtθ(t), we compute the Laplace transform of the time domain signal:

Vsrcs=LV0sinω0tθt=V0ω0s2+ω02

and multiply by the natural impulse response above. Here, θt denotes the Heaviside step function. Taking the inverse Laplace transform of this result:

Inats=VsrcsHnats=V0Lω0s2+ω02ss2+2αs+ω02

inatt=L-1Inats=V02αLsinωt-ω0-α2+ω02sin-α2+ω02 te-αtθ(t)

In the limiting case where α≪ω0, the radical terms approximately simplify as -α2+ω02≈ω0. After substitution of α=R/2L then the result above simplifies to:

inatt≈V02αL1-e-αtsin⁡(ω0t)θ(t)

=VR1-e-αtsinω0tθ(t)

where we observe that the input sinusoid is modulated by an exponentially decaying envelope function with time constant determined by α, and the steady state amplitude is the applied voltage signal amplitude divided by the resistance of the LCR circuit. There is no phase shift associated with the current waveform relative to the voltage waveform. From an intuitive point of view, operating the LCR circuit at resonance results in cancellation of the imaginary impedances presented by the inductor and capacitor. This results in only the real resistance of the circuit apparent to the amplifier, accounting for no phase in the current waveform relative to the voltage waveform, and further attenuation (amplification) of the voltage amplitude by the resistance value. The exponential envelope relates to the time required to ramp up the energy stored in the electric and magnetic fields in the capacitor and inductor respectively to the full oscillation amplitude.

- Tuned (Resonant) Pulse Shaping

Suppose we desire a current response of the form

idest=V02αLsinωt-ω0-γ2+ω02sin-γ2+ω02 te-γtθ(t)

=γαV02γLsinωt-ω0-γ2+ω02sin-γ2+ω02 te-γtθ(t)

which has the steady state behavior of the natural response (as t→∞, the magnitude of the sinusoid approaches V0/2αL=V0/R) but has the envelope characteristics governed by the time constant related to γ rather than α. This can be seen by again assuming that γ≪ω0, such that -γ2+ω02≈ω0. After applying this approximation, the desired current takes the form

idest≈V0R1-e- γtsin⁡(ω0t)θ(t)

Note that this system has impulse response given by:

Hdess=γα1Lss2+2γs+ω02

where the prefactor γ/α accounts for the difference in the steady state amplitude of the waveform. Our goal is to compute an envelope function to apply to the voltage waveform, such that the natural current waveform is shaped to achieve this desired waveform without feedback control. From the previous section, we saw that the natural current waveform in response to a sinusoidal voltage step input vsrct=V0sinω0tθ(t), at the resonant frequency of the LCR circuit is given by

inatt≈V0R1-e-αtsin⁡(ω0t)θ(t)

for α≪ω0. This current response corresponds to system with step response Hnats. We would thus like to compute venv(t) such that the current output of the LCR circuit matches that of the desired current waveform. In the Laplace domain, this relation can be expressed as the following identical expressions:

VenvsHnats=VsrcsEsHnats=VsrcsHdes(s)=Idess

From this analysis, we can inspect the envelope function:

Es=HdessHnats=γαs2+2αs+ω2s2+2γs+ω2

Because the envelope is defined as a multiplication in the Laplace domain, to apply the envelope in the time domain we convolve the computed envelope function with the sinusoidal step function to retrieve the final enveloped voltage function (or simply the inverse Laplace transform of Venvs=VsrcsEs):

venvt=V0α-γ2+ω2sinωt-ωα-γsin-γ2+ω2 te-γtα-γ2+ω2θt

Again, under the assumption that  γ≪ω, we can approximate the following term as -γ2+ω2 ≈ω, simplifying the above to:

venvt≈V0αωsinωt-ωα-γsinωte-γtαωθ(t)

=V01-α-γαe-γtsinωtθ(t)

where the approximate function appears as the original sinusoidal step input modulated by an envelope function:

eapproxt=1-α-γαe-γt=1-τγ-τατγe-t/τγ

where the subscript approx denotes the approximate multiplicative envelope function applied to the original sinusoidal step, which is not equivalent to the inverse Laplace transform of E(s), which is to be convolved with the original input. Note that as t→0, eapproxt→γα, and as t→∞, eapproxt→1. We can glean insights into pulse shaping by inspecting the form of this multiplicative envelope. If we desire the rise time of the LCR current to increase, such that γ>α (recall that the time constants for the envelopes correspond to the inverse of γ, α, such that this is equivalent to τγ<τα, resulting in a faster rise time with smaller time constant), we require an increase in the initial voltage amplitude. However, this voltage amplitude decays with the time constant τγ to the required steady state amplitude corresponding to the proper steady state current amplitude. On the other hand, if we desire to slow down the rise time such that γ<α, we begin with a lower voltage amplitude that rises to the final amplitude with time constant τγ.

These intuitions are valuable for implementation in the hardware. To eliminate the rampup as a confounding variable in experiments, we will choose a constant rise time in terms of number of cycles for each resonant condition. Depending on the natural time constants of the LCR circuits in the system, we must be careful to choose a rise time that is achievable based on the voltage limitation of the amplifiers in use. In practice, we will choose to “slow down” the resonant envelopes to the slowest natural LCR time constant, so that for each condition we do not require additional voltage at the beginning of each pulse. This method can generalize to any desired pulse shape, given adequate capabilities of amplifiers used for driving the circuit.

- Untuned Pulse Shaping

Neglecting parasitic effects, the coil impedance at the output of the amplifier can be expressed as:

Zω=jωL+R

where L is the inductance of the coil, and R is the series resistance of the coil. In general, the series resistance of the coil varies with frequency. Upon applying a sinusoidal voltage waveform, as using an amplifier in voltage control mode:

vt=V0sin⁡(ωt)

the current waveform will also take a sinusoidal response:

it=V0|Zω| sin⁡(ωt-∠Zω)

where Zω=ωL2+R2, and ∠Z(ω)=tan-1(ωL/R). In the limiting case of ωL≫R, note that Zω≈ωL and ∠Zω≈π/2, resulting in the current waveform:

it=V0ωL cos⁡(ωt)

To reduce the impact of pulse envelope shape on the experimental results, an exponential envelope is applied to the voltage waveform to mimic the dynamics of the second order LCR circuit in the tuned case from the previous section. The voltage waveform from the amplifier is then:

venvt=V01-exp-tτdessin⁡(ωt)

where τdes is the desired time constant of the envelope, resembling the natural step response dynamics of the second order LCR circuit in the preceding section (the inverse of γ from the previous section). This rampup envelope also alleviates the burden on the amplifiers – rather than desiring full current amplitude at t=0 as in the case without the voltage envelope, the amplifier can accordingly rampup the current waveform. The current then takes the form:

it=V0|Zω|1-exp-tτdessin⁡(ωt-∠Zω)