Supplementary Information

MicroeElectroThermoforming (ETF): One-step Versatile 3D Shaping of Thin-Film Microelectronics for Enhanced Neural Interfaces

Dong Hyeon Lee1,7, Younghoon Park2,7, Yoon Seo3, Hyunbeen Jeong4, Jong-Mo Seo4,5, Min-Ho Seo3,6, Kyungsik Eom2*, and Joonsoo Jeong3,6*

1 School of Mechanical Engineering, Pusan National University, Busan 46241, Republic of Korea

2 Department of Electronics Engineering, College of Engineering, Pusan National University, Busan 46241, Republic of Korea

3 Department of Information Convergence Engineering, Pusan National University, Yangsan 50612, Republic of Korea

4 Department of Electrical and Computer Engineering, Seoul National University, Seoul 08826, Republic of Korea

5 Biomedical Research Institute, Seoul National University Hospital, Seoul 03080, Republic of Korea

6 School of Biomedical Convergence Engineering, Pusan National University, Yangsan 50612, Republic of Korea

7 These authors contributed equally: Dong Hyeon Lee, Younghoon Park

* Corresponding author: kseom@pusan.ac.kr, joonsoo_jeong@pusan.ac.kr

Keywords: microelectrode arrays, liquid crystal polymer, neural interface, microthermoforming, 3D electrode

Fig. S1: Microfabrication process of planar and microelectrothermoformed (ETF) LCP MEA. (a) Attach an LCP film to a PDMS-coated carrier silicon wafer. (b) Deposit a metal seed layer (Ti 50 nm/Au 50 nm) on the LCP film using E-beam evaporation. (c) Spin-coat a positive photoresist (PR, 8 m thickness) and pattern electroplating mold via photolithography. (d) Electroplate gold layer. (e) Strip PR and remove seed layers by wet etching. (f) Thermally laminate a low-melting-temperature LCP cover layer. (g) Perform UV laser machining for outlining and site opening to complete a (h) planar LCP MEA. (i) Apply ETF to planar LCP MEA for creating 3D structures. (j) Apply “macro”thermoforming to 3D LCP MEA for conformable fitting to the surrounding tissues

Fig. S2: Preparation of a high-temperature epoxy ETF mold via double casting of a 3D-printed mold for enhanced thermal tolerance during ETF. (a) 3D-printed mold is attached on a disposable dish using double-sided tape. (b) PDMS (10:1 ratio) is poured in the dish and cured at 100 ℃ for 30 minutes. (c) Release the PDMS negative mold. (d) High-temperature-resistant epoxy is poured over the PDMS negative mold, which is cured via two steps: first at 120 ℃ for 4 hours, followed by post-curing at 230 ℃ for 16 hours to reinforce its thermal and mechanical strength. (e) Release the epoxy mold. (f) Trim to fit into the socket in the bottom jig.

Fig. S3: A ETF setup for aligning planar MEA with 3D mold and applying uniform pressure. a) Setup consists of a pair of top and bottom metal plates; a metal frame housing a rubber sheet for uniform pressure; another metal frame housing a rubber stamp for pressing the electrode array against mold. b) Magnified view around 3D mold and MEA, showing alignment scheme of planar MEA with 3D mold fixed in bottom metal plate. Precise alignment is secured by matching laser-drilled align holes on planar MEA with align pins on mold.

Fig. S4: “Macro” thermoforming (mETF) process to create a gross eye-conformable curvature of the ETF 3D MEA. a) Schematic representation of the MEA before and after mETF. b) Schematic illustration and c) photographs of mETF process, fixing a planar or 3D MEA into a concave metal mold having 24 mm-diameter of human eyeball, which is heated at 200℃ for 10 min without applying pressure for permanent deformation. Scale bar: 10 mm.

Fig. S5: 3D-printed master mold for 80-m-high ETF. a) CAD drawing and b) photograph of 3D-printed master mold carrying pillars with a diameter of 200 m and a height of 80 m. Scale bars: 1 mm. c) Cross-sectional drawing (top) and surface profile (bottom) of the 3D mold. Fig. S6: ETF for 100 m-diameter MEA. a) 3D mold for ETF MEA for ex vivo calcium imaging carrying mixed protruding and planar electrodes. Scale bars: 200 m. b) SEM image of µETF electrodes and planar electrodes with 100 m diameter. Laser-engraved marks around electrodes serve as location markers during ex vivo experiments. Scale bars: 200 m.

Fig. S7: Design details of 3D molds used for creating versatile 3D ETF structures. a) 3D mold used for protruding or recessed structures with varying heights on a single MEA and their cross-sectional profiles. b) Photograph of corresponding 3D-printed master mold. c) 3D mold used for creating versatile 3D structures including polygons (triangle, rectangle and hexagram), ovals (sunken, plateau, and walled), domes, and S-shaped walls, along with their geometric parameters. d) Photograph of corresponding 3D-printed master mold. Scale bars in a–d: 1 mm.

Fig. S8: Comparison of mechanical stress induced on 4 m-thick gold tracks by ETF using dome-shape mold and plateau-shape. Cross-sectional stress distribution within gold lines during ETF with a) dome-shaped mold and b) plateau-shaped mold, both with 200 m diameter and 80 m height. Stress values were normalized to a maximum value. SEM images of gold electrodes after ETF using c) dome-shaped mold, which resulted in cracks in the center of the electrode, compared to d) intact electrode thermoformed by plateau-shaped mold. Scale bars: 100 m.

Fig. S9: Comparison of mechanical stress induced on gold tracks with in-plane serpentine and out-of-plane serpentine with varying geometric parameters under 15% strain. a) Stress distributions in serpentines with varying w and  Normalized color scales to maximum mechanical stress. Buckling strain represents out-of-plane displacement. b) Comparison of maximum stress on in-plane serpentine and out-of-plane serpentine with varying w and .

Fig. S10: Robustness of gold tracks for 80-m-ETF depending on thickness. a) Representative SEM images of electrodes with varying thicknesses and line shapes after ETF, showing two distinct cracking modes. Scale bars: 200 m. b) Failure rates of gold lines after ETF with respect to gold thicknesses and line shapes.

Fig. S11: Electrochemical characterization of MEA before and after ETF for both gold and IrOx electrodes. a) Impedance magnitude and b) phase, and c) CV curves of planar MEA and 80-m-ETF MEA, confirming preserved electromechanical performance after ETF. Impedance data are presented as mean ± standard deviation. Number of samples for electrochemical impedance: 56 for planar-gold, 31 for protruding-gold, 43 for planar-IrOx, and 63 for protruding IrOx.

Table S1. Electrical parameters of each layer and material in a human retina model used for FEA analysis.

| Structures | Component | Conductivity (S/m) | Thickness (m) |
| --- | --- | --- | --- |
| Microelectrode | Gold | 4.1×107 | - |
|  | LCP | 1e-17 | - |
| Retinal layers | VB | 1.2821 | 60 |
|  | NFL | 0.0126 | 24 |
|  | GCL | 0.0126 | 49 |
|  | IPL | 0.0571 | 40 |
|  | INL | 0.0147 | 38 |
|  | OPL | 0.0571 | 30 |
|  | ONL | 0.0147 | 91 |
|  | OS | 0.10309 | 25 |
|  | RPE | 0.001 | 35 |

Fig. S12: Retinal and MEA models used in FEA for evaluating benefits of protruding MEA in subretinal stimulation. a) Layer structure of human retinal model with planar and protruding MEAs. b) A 3D view and cross-sectional view of 3×3 MEA with variable heights and pitches. The electrical parameters for each model are presented in Table. S1.

Fig. S13: FEA E-field distributions at INL generated by stimulations from MEAs with varying protruding heights. a) Cross-sectional FEA retinal models of MEA with varying protruding heights. Red dashed lines represent INL plane. Transversal E-field distribution at INL induced by stimulation currents of b) Ith (lower limit of DR) and c) Imax (upper limit of DR). At the onset of inter-channel interference (Imax), activated INL areas (contoured by white dashed lines) by a channel overlap with the area activated by adjacent channels.

Fig. S14: FEA E-field distributions at INL generated by 3D ETF MEA with varying heights and pitches. a) 3D model used in FEA with variable geometric parameters. b) E-field profiles at INL induced by simultaneous injection of 10 A stimulating currents from three adjacent electrodes with varying heights (from 0 to 100 m) and electrode pitches (from 300 to 600 m).

Fig. S15: Stimulation parameters and calcium responses from mouse retina. a) Stimulation pulses consisting of a burst of seven cathodic-first biphasic pulses with inter-phase delay. b) Typical calcium transient of an activated RGC, represented by ΔF/F0 = (Fpeak – F0)/F0. Fig. S16: Spatial extent of subretinal stimulation by planar and protruding MEAs (I = 10 μA). Spatial extent of electrical activation of RGCs depending on distances from center of stimulating electrode with a stimulation current, I = 10 μA. The calcium intensity was normalized to maximum fluorescence at electrode center.

Fig. S17: FEA for mechanical stress on the electrode and the surrounding tissue during insertion of the protruding MEAs. a) Comparisons of stress distribution within MEA structures and surrounding tissues, induced by insertion of MEAs with two different materials (silicon and LCP) and edge shapes (sharp and round). Mechanical properties of materials adopted for FEA analysis are summarized in Table S2. Comparison of maximum stress b) within MEA and c) on surrounding tissue.

Table S2. Mechanical properties of materials used for FEA analysis of stress during insertion.

| Material | Density (kg/m3) | Young’s modulus (Pa) | Poisson’s ratio |
| --- | --- | --- | --- |
| Silicon | 2,329 | 170×109 | 0.28 |
| LCP | 1,600 | 3.6×109 | 0.4 |
| Retina | 1,040 | 431×103 | 0.416 |