Dragonfly: A Portable Molecular Diagnostic Platform for Rapid Point-of-Care Detection of Mpox and Other Diseases

Matthew L. Cavuto1,2,&, Kenny Malpartida-Cardenas1,2,&, Ivana Pennisi1,2, Marcus Pond3, Sohail Mirza1,2, Nicholas Moser2,4, Mark Comer2, Isobel Stokes5, Lucy Eke5, Sian Lant5, Katarzyna M. Szostak2, Luca Miglietta1, Oliver W. Stringer1,2, Rebecca P. Sumner5, Frances Bolt1,2, Shiranee Sriskandan1, Alison Holmes1, Pantelis Georgiou2,4, David O. Ulaeto6, Carlos Maluquer de Motes5, Jesus Rodriguez-Manzano1,2,*

1 Department of Infectious Disease, Faculty of Medicine, Imperial College London, London, UK

2 ProtonDx Ltd, Translation & Innovation Hub, Imperial College London, London, UK

3 Department of Infection and Immunity, Imperial College Healthcare NHS Trust, London, UK

4 Department of Electrical and Electronic Engineering, Imperial College London, London, UK

5 Department of Microbial Sciences, School of Biosciences and Medicine, University of Surrey, Guildford, UK

6 CBR Division, Defence Science and Technology Laboratory, Salisbury, UK

& These authors contributed equally

* Corresponding author: j.rodriguez-manzano@imperial.ac.uk

Table of Contents

Table S1. LAMP assays sequences used in this study							2

Table S2. Summary of the LAMP assays used in this study						4

Table S3. Sample-to-result analytical sensitivity performance						5

Table S4. PCR assays sequences used in this study							6

Figure S1. Standard curves of LAMP assays used in this study						7

Figure S2. Confusion matrices of diagnostic performance						8

Figure S3. Histogram showing the distribution of the Ct values obtained from the all the clinical samples	9

Table S1. LAMP assays sequences used in this study.

| Target | LAMP assay ID | Sequence (5' to 3') |
| --- | --- | --- |
| OPXV | LAMP-OPV-E9L-F3 | AACATTTTTGGCAGAGAGAG |
|  | LAMP-OPV-E9L-B3 | ATAGATGGCCTTTTCAGTTGAAC |
|  | LAMP-OPV-E9L-LF | GCTAAGAGTTGCACATCCATAGG |
|  | LAMP-OPV-E9L-LB | GATAACTCTGCTCCATTTAGTACC |
|  | LAMP-OPV-E9L-FIP | ATCTTGTACGTATAYTGCATGGAATC-AAAGATGYTAAAACAGGCTACC |
|  | LAMP-OPV-E9L-BIP | AGTGCTCTATACTCATACGCTTCG-TCTAGATACAAKATCATTCTACGTCC |
|  | LAMP-OPV-F13L-F3 | ATTTGGCCATAGTTCCCAC |
|  | LAMP-OPV-F13L-B3 | CRTCGTCGACTATYAACAAT |
|  | LAMP-OPV-F13L-LF | ACTCCTCTATTAATGGCTGCTT |
|  | LAMP-OPV-F13L-LB | CCGCCAGAAGTCTAGACGC |
|  | LAMP-OPV-F13L-FIP | CCAATTACCRACTARAAGTCTGA-ATTGGCCYGACATTTACAAC |
|  | LAMP-OPV-F13L-BIP | AACGACGTATATTCTATGGCAAC-ACCTTCACAGATAGATCATTTTGRA |
| MPXV | LAMP-MPXV-G2R-F3 | ACGAAAGACTGGATCACAATC |
|  | LAMP-MPXV-G2R-B3 | TCTAAAACAAAGTGTGGAATAGG |
|  | LAMP-MPXV-G2R-LF | CACATCGTGTACTCGGAC |
|  | LAMP-MPXV-G2R-LB | GAAGAGACGGTGTGAGAATATG |
|  | LAMP-MPXV-G2R-FIP | GATGTGGAAATTAACCTGTATCCAGTC-GACGTTGAGATGGATTCGC |
|  | LAMP-MPXV-G2R-BIP | GTATTGCTGGTTACGACGGG-ACGTCATCTGTTCTCCGTG |
|  | LAMP-MPXV-A9L-F3 | GATTGATATCGCATAGAAATAGAAAAG |
|  | LAMP-MPXV-A9L-B3 | CAGCACAAGATCRACCCC |
|  | LAMP-MPXV-A9L-LB | CGTTTGGTAATGGCAATGTATTAAG |
|  | LAMP-MPXV-A9L-FIP | TGTATGAACAGTACTTTGTCGATG-CTCATCATTGAAGATTACTCTGTTAC |
|  | LAMP-MPXV-A9L-BIP | GATGAGATGTTTATATGTTGGCATAGTAG-GATGCTGCTAGACATTATGGAG |
| VZV | LAMP-VZV-ORF62-F3 | TCAGAAGCCTCACATCCTCC |
|  | LAMP-VZV-ORF62-B3 | CGTACTGTACCCCCGAAAC |
|  | LAMP-VZV-ORF62-LF | AGTGGAGGCGCTGCGACGGA |
|  | LAMP-VZV-ORF62-LB | CGCAGGGCGCCAGGCCGTGG |
|  | LAMP-VZV-ORF62-FIP | GACGGTTTGGTCCACCCAGC-TCTGGGATCTGCCGCATC |
|  | LAMP-VZV-ORF62-BIP | GGCGCCGGGATCAAAGCTTA-GGTCGACGACCCATTGTTTC |
|  | LAMP-VZV-ORF28-F3 | CCAATACGACCACCGGATC |
|  | LAMP-VZV-ORF28-B3 | GGACTGGCTTCGTCTCG |
|  | LAMP-VZV-ORF28-LF | CGAAATGTAGGATATAAAGG |
|  | LAMP-VZV-ORF28-LB | CATCGGGAATAACATCCTTATATTC |
|  | LAMP-VZV-ORF28-FIP | CTCCCACTGGTACGTCAAGTG-AGGGTCAAAAACCCTGGC |
|  | LAMP-VZV-ORF28-BIP | CATCTCTTCCTCAACATCCCCG-TACCCGATGGGGGATACC |
| HSV-1 | LAMP-HSV1-UL1-2-F3 | CAGCCACACACCTGTGAA |
|  | LAMP-HSV1-UL1-2-B3 | TCCGTCGAGGCATCGTTAG |
|  | LAMP-HSV1-UL1-2-LF | AAATCCTGTCGCCCTACACAGCGG |
|  | LAMP-HSV1-UL1-2-LB | CACCCCGCGACGGGACGCCG |
|  | LAMP-HSV1-UL1-2-FIP | CCAGACGTTCCGTTGGTAGGTC-ACTTTGACTATTCGCGCACC |
|  | LAMP-HSV1-UL1-2-BIP | CCATCATCGCCACGTCGGAC-TCGGCGTCTGCTTTTTGTG |
| HSV-2 | LAMP-HSV2-US4-F3 | CCGTCAGCCCATCCT |
|  | LAMP-HSV2-US4-B3 | GCCCACCTCTACCCAC |
|  | LAMP-HSV2-US4-LF | GTCTTTGGGGACGGCG |
|  | LAMP-HSV2-US4-LB | GCGGAGACATTCGAGTACC |
|  | LAMP-HSV2-US4-FIP | CCGCCCTGGTACGTGTA-AGTATGGAGGGTGTCGC |
|  | LAMP-HSV2-US4-BIP | TCGTAAATGCTTCCCTGCTG-TCGCCGCCGAGTTC |

Table S2. Summary of the LAMP assays used in this study.

| Assay _ID | Target | Gene | LOD (copies/reaction) | TTP (min) | Ref. |
| --- | --- | --- | --- | --- | --- |
| LAMP-OPV_E9L | OPXV | E9L | 10 | < 10 | This study |
| LAMP-OPV_F13L | OPXV | F13L | 500 | < 15 | This study |
| LAMP-MPXV_G2R | MPXV | G2R | 100 | < 15 | This study |
| LAMP-MPXV_A9L | MPXV | A9L | 500 | < 15 | This study |
| LAMP-VZV_ORF62 | VZV | ORF62 | 100 | < 15 | Okamoto et al. 2004 |
| LAMP-VZV_ORF28 | VZV | ORF28 | 100 | < 15 | This study |
| LAMP-HSV1_UL1 | HSV1 | UL1-2 | 500 | < 15 | Kaneko et al. 2005 |
| LAMP-HSV2_US4 | HSV2 | US4 | 500 | <25 | This study |

Table S3. Input concentrations of viral particles into Dragonfly for sample-to-result analytical sensitivity performance.

| MPXV Viral particles |  |  | HSV1/VZV Viral particles |  |  | HSV2 Viral particles |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ID tube | copies per mL | copies per reaction | ID tube | copies per mL | copies per reaction | ID tube | copies per mL | copies per reaction |
| Tube 1 | 5.00E+04 | 2.00E+03 | Tube 1 | 2.00E+04 | 8.00E+02 | Tube 1 | 2.87E+04 | 1.15E+03 |
| Tube 2 | 1.00E+04 | 4.00E+02 | Tube 2 | 7.00E+03 | 2.80E+02 | Tube 2 | 1.00E+04 | 4.01E+02 |
| Tube 3 | 7.00E+03 | 2.80E+02 | Tube 3 | 2.50E+03 | 1.00E+02 | Tube 3 | 3.58E+03 | 1.43E+02 |
| Tube 4 | 2.50E+03 | 1.00E+02 | Tube 4 | 1.25E+03 | 5.00E+01 | Tube 4 | 1.79E+03 | 7.17E+01 |
| Tube 5 | 1.25E+03 | 5.00E+01 | Tube 5 | 7.50E+02 | 3.00E+01 | Tube 5 | 1.08E+03 | 4.30E+01 |
| Tube 6 | 7.50E+02 | 3.00E+01 |  |  |  |  |  |  |

Table S4. PCR assays sequences used in this study.

| PCR assay ID | Sequence (5′ to 3′) | Reference |
| --- | --- | --- |
| HSV1-F | CGGCCGTGTGACACTATCG | Weidmann et al., 2003 |
| HSV1-R | CTCGTAAAATGGCCCCTCC |  |
| HSV1-P | FAM-CCATACCGACCACACCGACGAACC-TAMRA |  |
| HSV2-F | CGCTCTCGTAAATGCTTCCCT |  |
| HSV2-R | TCTACCCACAACAGACCCACG |  |
| HSV2-P | FAM-CGCGGAGACATTCGAGTACCAGATCG-TAMRA |  |
| VZV-F | CGGCATGGCCCGTCTAT |  |
| VZV-R | TCGCGTGCTGCGGC |  |
| VZV-P | FAM-ATTCAGCAATGGAAACACACGACGCC-TAMRA |  |
| MPXV-WA-F | CACACCGTCTCTTCCACAGA | Li et al., 2010 |
| MPXV-WA-R | GATACAGGTTAATTTCCACATCG |  |
| MPXV-WA-P | FAM-AACCCGTCGTAACCAGCAATACATTT-BHQ1 |  |
| OPXV-E9L-F | TCAACTGAAAAGGCCATCTATGA | Li et al., 2006 |
| OPXV-E9L-R | GAGTATAGAGCACTATTTCTAAATCCC |  |
| OPXV-E9L-P | Cy5-CCATGCAATATACGTACAAGATAGTAGCCAAC-BHQ2 |  |

Figure S1. Standard curves of LAMP assays used in this study.

Figure S2. Confusion matrices of diagnostic performance.

Figure S3. Histogram showing the distribution of the Ct values obtained from the  all the clinical samples, using the TaqMan assay for RNaseP.