Table S1: Validation of the Xpert® SA Nasal Complete for the use of urine as a matrix. Target gene: spa.

| Sample | Day 1 Ct-value | Day 2 Ct-value | Day 3 Ct-value | Robustness (Inter-Assay precision) | Target value |
| --- | --- | --- | --- | --- | --- |
| No_76 | 18.7 | 20.4 | 19.6 | 100% | 100% |
| No_99 | 0.0 | 0.0 | 0.0 | 100% | 100% |

Table S2: Serial dilution (Staphylococcus aureus ATCC 29213) to calculate the limit of detection of the Xpert® SA Nasal Complete for the use of urine as a matrix. Target gene: spa.

| Day 1 |  |  | Day 2 |  |  | Day 3 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CFU /sample | PCR result | Ct-valuea | CFU /sample | PCR result | Ct-valuea | CFU /sample | PCR result | Ct-valuea |
| 0 | neg | 0.0 | - | - | - | - | - | - |
| 9 | neg | 0.0 | - | - | - | - | - | - |
| 23 | pos | 35.1 | 23 | pos | 33.7 | 23 | pos | 35.3 |
| 59 | pos | 35.0 | - | - | - | - | - | - |
| 117 | pos | 33.2 | - | - | - | - | - | - |

aCt values ≤37.4 were interpreted as positive. neg = negative; pos=positive