## Table of contents
Table of contents
Weinschenk et al. Guiding treatment and clinical management of patients with CNS lymphomas by minimal-invasive detection of circulating tumor DNA in cerebrospinal fluid
Supplementary Table 1 | Primer, probes, and PCR conditions used for the design of the MYD88 L265P ddPCR assay.
Supplementary Table 2 | Determining limit of detection for the MYD88 L265P ddPCR assay.
Supplementary Table 3 | Determining limit of blank for the MYD88 L265P ddPCR assay in 21 plasma samples from healthy controls.
Supplementary Table 4 | Determining analytical specificity by analyzing 13 plasma samples from elderly healthy individuals.
Supplementary Table 5 | Plasma samples from CNSL patients with known MYD88 L265P status used for assessment of inter-assay reliability.
Supplementary Table 6 | Plasma samples from CNSL patients used for the assessment of intra-assay reliability.
Supplementary Table 7 | Overview and characteristics of CSF / plasma samples from patients included in the independent validation cohort.
Supplementary Table 8 | Overview and characteristics of CSF samples submitted for routine ctDNA testing by the MYD88 L265P ddPCR assay.
Supplementary Table 9 | Overview and characteristics of CSF samples and patients participating in the observational study.

## Supplementary Table 1
Supplementary Table 1. Primer, probes, and PCR conditions used for the design of the MYD88 L265P ddPCR assay.
MYD88 L265P assay primer/probes
Primer
Gene | Target | Direction | Sequence
MYD88 | L265P | forward | 5'-TGG GCT TGT CCC ACC ATG-3'
|  | reverse | 5'-TGG GGA ACT CTT TCT TCA TTG C-3'
Probes
Gene | Target
MYD88 | L265P | Wildtype | 5'- /5HEX/TGGGG+ATC+AGT+CGC+TTCT/3IABkFQ/ -3'
|  | chr3:38182641 (hg19) T>C | 5'- /56-FAM/TGGGG+ATC+GGT+CGC+TTCT/3IABkFQ/ -3'
PCR conditions
Cycling step | Temperature (Celsius) | Time | Number of cycles
Enzyme activation | 95 | 10 Min | 1
Denaturation | 95 | 30 Sec | 40
Annealing/Extension | 55 | 60 Sec
Enzyme deactivation | 98 | 10 Min | 1
Hold (optional) | 12 | Infinite | 1
ddPCR, digital droplet PCR

## Supplementary Table 2
Supplementary Table 2. Determining limit of detection for the MYD88 L265P ddPCR assay.
Dilution | Expected AF (%) | Measured Ø WT copies / well | Measured Ø MUT copies / well | Measured MUT:WT | Measured AF (%) | SD (MUT copies / well) | p value Unpaired t-test
NTC |  | 0 | 0
WT |  | 18455 | 0.0025
1:20 | 5 | 18510 | 724 | 1:26 | 3.9114 | 23.6643 | <0.0001
1:40 | 2.5 | 18870 | 391.5 | 1:48 | 2.0747 | 28.9266 | <0.0001
1:100 | 1 | 19020 | 191 | 1:100 | 1.0042 | 33.0303 | <0.0001
1:200 | 0.5 | 18995 | 119.5 | 1:159 | 0.62911 | 32.9052 | 0.0008
1:400 | 0.25 | 19245 | 56 | 1:344 | 0.29098 | 11.0454 | 0.0001
1:800 | 0.125 | 18415 | 35.45 | 1:519 | 0.19251 | 11.6665 | 0.0019
1:1000 | 0.1 | 18500 | 16.15 | 1:1145 | 0.087297 | 1.8296 | <0.0001
1:2000 | 0.05 | 17980 | 9.1 | 1:1976 | 0.050612 | 1.8841 | <0.0001
1:10000 | 0.01 | 18532.5 | 1.9 | 1:9754 | 0.010252 | 1.6912 | 0.068
AF, allele frequency
ddPCR, digital droplet PCR
WT, wildtype
MUT, mutated
SD, standard deviation
NTC, non template control

## Supplementary Table 3
Supplementary Table 3. Determining limit of blank for the MYD88 L265P ddPCR assay in 21 plasma samples from healthy controls.
ID | Plasma volume (mL) | Volume elaute (µl) | mL plasma / well | Ø WT copies / well | Ø WT copies / mL plasma | Ø MUT copies / well | Ø MUT copies / mL plasma | Age
NTC |  |  |  | 0 |  | 0
WT |  |  |  | 18455 |  | 0
MIX |  |  |  | 3360 |  | 422.5
HEAL_001 | 3.8 | 68.2 | 0.39003 | 1063 | 2725.4361 | 0 | 0 | 26
HEAL_002 | 3.9 | 67.1 | 0.40686 | 436 | 1071.6337 | 0 | 0 | 27
HEAL_003 | 3.9 | 67.7 | 0.40325 | 216.5 | 536.8883 | 0 | 0 | 28
HEAL_004 | 3.8 | 61.5 | 0.43252 | 200.5 | 463.562 | 0 | 0 | 26
HEAL_005 | 3.9 | 68.4 | 0.39912 | 370 | 927.033 | 0 | 0 | 28
HEAL_006 | 4 | 62.8 | 0.44586 | 306.5 | 687.4357 | 0 | 0 | 25
HEAL_007 | 4 | 60.9 | 0.45977 | 480.5 | 1045.0875 | 0 | 0 | 28
HEAL_008 | 3.8 | 60.8 | 0.4375 | 200.5 | 458.2857 | 0 | 0 | 27
HEAL_009 | 3.9 | 59.6 | 0.45805 | 629.5 | 1374.293 | 0 | 0 | 29
HEAL_010 | 4 | 67.8 | 0.41298 | 813 | 1968.6214 | 0 | 0 | 32
HEAL_011 | 4 | 62.4 | 0.44872 | 1061 | 2364.5143 | 0 | 0 | 23
HEAL_012 | 4 | 62.3 | 0.44944 | 1061 | 2360.725 | 0 | 0 | 20
HEAL_013 | 4 | 68.4 | 0.40936 | 453 | 1106.6143 | 0 | 0 | 22
HEAL_014 | 4 | 43.8 | 0.63927 | 200 | 312.8571 | 0 | 0 | 23
HEAL_015 | 4 | 67.8 | 0.41298 | 569 | 1377.7929 | 0 | 0 | 23
HEAL_016 | 4 | 66.6 | 0.42042 | 1085.5 | 2581.9393 | 0 | 0 | 26
HEAL_017 | 4 | 68.6 | 0.40816 | 515 | 1261.75 | 0 | 0 | 30
HEAL_018 | 4 | 64 | 0.4375 | 636.5 | 1454.8571 | 0.3 | 0.68571 | 23
HEAL_019 | 4 | 69.6 | 0.4023 | 298 | 740.7429 | 0 | 0 | 19
HEAL_020 | 4 | 69.8 | 0.40115 | 288 | 717.9429 | 0 | 0 | 20
HEAL_021 | 4 | 65.7 | 0.42618 | 408 | 957.3429 | 0 | 0 | 19
Mean | [5 empty] | 0.014 | 0.033
σ | [5 empty] | 0.064 | 0.146
LOB | [5 empty] | 0.2 | 0.5
ddPCR, digital droplet PCR
ID, identifier
NTC, non template control
WT, wildtype
mL, milliliters
µl, microliters
MUT, mutated
LOB, limit of blank

## Supplementary Table 4
Supplementary Table 4. Determining analytical specificity by analyzing 13 plasma samples from elderly healthy individuals.
ID | Age | Sample type | Sample volume (mL) | MUT copies / mL | WT copies / mL | ddPCR result (pos/neg) | Allele frequency (%)
HEAL_022 | 52 | Plasma | 4 | 0 | 843.9286 | neg | 0
HEAL_023 | 51 | Plasma | 4 | 0 | 1221.7875 | neg | 0
HEAL_024 | 61 | Plasma | 4 | 0 | 1351.5 | neg | 0
HEAL_025 | 54 | Plasma | 4 | 0 | 872.5 | neg | 0
HEAL_026 | 56 | Plasma | 4 | 0 | 1323.1339 | neg | 0
HEAL_027 | 22 | Plasma | 4 | 0 | 1660.8679 | neg | 0
HEAL_028 | 25 | Plasma | 4 | 0 | 1721.4214 | neg | 0
HEAL_029 | 62 | Plasma | 3.8 | 0 | 4477.188 | neg | 0
HEAL_030 | 58 | Plasma | 2.2 | 0 | 2726.3929 | neg | 0
HEAL_031 | 55 | Plasma | 3.1 | 0 | 2091.4908 | neg | 0
HEAL_032 | 83 | Plasma | 4 | 0 | 1707.225 | neg | 0
HEAL_033 | 63 | Plasma | 3.6 | 0 | 1798.9484 | neg | 0
HEAL_034 | 61 | Plasma | 3.8 | 0 | 1758.6842 | neg | 0
Median age: | 56
ID, identifier
mL, milliliters
WT, wildtype
MUT, mutated
ddPCR, digital droplet PCR

## Supplementary Table 5
Supplementary Table 5. Plasma samples from CNSL patients with known MYD88 L265P status used for assessment of inter-assay reliability.
ID | Diagnosis | ddPCR result Operator 1 (pos/neg) | Allele frequency (%), Operator 1 | ddPCR result Operator 2 (pos/neg) | Allele frequency (%), Operator 2
P01 | CNSL | pos | 0.67 | pos | 0.74
P02 | CNSL | neg | 0 | neg | 0
P03 | CNSL | pos | 0.22 | pos | 0.13
P04 | CNSL | pos | 2.27 | pos | 2.31
P05 | CNSL | pos | 0.21 | pos | 0.1
P06 | CNSL | pos | 1.75 | pos | 2.17
P07 | CNSL | pos | 1.74 | pos | 2.11
P08 | CNSL | pos | 1.78 | pos | 2.1
P09 | CNSL | neg | 0 | neg | 0
P10 | CNSL | neg | 0 | neg | 0
P11 | CNSL | neg | 0 | neg | 0
P12 | CNSL | neg | 0 | neg | 0
P13 | CNSL | neg | 0 | neg | 0
CNSL, central nervous system lymphoms
ID, identifier
ddPCR, digital droplet PCR
pos, positive
neg, negative

## Supplementary Table 6
Supplementary Table 6. Plasma samples from CNSL patients used for the assessment of intra-assay reliability.
ID | Diagnosis | Total copies well 1 | Total copies well 2 | Total copies well 3 | Total copies well 4 | Variation coefficient (%)
P14 | CNSL | 798.4 | 808.6 | 842.6 | 802.4 | 2.48
P15 | CNSL | 1282 | 1393.2 | 1350 | 1392 | 3.86
P16 | CNSL | 708.4 | 675.2 | 648 | 656 | 4
P17 | CNSL | 834.6 | 794 | 819 | 793.4 | 2.49
P18 | CNSL | 860 | 891.2 | 877.2 | 951.2 | 4.43
P19 | CNSL | 2046 | 2158 | 2028 | 2060 | 2.81
P20 | CNSL | 2548 | 2552 | 2680 | 2658 | 2.66
P21 | CNSL | 3004 | 2914 | 3010 | 2898 | 1.99
CNSL, central nervous system lymphoms
ID, identifier

## Supplementary Table 7
Supplementary Table 7. Overview and characterisitcs of CSF / plasma samples from patients included in the independent validation cohort.
Patient ID | Diagnosis | Sex (M/F) | Sample type | Sample time point | Sample volume (mL) | ddPCR result (pos/neg) | ddPCR allele frequency (%) | NGS results available (yes/no) | NGS result (pos/neg) | NGS allele frequency (%) | Matched tumor MYD88 L265P status (pos/neg/NA) | Conventional CSF CP/FC analysis (pos/neg/NA)
CNSL_001 | PCNSL | F | Plasma | Diagnosis | 9 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_002 | PCNSL | M | Plasma | 1st relapse | 9 | neg | 0 | yes | pos | 0.155 | NA | NA
CNSL_003 | iSCNSL | M | CSF | Diagnosis | 1.1 | pos | 2.492 | yes | pos | 1.931 | pos | neg
CNSL_003 | iSCNSL | M | Plasma | Diagnosis | 9 | neg | 0 | yes | pos | 0.103 | pos | NA
CNSL_003 | iSCNSL | M | Plasma | 1st relapse | 9.5 | pos | 0.074 | yes | neg | 0 | NA | NA
CNSL_004 | iSCNSL | M | Plasma | 1st relapse | 6.54 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_005 | PCNSL | M | CSF | Diagnosis | 2 | pos | 4.156 | no | NA | NA | pos | pos
CNSL_005 | PCNSL | M | Plasma | Diagnosis | 2.7 | neg | 0 | yes | pos | 0.113 | pos | NA
CNSL_005 | PCNSL | M | Plasma | 1st relapse | 7.2 | pos | 0.22977 | yes | pos | 0.096 | NA | NA
CNSL_005 | PCNSL | M | Plasma | 2nd relapse | 3.6 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_005 | PCNSL | M | Plasma | 3rd relapse | 8.8 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_006 | iSCNSL | F | Plasma | 3rd relapse | 8.2 | neg | 0 | no | NA | NA | NA | NA
CNSL_007 | PCNSL | M | Plasma | 2nd relapse | 8.5 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_007 | PCNSL | M | Plasma | 3rd relapse | 8.4 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_008 | PCNSL | F | Plasma | Diagnosis | 4 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_008 | PCNSL | F | Plasma | 1st relapse | 8.7 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_009 | PCNSL | F | Plasma | 3rd relapse | 9.7 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_010 | PCNSL | M | Plasma | Diagnosis | 9.1 | neg | 0 | no | NA | NA | neg | NA
CNSL_011 | PCNSL | F | Plasma | Diagnosis | 6.1 | pos | 0.077459 | yes | pos | 0.154 | pos | NA
CNSL_012 | PCNSL | M | Plasma | Diagnosis | 6.7 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_013 | PCNSL | F | Plasma | 1st relapse | 9.8 | pos | 0.23713 | yes | pos | 0.241 | pos | NA
CNSL_014 | PCNSL | F | Plasma | Diagnosis | 3.1 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_015 | PCNSL | M | CSF | Diagnosis | 1.5 | pos | 1.2724 | yes | pos | 3.198 | pos | pos
CNSL_015 | PCNSL | M | Plasma | Diagnosis | 3.1 | pos | 0.15985 | yes | pos | 0.293 | pos | NA
CNSL_016 | iSCNSL | M | Plasma | 1st relapse | 8.5 | neg | 0 | no | NA | NA | pos | NA
CNSL_017 | PCNSL | F | Plasma | Diagnosis | 8.3 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_018 | PCNSL | F | Plasma | Diagnosis | 9.2 | neg | 0 | no | NA | NA | pos | NA
CNSL_018 | PCNSL | F | Plasma | 1st relapse | 8.5 | neg | 0 | no | NA | NA | NA | NA
CNSL_019 | PCNSL | M | Plasma | Diagnosis | 8.8 | neg | 0 | no | NA | NA | neg | NA
CNSL_019 | PCNSL | M | Plasma | 1st relapse | 6.3 | neg | 0 | no | NA | NA | NA | NA
CNSL_020 | PCNSL | F | CSF | Diagnosis | 3 | neg | 0 | no | NA | NA | neg | neg
CNSL_020 | PCNSL | F | Plasma | Diagnosis | 8.4 | neg | 0 | no | NA | NA | neg | NA
CNSL_021 | PCNSL | M | Plasma | 1st relapse | 7.6 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_021 | PCNSL | M | Plasma | 2nd relapse | 10.5 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_021 | PCNSL | M | Plasma | 3rd relapse | 8.4 | pos | 0.06411 | yes | pos | 0.371 | NA | NA
CNSL_022 | PCNSL | F | CSF | 2nd relapse | 3 | neg | 0 | yes | neg | 0 | pos | neg
CNSL_022 | PCNSL | F | Plasma | Diagnosis | 8.8 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_022 | PCNSL | F | Plasma | 1st relapse | 8.4 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_022 | PCNSL | F | Plasma | 2nd relapse | 7.8 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_022 | PCNSL | F | Plasma | 3rd relapse | 9.9 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_023 | PCNSL | F | Plasma | Diagnosis | 4 | neg | 0 | no | NA | NA | NA | NA
CNSL_024 | PCNSL | F | Plasma | Diagnosis | 9.2 | pos | 41.2907 | yes | pos | 42.735 | pos | NA
CNSL_025 | PCNSL | M | Plasma | Diagnosis | 10 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_026 | PCNSL | M | Plasma | Diagnosis | 9.4 | neg | 0 | no | NA | NA | neg | NA
CNSL_026 | PCNSL | M | Plasma | 1st relapse | 10.4 | neg | 0 | no | NA | NA | NA | NA
CNSL_027 | PCNSL | F | Plasma | Diagnosis | 5.6 | pos | 0.052046 | yes | pos | 0.057 | pos | NA
CNSL_028 | PCNSL | M | Plasma | Diagnosis | 7 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_029 | PCNSL | M | Plasma | Diagnosis | 6.4 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_029 | PCNSL | M | Plasma | 1st relapse | 7.6 | neg | 0 | yes | pos | 0.023 | NA | NA
CNSL_029 | PCNSL | M | Plasma | 2nd relapse | 9.9 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_029 | PCNSL | M | Plasma | 3rd relapse | 11.8 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_030 | PCNSL | F | CSF | 2nd relapse | 1.9 | pos | 29.9244 | yes | pos | 34.259 | pos | neg
CNSL_030 | PCNSL | F | Plasma | Diagnosis | 10.6 | pos | 16.8093 | yes | pos | 19.772 | NA | NA
CNSL_030 | PCNSL | F | Plasma | 2nd relapse | 9.8 | neg | 0 | yes | pos | 0.036 | NA | NA
CNSL_031 | PCNSL | M | Plasma | Diagnosis | 6.1 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_032 | PCNSL | M | Plasma | Diagnosis | 11.3 | pos | 0.056505 | yes | pos | 0.166 | pos | NA
CNSL_033 | PCNSL | F | Plasma | Diagnosis | 11.6 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_034 | PCNSL | M | Plasma | Diagnosis | 11.8 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_035 | PCNSL | M | Plasma | Diagnosis | 5.9 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_036 | iSCNSL | M | CSF | 1st relapse | 1.8 | neg | 0 | yes | neg | 0 | NA | neg
CNSL_036 | iSCNSL | M | Plasma | Diagnosis | 8.2 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_036 | iSCNSL | M | Plasma | 1st relapse | 9.1 | neg | 0 | yes | pos | 0.189 | NA | NA
CNSL_037 | iSCNSL | M | Plasma | Diagnosis | 8.8 | pos | 0.50862 | yes | pos | 0.843 | pos | NA
CNSL_038 | PCNSL | F | Plasma | 1st relapse | 7 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_039 | PCNSL | F | Plasma | Diagnosis | 9 | pos | 3.0797 | yes | pos | 4.279 | pos | NA
CNSL_039 | PCNSL | F | Plasma | 1st relapse | 9.9 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_040 | iSCNSL | M | Plasma | Diagnosis | 9.1 | pos | 0.33032 | yes | pos | 0.699 | pos | NA
CNSL_040 | iSCNSL | M | Plasma | 1st relapse | 11 | pos | 0.4793 | yes | pos | 0.506 | NA | NA
CNSL_040 | iSCNSL | M | Plasma | 3rd relapse | 8.4 | neg | 0 | yes | neg | 0 | NA | NA
CNSL_041 | PCNSL | F | Plasma | Diagnosis | 3.4 | neg | 0 | no | NA | NA | pos | NA
CNSL_042 | iSCNSL | M | Plasma | Diagnosis | 8.8 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_043 | PCNSL | M | CSF | Diagnosis | 2.7 | pos | 1.0521 | no | NA | NA | NA | neg
CNSL_044 | PCNSL | F | CSF | Diagnosis | 4.1 | neg | 0 | no | NA | NA | NA | neg
CNSL_045 | PCNSL | M | CSF | Diagnosis | 3.2 | pos | 0.60771 | no | NA | NA | NA | neg
CNSL_046 | PCNSL | F | CSF | Diagnosis | 1.8 | neg | 0 | no | NA | NA | NA | neg
CNSL_047 | PCNSL | F | CSF | Diagnosis | 3.2 | neg | 0 | no | NA | NA | NA | neg
CNSL_048 | PCNSL | M | Plasma | 1st relapse | 9.2 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_049 | PCNSL | M | CSF | Diagnosis | 3 | pos | 0.17857 | no | NA | NA | NA | neg
CNSL_050 | PCNSL | M | CSF | Diagnosis | 3.1 | pos | 0.98039 | no | NA | NA | NA | neg
CNSL_051 | PCNSL | M | CSF | Diagnosis | 4.1 | neg | 0 | no | NA | NA | NA | pos
CNSL_052 | PCNSL | M | CSF | Diagnosis | 1.9 | pos | 0.072763 | no | NA | NA | pos | neg
CNSL_053 | PCNSL | F | CSF | Diagnosis | 2.2 | pos | 7.3467 | no | NA | NA | NA | pos
CNSL_054 | PCNSL | M | CSF | Diagnosis | 5.1 | pos | 11.033 | no | NA | NA | NA | neg
CNSL_055 | PCNSL | F | CSF | Diagnosis | 2.7 | pos | 15.346 | no | NA | NA | NA | neg
CNSL_056 | PCNSL | M | CSF | Diagnosis | 2.3 | pos | 60.8579 | no | NA | NA | NA | neg
CNSL_057 | PCNSL | F | CSF | Diagnosis | 3.2 | pos | 0.20204 | no | NA | NA | NA | neg
CNSL_058 | PCNSL | M | CSF | Diagnosis | 2.7 | neg | 0 | no | NA | NA | NA | neg
CNSL_059 | PCNSL | M | CSF | Diagnosis | 3.6 | pos | 0.81186 | no | NA | NA | NA | neg
CNSL_060 | PCNSL | F | CSF | Diagnosis | 4 | pos | 3.7256 | no | NA | NA | NA | neg
CNSL_061 | PCNSL | M | CSF | Diagnosis | 3.6 | pos | 0.095958 | no | NA | NA | NA | neg
CNSL_062 | PCNSL | F | Plasma | Diagnosis | 7.4 | pos | 0.093692 | yes | pos | 0.146 | pos | NA
CNSL_063 | PCNSL | F | CSF | Diagnosis | 0.9 | neg | 0 | no | NA | NA | NA | pos
CNSL_064 | PCNSL | M | CSF | Diagnosis | 2.1 | pos | 25.3632 | no | NA | NA | NA | neg
CNSL_065 | PCNSL | F | CSF | Diagnosis | 1.2 | pos | 5.5334 | yes | pos | 11.494 | NA | pos
CNSL_066 | PCNSL | M | CSF | Diagnosis | 1.7 | neg | 0 | no | NA | NA | NA | neg
CNSL_067 | PCNSL | M | CSF | Diagnosis | 2.1 | neg | 0 | no | NA | NA | neg | neg
CNSL_068 | PCNSL | M | CSF | Diagnosis | 1.1 | neg | 0 | no | NA | NA | neg | neg
CNSL_068 | PCNSL | M | Plasma | Diagnosis | 7.9 | pos | 0.40494 | no | NA | NA | NA | NA
CNSL_069 | PCNSL | F | Plasma | Diagnosis | 6.5 | neg | 0 | no | NA | NA | neg | NA
CNSL_070 | PCNSL | F | CSF | 2nd relapse | 2.5 | pos | 71.8075 | yes | pos | 76.224 | NA | neg
CNSL_070 | PCNSL | F | CSF | 3rd relapse | 0.8 | pos | 82.9787 | yes | pos | 77.338 | NA | neg
CNSL_070 | PCNSL | F | Plasma | Diagnosis | 9.8 | pos | 4.5482 | yes | pos | 8.573 | pos | NA
CNSL_070 | PCNSL | F | Plasma | 1st relapse | 9.9 | pos | 1.0417 | yes | pos | 1.235 | NA | NA
CNSL_071 | PCNSL | F | Plasma | Diagnosis | 8.3 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_072 | iSCNSL | M | Plasma | Diagnosis | 11.3 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_073 | PCNSL | M | Plasma | Diagnosis | 10.1 | neg | 0 | yes | neg | 0 | pos | NA
CNSL_074 | PCNSL | F | CSF | Diagnosis | 2.5 | pos | 3.9201 | yes | pos | 1.195 | pos | neg
CNSL_074 | PCNSL | F | Plasma | Diagnosis | 8.5 | pos | 0.1313 | yes | pos | 0.097 | pos | NA
CNSL_075 | PCNSL | F | Plasma | Diagnosis | 10.3 | pos | 0.062536 | yes | pos | 0.034 | pos | NA
CNSL_076 | PCNSL | M | CSF | Diagnosis | 1.5 | pos | 0.34186 | yes | pos | 0.384 | pos | neg
CNSL_076 | PCNSL | M | Plasma | Diagnosis | 7.7 | pos | 0.092421 | yes | pos | 0.092 | pos | NA
CNSL_077 | PCNSL | M | CSF | Diagnosis | 9.9 | neg | 0 | yes | neg | 0 | pos | neg
CNSL_078 | PCNSL | M | CSF | Diagnosis | 2.7 | neg | 0 | no | NA | NA | neg | neg
CNSL_079 | PCNSL | F | CSF | Diagnosis | 1.5 | pos | 32.6547 | yes | pos | 31.897 | pos | neg
CNSL_080 | PCNSL | F | CSF | Diagnosis | 4.9 | pos | 0.10241 | yes | neg | 0 | pos | neg
CNSL_081 | iSCNSL | F | CSF | Diagnosis | 5 | neg | 0 | no | NA | NA | NA | neg
CNSL_082 | PCNSL | F | CSF | Diagnosis | 6.3 | pos | 6.8939 | no | NA | NA | NA | neg
CNSL_083 | PCNSL | F | CSF | Diagnosis | 1 | pos | 47.5138 | no | NA | NA | NA | pos
CNSL_084 | PCNSL | M | CSF | Diagnosis | 1.4 | pos | 31.299 | no | NA | NA | NA | neg
CNSL_085 | PCNSL | M | CSF | Diagnosis | 1.9 | pos | 31.9877 | no | NA | NA | NA | pos
CNSL_086 | PCNSL | F | CSF | Diagnosis | 4.3 | pos | 22.7865 | no | NA | NA | NA | pos
CNSL_087 | PCNSL | M | CSF | Diagnosis | 3.1 | pos | 29.5102 | no | NA | NA | NA | pos
Non-CNSL_001 | GBM | M | Plasma | Diagnosis | 9.4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_002 | GBM | M | Plasma | Diagnosis | 3.6 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_003 | GBM | M | Plasma | Diagnosis | 6.1 | neg | 0 | no | NA | NA | neg | NA
Non-CNSL_004 | GBM | M | Plasma | Diagnosis | 6.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_005 | Astrocytoma | M | Plasma | Diagnosis | 7.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_006 | Atypical plexus papilloma | M | CSF | Diagnosis | 6.4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_006 | Atypical plexus papilloma | M | Plasma | Diagnosis | 7 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_007 | GBM | F | Plasma | Diagnosis | 8.9 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_008 | GBM | M | Plasma | Diagnosis | 9.4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_009 | GBM | M | Plasma | Diagnosis | 9 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_010 | GBM | F | Plasma | Diagnosis | 8.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_011 | GBM | F | Plasma | Diagnosis | 8.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_012 | NSCLC metastasis | F | CSF | Diagnosis | 1.4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_013 | NSCLC metastasis | M | CSF | Diagnosis | 0.9 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_014 | Astrocytoma | M | CSF | Diagnosis | 4.3 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_014 | Astrocytoma | M | Plasma | Diagnosis | 4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_015 | Astrocytoma | F | CSF | Diagnosis | 0.7 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_015 | Astrocytoma | F | Plasma | Diagnosis | 8.1 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_016 | Pleomorphic Xanthoastrocytoma | F | CSF | Diagnosis | 4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_016 | Pleomorphic Xanthoastrocytoma | F | Plasma | Diagnosis | 9.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_017 | GBM | F | CSF | Diagnosis | 0.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_017 | GBM | F | Plasma | Diagnosis | 9.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_018 | Melanoma metastasis | M | CSF | Diagnosis | 1.5 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_019 | Melanoma metastasis | M | CSF | Diagnosis | 1.4 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_020 | Neurotuberculosis | F | CSF | Diagnosis | 0.9 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_021 | Optic neuritis | M | CSF | Diagnosis | 1.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_022 | Multiple sclerosis | F | CSF | Diagnosis | 0.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_023 | Autoimmune Pachy- / Leptomeningitis | F | CSF | Diagnosis | 1.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_024 | Multiple sclerosis | M | CSF | Diagnosis | 1.1 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_025 | Neurosarcoidosis | F | CSF | Diagnosis | 0.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_026 | Neuromyelitis | F | CSF | Diagnosis | 1.5 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_027 | Multiple sclerosis | M | CSF | Diagnosis | 1.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_028 | Acute disseminated encephalomyelitis (ADEM) | F | CSF | Diagnosis | 0.3 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_029 | Multiple sclerosis | F | CSF | Diagnosis | 0.6 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_030 | Bacterial meningitis | F | CSF | Diagnosis | 1.3 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_031 | Breast cancer metastasis (meningeal carcinomatosis) | F | CSF | Diagnosis | 2.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_032 | Breast cancer metastasis (meningeal carcinomatosis) | F | CSF | Diagnosis | 1.5 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_033 | Cancer of unknown primary metastasis (meningeal carcinomatosis) | F | CSF | Diagnosis | 3 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_034 | SCLC metastasis (meningeal carcinomatosis) | M | CSF | Diagnosis | 1.5 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_035 | Breast cancer metastasis (meningeal carcinomatosis) | F | CSF | Diagnosis | 0.6 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_036 | NSCLC metastasis (meningeal carcinomatosis) | M | CSF | Diagnosis | 2.8 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_037 | GBM | M | CSF | Diagnosis | 0.7 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_038 | GBM | M | CSF | Diagnosis | 0.6 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_039 | Desmoplastic Medulloblastoma | F | CSF | Diagnosis | 1.2 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_040 | GBM | F | CSF | Diagnosis | 0.5 | neg | 0 | no | NA | NA | NA | NA
Non-CNSL_041 | Encephalitis | F | CSF | Diagnosis | 5.8 | neg | 0 | no | NA | NA | NA | NA
CNSL, central nervous system lymphoms
PCNSL, primary CNSL
iSCNSL, isolated secondary CNSL
ID, identifier
CSF, cerebrospinal fluid
M, male
F, female
mL, milliliters
pos, positive
neg, negative
ddPCR, digital droplet PCR
NGS, next-generation sequencing
NA, not applicable or not assessed
CP, cytology
FC, flow cytometry
GBM, glioblastoma
NSCLC, non-small cell lung cancer
SCLC, small cell lung cancer

## Supplementary Table 8
Supplementary Table 8. Overview and characteristics of CSF samples submitted for routine ctDNA testing by the MYD88 L265P ddPCR assay.
Patient ID | Sample ID | Sex (M/F) | Age at lumbar puncture | Submitter (internal/external) | Submitted CSF volume (mL) | Turnover (days) | ddPCR result (pos/neg) | ddPCR allele frequency (%) | Participation in observational study (yes/no)
P001 | P001_p1 | M | 67 | external | 2.1 | 5 | pos | 0.1064 | no
P002 | P002_p1 | F | 68 | external | 6.3 | 2 | pos | 0.22631 | yes
P003 | P003_p1 | F | 78 | internal | 0.95 | 4 | pos | 0.48731 | yes
P003 | P003_p2 | F | 78 | internal | 2.5 | 2 | pos | 0.29349 | yes
P004 | P004_p1 | F | 76 | external | 5.5 | 5 | neg | 0 | no
P005 | P005_p1 | M | 79 | internal | 0.8 | 1 | pos | 4.6243 | yes
P005 | P005_p2 | M | 80 | internal | 1.9 | 7 | pos | 0.2259 | yes
P006 | P006_p1 | M | 57 | external | 2.3 | 11 | neg | 0 | no
P006 | P006_p2 | M | 51 | external | 5.65 | 11 | neg | 0 | yes
P007 | P007_p1 | F | 72 | internal | 2.7 | 6 | neg | 0 | yes
P008 | P008_p1 | M | 78 | internal | 5.6 | 13 | neg | 0 | yes
P009 | P009_p1 | F | 58 | internal | 2.1 | 9 | neg | 0 | yes
P010 | P010_p1 | M | 72 | external | 4 | 2 | neg | 0 | yes
P011 | P011_p1 | M | 61 | external | 2.5 | 1 | neg | 0 | no
P012 | P012_p1 | M | 81 | external | 1.7 | 9 | neg | 0 | no
P013 | P013_p1 | M | 47 | internal | 1.2 | 9 | neg | 0 | yes
P014 | P014_p1 | M | 75 | internal | 0.6 | 1 | neg | 0 | yes
P015 | P015_p1 | F | 65 | internal | 2.9 | 4 | neg | 0 | yes
P016 | P016_p1 | F | 57 | internal | 1.55 | 2 | pos | 0.30811 | yes
P017 | P017_p1 | F | 75 | external | 4.8 | 1 | pos | 0.11521 | yes
P018 | P018_p1 | F | 73 | external | 5.1 | 3 | neg | 0 | yes
P019 | P019_p1 | M | 83 | internal | 3.8 | 6 | neg | 0 | yes
P020 | P020_p1 | F | 46 | external | 2.3 | 2 | neg | 0 | no
P021 | P021_p1 | M | 66 | internal | 1.6 | 6 | neg | 0 | yes
P022 | P022_p1 | M | 57 | internal | 1.8 | 0 | neg | 0 | yes
P023 | P023_p1 | M | 25 | external | 2.5 | 2 | neg | 0.0034687 | no
P024 | P024_p1 | M | 62 | internal | 2.5 | 3 | neg | 0 | yes
P025 | P025_p1 | M | 75 | external | 3 | 4 | pos | 0.58102 | no
P026 | P026_p1 | M | 48 | internal | 2.9 | 21 | neg | 0 | yes
P027 | P027_p1 | M | 66 | internal | 1.65 | 4 | pos | 10.7981 | yes
P028 | P028_p1 | F | 81 | external | 4 | 11 | pos | 0.50044 | no
P029 | P029_p1 | M | 60 | external | 3 | 1 | neg | 0 | no
P030 | P030_p1 | F | 75 | external | 5 | 2 | neg | 0 | yes
P031 | P031_p1 | M | 72 | external | 4.6 | 3 | neg | 0 | yes
P032 | P032_p1 | M | 57 | internal | 3.5 | 4 | neg | 0 | yes
P033 | P033_p1 | M | 68 | external | 5.5 | 5 | neg | 0 | no
P034 | P034_p1 | F | 81 | external | 1.8 | 9 | pos | 0.081833 | yes
P035 | P035_p1 | F | 69 | internal | 0.5 | 3 | pos | 2.2556 | yes
P036 | P036_p1 | F | 44 | external | 5.9 | 7 | neg | 0.0036514 | no
P037 | P037_p1 | F | 36 | external | 2.1 | 5 | neg | 0 | no
P038 | P038_p1 | M | 44 | internal | 2.4 | 3 | neg | 0 | yes
P039 | P039_p1 | M | 72 | internal | 2.5 | 7 | pos | 4.6296 | yes
P040 | P040_p1 | F | 76 | external | 5.05 | 2 | neg | 0 | yes
P041 | P041_p1 | F | 79 | external | 4.4 | 0 | neg | 0 | yes
P042 | P042_p1 | F | 77 | internal | 2 | 7 | neg | 0 | yes
P043 | P043_p1 | M | 66 | external | 3.5 | 2 | neg | 0 | yes
P044 | P044_p1 | M | 67 | internal | 7.6 | 4 | neg | 0 | yes
P045 | P045_p1 | M | 52 | internal | 1.8 | 6 | neg | 0 | yes
P046 | P046_p1 | F | 78 | internal | 0.9 | 8 | pos | 13.0958 | yes
P047 | P047_p1 | M | 78 | external | 4.5 | 2 | pos | 0.19845 | yes
P048 | P048_p1 | M | 69 | internal | 3 | 7 | neg | 0 | yes
P049 | P049_p1 | F | 60 | external | 5.55 | 7 | neg | 0 | yes
P050 | P050_p1 | M | 64 | internal | 1.35 | 3 | neg | 0 | yes
P051 | P051_p1 | F | 61 | external | 4.3 | 1 | neg | 0 | yes
P051 | P051_p2 | F | 61 | external | 1.15 | 2 | neg | 0 | yes
P052 | P052_p1 | F | 76 | external | 1.5 | 1 | neg | 0 | yes
P053 | P053_p1 | M | 55 | external | 5 | 9 | neg | 0 | yes
P054 | P054_p1 | F | 67 | internal | 2.75 | 3 | neg | 0 | yes
P055 | P055_p1 | M | 79 | external | 6.1 | 6 | neg | 0 | yes
P056 | P056_p1 | F | 78 | internal | 1.1 | 6 | neg | 0 | yes
P057 | P057_p1 | M | 66 | external | 3 | 4 | neg | 0 | no
P058 | P058_p1 | M | 76 | internal | 2.2 | 6 | pos | 0.1673 | yes
P059 | P059_p1 | F | 43 | external | 2.5 | 1 | neg | 0 | no
P060 | P060_p1 | M | 56 | external | 2 | 9 | neg | 0 | yes
P061 | P061_p1 | F | 48 | external | 3 | 7 | neg | 0 | yes
P062 | P062_p1 | M | 63 | internal | 3.1 | 5 | neg | 0 | yes
P063 | P063_p1 | F | 56 | internal | 1 | 5 | neg | 0 | yes
P064 | P064_p1 | M | 79 | external | 4.1 | 1 | pos | 0.079212 | no
P065 | P065_p1 | F | 85 | external | 3.8 | 1 | pos | 2.7286 | yes
P066 | P066_p1 | M | 80 | internal | 3.75 | 1 | pos | 4.1233 | yes
P067 | P067_p1 | F | 74 | external | 5.7 | 7 | neg | 0 | yes
P068 | P068_p1 | F | 83 | external | 5 | 5 | pos | 1.4592 | yes
P069 | P069_p1 | F | 71 | internal | 1.8 | 18 | pos | 0.55249 | yes
P070 | P070_p1 | F | 69 | internal | 1.7 | 8 | pos | 3.0188 | yes
P071 | P071_p1 | F | 52 | external | 5 | 2 | neg | 0 | no
P072 | P072_p1 | M | 80 | external | 1.8 | 7 | neg | 0 | yes
P073 | P073_p1 | M | 73 | external | 1.4 | 7 | neg | 0 | no
P074 | P074_p1 | F | 74 | external | 4.3 | 2 | neg | 0 | no
P075 | P075_p1 | F | 53 | internal | 0.9 | 5 | neg | 0 | yes
P076 | P076_p1 | M | 52 | internal | 2.3 | 6 | neg | 0 | yes
P076 | P076_p1 | M | 52 | internal | 2.4 | 5 | neg | 0 | yes
P077 | P077_p1 | M | 55 | internal | 1.4 | 6 | neg | 0 | no
P078 | P078_p1 | F | 67 | external | 3 | 2 | neg | 0 | no
P079 | P079_p1 | M | 55 | internal | 2.3 | 1 | pos | 1.3015 | yes
P079 | P079_p2 | M | 55 | internal | 1.6 | 10 | neg | 0 | yes
P080 | P080_p1 | M | 74 | external | 5.7 | 4 | neg | 0 | yes
P081 | P081_p1 | M | 64 | external | 3.15 | 12 | pos | 0.1227 | yes
P082 | P082_p1 | F | 62 | external | 2.2 | 1 | neg | 0 | yes
P083 | P083_p1 | M | 55 | external | 3 | 5 | neg | 0 | no
P084 | P084_p1 | M | 72 | internal | 3.6 | 4 | pos | 0.21672 | yes
P084 | P084_p2 | M | 72 | internal | 0.8 | 2 | pos | 0.39216 | yes
P085 | P085_p1 | M | 70 | internal | 1 | 2 | neg | 0 | yes
P086 | P086_p1 | F | 83 | internal | 1 | 4 | neg | 0 | yes
P087 | P087_p1 | F | 60 | internal | 1.5 | 6 | neg | 0 | yes
P088 | P088_p1 | M | 80 | external | 5.3 | 7 | pos | 0.22749 | yes
P089 | P089_p1 | F | 38 | external | 3.8 | 8 | neg | 0 | no
P090 | P090_p1 | M | 63 | internal | 1.5 | 5 | neg | 0 | yes
P091 | P091_p1 | M | 77 | internal | 6.5 | 6 | neg | 0 | yes
P092 | P092_p1 | M | 56 | external | 2.15 | 3 | neg | 0 | no
P093 | P093_p1 | M | 66 | external | 6.4 | 4 | pos | 7.4995 | no
P093 | P093_p2 | M | 66 | external | 3.5 | 6 | neg | 0 | no
P094 | P094_p1 | M | 27 | external | 2.9 | 5 | neg | 0 | no
P095 | P095_p1 | F | 49 | external | 2.9 | 7 | neg | 0 | no
P095 | P095_p2 | F | 49 | external | 4 | 3 | pos | 0.074248 | no
P095 | P095_p3 | F | 49 | external | 2.8 | 20 | neg | 0 | no
P096 | P096_p1 | M | 73 | internal | 1.5 | 7 | neg | 0 | yes
P097 | P097_p1 | F | 74 | external | 5.4 | 7 | neg | 0 | yes
P098 | P098_p1 | M | 69 | external | 4.1 | 8 | neg | 0 | yes
P099 | P099_p1 | F | 70 | external | 3.8 | 3 | neg | 0 | no
P100 | P100_p1 | M | 53 | internal | 1.4 | 6 | neg | 0 | yes
P101 | P101_p1 | F | 72 | external | 4 | 5 | pos | 0.54212 | yes
P101 | P101_p2 | F | 72 | external | 5.82 | 3 | pos | 1.9314 | yes
P102 | P102_p1 | M | 42 | internal | 8.9 | 5 | pos | 0.29104 | yes
P103 | P103_p1 | F | 81 | external | 1.3 | 7 | pos | 0.055473 | yes
P104 | P104_p1 | M | 34 | external | 3 | 7 | neg | 0 | no
P104 | P104_p1 | M | 34 | external | 4.5 | 2 | neg | 0 | no
P105 | P105_p1 | F | 64 | external | 3.8 | 9 | pos | 0.074111 | no
P106 | P106_p1 | F | 39 | external | 4.9 | 2 | neg | 0 | no
P107 | P107_p1 | M | 55 | internal | 4.4 | 10 | neg | 0 | yes
P108 | P108_p1 | M | 57 | external | 10 | 5 | neg | 0 | no
P108 | P108_p2 | M | 57 | external | 2.8 | 24 | neg | 0 | no
P109 | P109_p1 | M | 67 | internal | 1.5 | 6 | pos | 0.48369 | yes
P110 | P110_p1 | F | 68 | internal | 1.1 | 10 | neg | 0 | yes
P111 | P111_p1 | M | 66 | external | 2.2 | 4 | neg | 0 | yes
P112 | P112_p1 | F | 76 | external | 4.6 | 2 | neg | 0 | yes
P113 | P113_p1 | M | 75 | external | 2.7 | 5 | pos | 0.76592 | no
P114 | P114_p1 | M | 57 | external | 3.5 | 1 | pos | 1.4358 | no
P115 | P115_p1 | F | 61 | internal | 2.4 | 7 | neg | 0 | yes
P116 | P116_p1 | F | 62 | external | 3.8 | 10 | neg | 0 | no
P117 | P117_p1 | M | 28 | external | 2.8 | 3 | neg | 0 | no
P118 | P118_p1 | M | 20 | external | 1.8 | 2 | neg | 0 | no
P119 | P119_p1 | F | 44 | external | 4 | 1 | neg | 0 | yes
P120 | P120_p1 | M | 48 | external | 2.2 | 5 | pos | 3.3019 | no
P121 | P121_p1 | M | 69 | external | 5 | 2 | pos | 3.6334 | yes
P121 | P121_p2 | M | 69 | external | 5 | 5 | pos | 4.8128 | yes
P121 | P121_p3 | M | 69 | external | 4.6 | 4 | pos | 9.8916 | yes
P122 | P122_p1 | F | 49 | internal | 1.4 | 5 | neg | 0 | no
P123 | P123_p1 | F | 73 | external | 4 | 16 | pos | 0.1405 | no
P124 | P124_p1 | M | 63 | external | 3.6 | 3 | neg | 0 | yes
P125 | P125_p1 | F | 78 | external | 1.5 | 9 | pos | 0.054833 | yes
P126 | P126_p1 | F | 61 | external | 3.5 | 5 | neg | 0 | no
P127 | P127_p1 | M | 66 | internal | 0.7 | 2 | neg | 0 | yes
P128 | P128_p1 | F | 88 | external | 3.3 | 5 | neg | 0 | yes
P129 | P129_p1 | M | 31 | internal | 1.5 | 6 | neg | 0 | yes
P130 | P130_p1 | M | 47 | internal | 1.6 | 4 | neg | 0 | yes
P131 | P131_p1 | F | 81 | external | 3.7 | 5 | neg | 0 | yes
P132 | P132_p1 | F | 60 | internal | 0.85 | 13 | neg | 0 | yes
P133 | P133_p1 | F | 42 | internal | 5.6 | 2 | neg | 0 | yes
P134 | P134_p1 | F | 69 | external | 1.9 | 6 | neg | 0 | no
P135 | P135_p1 | M | 70 | external | 3.85 | 1 | pos | 6.7879 | yes
P136 | P136_p1 | M | 66 | internal | 1.8 | 6 | neg | 0 | no
P137 | P137_p1 | F | 72 | external | 4 | 15 | neg | 0 | no
P138 | P138_p1 | F | 76 | external | 1.4 | 2 | pos | 2.4347 | yes
P139 | P139_p1 | F | 73 | external | 3 | 1 | pos | 0.67791 | no
P140 | P140_p1 | M | 65 | internal | 2 | 9 | pos | 0.056557 | yes
P141 | P141_p1 | F | 74 | internal | 3 | 7 | pos | 0.34483 | yes
P142 | P142_p1 | F | 82 | internal | 5.75 | 20 | pos | 12.3234 | yes
P143 | P143_p1 | M | 41 | external | 1.83 | 4 | neg | 0 | yes
P144 | P144_p1 | M | 63 | internal | 3.8 | 7 | neg | 0 | yes
P144 | P144_p2 | M | 63 | internal | 1.6 | 3 | neg | 0 | yes
P144 | P144_p3 | M | 63 | internal | 2.05 | 11 | pos | 3.1285 | yes
P144 | P144_p4 | M | 63 | internal | 0.9 | 4 | neg | 0.0032424 | yes
P145 | P145_p1 | M | 84 | external | 2.85 | 2 | neg | 0 | no
P146 | P146_p1 | M | 83 | internal | 2.9 | 7 | neg | 0 | yes
P147 | P147_p1 | F | 23 | external | 5.15 | 7 | pos | 0.069589 | yes
P148 | P148_p1 | F | 57 | external | 4.6 | 8 | neg | 0 | yes
P149 | P149_p1 | M | 75 | external | 3.72 | 3 | pos | 1.0522 | no
P150 | P150_p1 | F | 71 | external | 2 | 4 | neg | 0 | no
P151 | P151_p1 | F | 67 | external | 1.6 | 5 | pos | 0.071597 | no
P152 | P152_p1 | F | 73 | external | 2.5 | 5 | pos | 15.6205 | yes
P153 | P153_p1 | F | 63 | external | 4.3 | 3 | pos | 0.38171 | no
P154 | P154_p1 | M | 50 | external | 3.1 | 2 | neg | 0 | no
P155 | P155_p1 | M | 59 | internal | 3.3 | 6 | neg | 0 | yes
P156 | P156_p1 | F | 83 | internal | 3.9 | 0 | pos | 25.2255 | yes
P157 | P157_p1 | F | 62 | internal | 3.6 | 16 | neg | 0 | yes
P158 | P158_p1 | M | 63 | external | 2.4 | 3 | neg | 0 | yes
P159 | P159_p1 | F | 64 | external | 3 | 5 | neg | 0 | yes
P160 | P160_p1 | F | 67 | external | 5.7 | 4 | neg | 0 | yes
P161 | P161_p1 | M | 75 | internal | 3.9 | 1 | neg | 0 | yes
P162 | P162_p1 | M | 84 | internal | 3 | 4 | neg | 0 | yes
P163 | P163_p1 | M | 60 | external | 2.9 | 7 | neg | 0 | no
P164 | P164_p1 | F | 80 | internal | 2.7 | 6 | neg | 0 | yes
P165 | P165_p1 | F | 80 | external | 1.5 | 2 | pos | 1.8817 | yes
P166 | P166_p1 | M | 62 | external | 1.8 | 2 | pos | 1.133 | yes
P167 | P167_p1 | M | 32 | external | 1.3 | 19 | neg | 0.00054146 | yes
P168 | P168_p1 | M | 84 | external | 5.1 | 10 | pos | 0.56462 | yes
P168 | P168_p2 | M | 85 | external | 5.1 | 8 | pos | 0.55504 | yes
P169 | P169_p1 | F | 80 | external | 3.1 | 3 | pos | 1.0225 | no
P170 | P170_p1 | M | 84 | internal | 6 | 2 | neg | 0 | yes
P171 | P171_p1 | F | 70 | internal | 2.4 | 4 | neg | 0 | yes
P172 | P172_p1 | M | 37 | external | 4.4 | 7 | neg | 0 | no
P172 | P172_p2 | M | 37 | external | 4 | 7 | neg | 0 | no
P173 | P173_p1 | M | 75 | external | 4.8 | 2 | neg | 0 | no
P174 | P174_p1 | F | 67 | external | 2.8 | 5 | pos | 0.13657 | yes
P175 | P175_p1 | M | 74 | external | 1.8 | 3 | neg | 0 | yes
P176 | P176_p1 | M | 74 | external | 2 | 1 | pos | 3.4635 | yes
P177 | P177_p1 | M | 75 | internal | 3.6 | 15 | pos | 0.16685 | yes
P177 | P177_p2 | M | 76 | internal | 1.5 | 5 | neg | 0 | yes
P178 | P178_p1 | F | 56 | internal | 1.3 | 2 | pos | 0.19637 | yes
P178 | P178_p2 | F | 57 | internal | 1.1 | 7 | neg | 0 | yes
P178 | P178_p3 | F | 57 | internal | 0.95 | 5 | pos | 4.2182 | yes
P179 | P179_p1 | M | 76 | external | 8.4 | 7 | neg | 0 | yes
P180 | P180_p1 | F | 51 | internal | 1.8 | 3 | pos | 0.1224 | yes
P181 | P181_p1 | F | 68 | external | 3.7 | 7 | neg | 0 | no
P182 | P182_p1 | M | 49 | internal | 9.1 | 1 | neg | 0 | yes
ID, identifier
CSF, cerebrospinal fluid
ddPCR, digital droplet PCR
ctDNA, circulating tumor DNA
M, male
F, female
mL, milliliters
pos, positive
neg, negative

## Supplementary Table 9
Supplementary Table 9. Overview and characteristics of CSF samples and patients participating in the observational study.
Patient ID | Sample ID | Reason for submission Category 1/2/3 | Reason for submission Sub-category | ddPCR result (pos/neg) | Surgical biopsy of brain lesion (yes/no/NA) | Diagnosis brain lesion | Known systemic lymphoma (yes/no/NA) | Systemic lymphoma entity | Steroids received before CSF analysis (yes/no/Unknown) | Cumulative steroid dose (Dexamethasone equivalent, mg) | Conventional analysis of CSF (yes/no) | CSF cytopathology (yes/no) | Result CSF cytopathology | CSF flow cytometry (yes/no) | Result CSF flow cytometry | Result CSF CP/FC combined | Effect of CSF-ctDNA analysis on surgical intervention, treatment, or clinical management (yes/no) | Effect on surgical intervention | Effect on treatemt | Other effect on clinical management | Used as MRD marker or potential use as MRD marker | Initiated CNS-specific therapy based on CSF-ctDNA analysis | Brain tumor response to initiated therapy (radiological or clinically) (yes/no) | Treatment response of brain lesion (Therapy may not be completed yet)
P002 | P002_p1 | 2 | 2d | pos | no | NA | no (Previous PCNSL) | NA | no | NA | yes | no | NA | yes | Unclear | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-DeVIC | yes | PR
P003 | P003_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | yes | 56 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P003 | P003_p2 | 3 | 3 | pos | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P005 | P005_p1 | 2 | 2a | pos | NA | NA | yes | DLBCL | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no (not fit for CNS-directed therapy) | no | yes | NA | NA | NA
P005 | P005_p2 | 3 | 3 | pos | no | NA | yes | DLBCL | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P006 | P006_p2 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P007 | P007_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | yes | 80 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P008 | P008_p1 | 1 | 1c | neg | yes | No definitive diagnosis / T-cell rich lymphocytic infiltrates | no | NA | yes | 6 | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P009 | P009_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P010 | P010_p1 | 1 | 1b | neg | yes | Glioma | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P013 | P013_p1 | 2 | 2a | neg | NA | NA | yes | Burkitt lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P014 | P014_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P015 | P015_p1 | 2 | 2c | neg | no | NA | yes | HGBCL | no | NA | yes | yes | Negative | yes | Positive | Positive | no | no | no (received HD-Cytarabine based on CSF FC result) | no | no | NA | NA | NA
P016 | P016_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | yes | 70 | yes | yes | Negative | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P017 | P017_p1 | 1 | 1a | pos | no | NA | no | NA | yes | 800 | yes | yes | Negative | yes | Positive | Positive | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-MTX + Procarbazine + Zanubrutinib | yes | CR
P018 | P018_p1 | 2 | 2b | neg | no | NA | yes | DLBCL | Unknown | NA | yes | no | NA | yes | Positive | Positive | no | no | no (received R-MTX based on CSF FC result) | no | no | NA | NA | NA
P019 | P019_p1 | 1 | 1b | neg | yes | GBM | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P021 | P021_p1 | 2 | 2a | neg | NA | NA | yes | HGBCL | yes | 32 | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P022 | P022_p1 | 2 | 2d | neg | no | NA | no (Previous PCNSL) | NA | yes | 12 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P024 | P024_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | yes | 80 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P026 | P026_p1 | 1 | 1c | neg | yes | No definitive diagnosis / T-cell rich inflammatory infiltrate | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P027 | P027_p1 | 3 | 3 | pos | yes | DLBCL | yes | DLBCL | yes | 114 | yes | yes | Negative | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P030 | P030_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P031 | P031_p1 | 2 | 2d | neg | no | NA | no (Previous PCNSL) | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P032 | P032_p1 | 1 | 1b | neg | yes | Malignant mesenchymal tumor | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P034 | P034_p1 | 1 | 1b | pos | yes | DLBCL | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | no | no (Therapy started after biopsy) | Accelerated diagnosis | yes | NA | NA | NA
P035 | P035_p1 | 2 | 2a | pos | NA | NA | yes | HGBCL | yes | 88 | yes | yes | Negative | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | R-DHAOx + autoSCT | Not evaluable | No evaluable brain lesion (occult involvement)
P038 | P038_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P039 | P039_p1 | 2 | 2a | pos | NA | NA | yes | Testicular DLBCL | yes | 32 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | MARTA + R-CHOP + autoSCT | Not evaluable | No evaluable brain lesion (occult involvement)
P040 | P040_p1 | 1 | 1b | neg | yes | GBM | no | NA | no | NA | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P041 | P041_p1 | 2 | 2c | neg | no | NA | yes | Follicular lymphoma | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P042 | P042_p1 | 2 | 2a | neg | NA | NA | yes | Transformed follicular lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P043 | P043_p1 | 1 | 1a | neg | no | NA | no | NA | Unknown | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P044 | P044_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P045 | P045_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P046 | P046_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | yes | 84 | yes | no | NA | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P047 | P047_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P048 | P048_p1 | 1 | 1c | neg | yes | No definitive diagnosis | no | NA | yes | 84 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P049 | P049_p1 | 2 | 2a | neg | NA | NA | yes | Marginal zone lymphoma | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P050 | P050_p1 | 2 | 2a | neg | NA | NA | yes | HGBCL | yes | 8 | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P051 | P051_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P051 | P051_p2 | 2 | 2c | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P052 | P052_p1 | 2 | 2b | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P053 | P053_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P054 | P054_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | yes | 10 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P055 | P055_p1 | 2 | 2c | neg | no | NA | yes | HGBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P056 | P056_p1 | 3 | 3 | neg | yes | DLBCL | no | NA | yes | 48 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P058 | P058_p1 | 3 | 3 | pos | yes | DLBCL | yes | DLBCL | yes | 28 | yes | no | NA | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P060 | P060_p1 | 3 | 3 | neg | yes | DLBCL | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P061 | P061_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P062 | P062_p1 | 1 | 1c | neg | yes | No definitive diagnosis / Suspicious for encephalitis | no | NA | yes | 36 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P063 | P063_p1 | 2 | 2d | neg | yes | DLBCL | no (Previous PCNSL) | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P065 | P065_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Negative | no | NA | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | Rituximab + Radiotherapy | no | PD
P066 | P066_p1 | 2 | 2c | pos | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-Temozolomide | yes | PR
P067 | P067_p1 | 1 | 1c | neg | yes | No definitive diagnosis | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P068 | P068_p1 | 2 | 2d | pos | no | NA | no (Previous PCNSL) | NA | no | NA | yes | no | NA | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-DeVIC | yes | CR
P069 | P069_p1 | 2 | 2a | pos | NA | NA | yes | DLBCL | yes | 80 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | Pola-R-CHP + MARTA + autoSCT | yes, clinically | Alleviating peripheral neuropathy / No evaluable brain lesion (occult involvement)
P070 | P070_p1 | 2 | 2b | pos | no | NA | yes | Splenic marginal zone lymphoma | yes | 96 | yes | no | NA | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MARTA + autoSCT | yes | PR
P072 | P072_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P075 | P075_p1 | 2 | 2d | neg | no | NA | no (Previous PCNSL) | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P076 | P076_p1 | 2 | 2a | neg | NA | NA | yes | HGBCL | yes | 32 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P076 | P076_p1 | 2 | 2a | neg | NA | NA | yes | HGBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P079 | P079_p1 | 1 | 1c | pos | yes | No definitive diagnosis | no | NA | yes | 42 | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | yes | no | Treatment-guiding | no | yes | MATRix + autoSCT | yes | CRu
P079 | P079_p2 | 3 | 3 | neg | yes | DLBCL | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P080 | P080_p1 | 2 | 2c | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P081 | P081_p1 | 2 | 2a | pos | NA | NA | yes | DLBCL | no | NA | yes | yes | Negative | no | NA | Negative | yes | no | Treatment-guiding | no | yes | R-CHOP + R-MTX | Not evaluable | No evaluable brain lesion (occult involvement)
P082 | P082_p1 | 2 | 2d | neg | no | NA | no (Previous PCNSL) | NA | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P084 | P084_p1 | 2 | 2c | pos | no | NA | yes | Intravascular large B-cell lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MARTA | yes | CR
P084 | P084_p2 | 3 | 3 | pos | no | NA | yes | Intravascular large B-cell lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P085 | P085_p1 | 3 | 3 | neg | yes | DLBCL | no | NA | yes | 52 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P086 | P086_p1 | 2 | 2a | neg | NA | NA | yes | Mantle cell lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P087 | P087_p1 | 3 | 3 | neg | yes | DLBCL | no | NA | yes | 12 | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P088 | P088_p1 | 1 | 1c | pos | yes | No definitive diagnosis | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | R-MTX | yes | CR
P090 | P090_p1 | 1 | 1b | neg | yes | GBM | no | NA | no | NA | yes | yes | Negative | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P091 | P091_p1 | 1 | 1a | neg | no | NA | no | NA | Unknown | NA | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P096 | P096_p1 | 1 | 1b | neg | yes | GBM | no | NA | no | NA | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P097 | P097_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P098 | P098_p1 | 1 | 1c | neg | yes | No definitive diagnosis | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P100 | P100_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P101 | P101_p1 | 1 | 1c | pos | yes | No definitive diagnosis | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | Dexamethasone | yes | PR
P101 | P101_p2 | 3 | 3 | pos | no | NA | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P102 | P102_p1 | 2 | 2a | pos | NA | NA | yes | HGBCL | yes | 72 | yes | yes | Negative | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | R-CHOP + MATRix | Not evaluable | No evaluable brain lesion (occult involvement)
P103 | P103_p1 | 2 | 2a | pos | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no (not fit for CNS-directed therapy) | no | yes | NA | NA | NA
P107 | P107_p1 | 2 | 2a | neg | NA | NA | yes | Transformed follicular lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P109 | P109_p1 | 3 | 3 | pos | yes | Lymphoplasmacytic lymphoma | no | NA | yes | 84 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P110 | P110_p1 | 2 | 2a | neg | NA | NA | yes | HGBCL | yes | 32 | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P111 | P111_p1 | 2 | 2c | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P112 | P112_p1 | 2 | 2b | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P115 | P115_p1 | 1 | 1a | neg | no | NA | no | NA | yes | 13 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P119 | P119_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | Unknown | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P121 | P121_p1 | 3 | 3 | pos | yes | Waldenstrom macroglobulinemia - Bing-Neel-Syndrome | yes | Waldenstrom macroglobulinemia | yes | NA | yes | no | NA | yes | Positive | Positive | no | no | no | no | yes | Ibrutinib | yes | PR
P121 | P121_p2 | 3 | 3 | pos | no | NA | yes | Waldenstrom macroglobulinemia | yes | NA | yes | no | NA | yes | Positive | Positive | no | no | no | no | yes | NA | NA | NA
P121 | P121_p3 | 3 | 3 | pos | no | NA | yes | Waldenstrom macroglobulinemia | Unknown | NA | no | NA | NA | NA | NA | NA | no | no | no | no | yes | NA | NA | NA
P124 | P124_p1 | 2 | 2c | neg | yes | DLBCL | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P125 | P125_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | yes | 448 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P127 | P127_p1 | 2 | 2a | neg | NA | NA | yes | Follicular lymphoma | Unknown | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P128 | P128_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P129 | P129_p1 | 1 | 1c | neg | yes | No definitive diagnosis / Suspicious for vasculitis | no | NA | yes | 180 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P130 | P130_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | yes | 96 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P131 | P131_p1 | 3 | 3 | neg | yes | DLBCL | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P132 | P132_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | yes | 80 | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P133 | P133_p1 | 1 | 1c | neg | yes | No definitive diagnosis / Suspicious for neuroepithelial lesion | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P135 | P135_p1 | 1 | 1c | pos | yes | No definitive diagnosis | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | MARTA + autoSCT | yes | PR
P138 | P138_p1 | 2 | 2c | pos | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | Radiatiotherapy | yes | PR
P140 | P140_p1 | 2 | 2c | pos | yes | No definitive diagnosis | yes | Transformed follicular lymphoma | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | no | no (decision for watch & wait) | Accelerated diagnosis | yes | NA | NA | NA
P141 | P141_p1 | 1 | 1c | pos | yes | No definitive diagnosis / Suspected cerebral amyloidoma on the basis of Waldenstrom macroglobulinemia | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | yes | no | no | Supported suspected diagnosis | yes | Watch & Wait | NA | NA
P142 | P142_p1 | 1 | 1b | pos | yes | DLBCL | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | no | no (Therapy started after biopsy) | Accelerated diagnosis | yes | NA | NA | NA
P143 | P143_p1 | 2 | 2a | neg | NA | NA | yes | Follicular lymphoma | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P144 | P144_p1 | 3 | 3 | neg | no | NA | yes | DLBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P144 | P144_p2 | 3 | 3 | neg | no | NA | yes | DLBCL | yes | 48 | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P144 | P144_p3 | 3 | 3 | pos | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P144 | P144_p4 | 3 | 3 | neg | no | NA | yes | DLBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P146 | P146_p1 | 1 | 1c | neg | yes | No definitive diagnosis / T-cell rich infiltrates | no | NA | yes | 10 | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P147 | P147_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | yes | Surgical biopsy obviated | Unknown due to LFU | no | yes | Unknown due to LFU | Not evaluable | NA
P148 | P148_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P152 | P152_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | Surgical biopsy obviated | Unknown due to LFU | no | yes | Unknown due to LFU | Not evaluable | NA
P155 | P155_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P156 | P156_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-Temozolomide | yes, clinically | Clinical response (radiologically LFU)
P157 | P157_p1 | 1 | 1a | neg | no | NA | no | NA | Unknown | NA | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P158 | P158_p1 | 2 | 2c | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P159 | P159_p1 | 1 | 1c | neg | yes | No definitive diagnosis / Suspicious for inflammatory disease | no | NA | no | NA | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P160 | P160_p1 | 2 | 2c | neg | no | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P161 | P161_p1 | 1 | 1a | neg | no | NA | no | NA | yes | 1000 | yes | yes | Negative | no | NA | Negative | No | no | no | no | no | NA | NA | NA
P162 | P162_p1 | 2 | 2b | neg | no | NA | yes | Marginal zone lymphoma | no | NA | yes | yes | Unclear cell infiltrate | yes | Positive | Positive | no | no | no (received R-DHA + Zanubrutinib based on CSF FC result) | no | no | NA | NA | NA
P164 | P164_p1 | 2 | 2a | neg | NA | NA | yes | DLBCL | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P165 | P165_p1 | 2 | 2d | pos | no | NA | no (Previous PCNSL) | NA | no | NA | yes | no | NA | yes | Unclear | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | R-MTX | yes | PR
P166 | P166_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MATRix + autoSCT | yes | CR
P167 | P167_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P168 | P168_p1 | 1 | 1c | pos | yes | No definitive diagnosis / Suspicious for intraocular lymphoma | no | NA | no | NA | yes | no | NA | yes | Unclear | Negative | yes | no | Treatment-guiding | no | yes | Intraocular MTX | yes | PR
P168 | P168_p2 | 3 | 3 | pos | no | NA | no | NA | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P170 | P170_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P171 | P171_p1 | 2 | 2c | neg | no | NA | yes | DLBCL | yes | 18 | yes | yes | Unclear cell infiltrate | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P174 | P174_p1 | 1 | 1a | pos | no | NA | no | NA | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MATRix + autoSCT | yes | CRu
P175 | P175_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | no | NA | yes | Negative | Negative | no | no | no | no | no | NA | NA | NA
P176 | P176_p1 | 3 | 3 | pos | yes | DLBCL | no | NA | no | NA | yes | no | NA | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P177 | P177_p1 | 2 | 2b | pos | no | NA | yes | DLBCL | yes | 10 | yes | yes | Unclear cell infiltrate | no | NA | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MARTA + R-CHOP + autoSCT | yes | CR
P177 | P177_p2 | 3 | 3 | neg | no | NA | yes | DLBCL | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P178 | P178_p1 | 2 | 2c | pos | no | NA | yes | DLBCL | no | NA | yes | yes | Negative | yes | Negative | Negative | yes | Surgical biopsy obviated | Treatment-guiding | no | yes | MATRix + autoSCT | yes | CR
P178 | P178_p2 | 3 | 3 | neg | no | NA | yes | DLBCL | no | NA | yes | yes | Negative | yes | Negative | Negative | no | no | no | no | yes | NA | NA | NA
P178 | P178_p3 | 3 | 3 | pos | no | NA | yes | DLBCL | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | yes | NA | NA | NA
P179 | P179_p1 | 1 | 1a | neg | no | NA | no | NA | no | NA | yes | yes | Unclear cell infiltrate | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
P180 | P180_p1 | 2 | 2a | pos | NA | NA | yes | HGBCL | yes | 64 | yes | yes | Negative | yes | Negative | Negative | yes | no | Treatment-guiding | no | yes | R-CHOP + MATRix + autoSCT | Not evaluable | No evaluable brain lesion (occult involvement)
P182 | P182_p1 | 1 | 1a | neg | no | NA | no | NA | yes | 50 | yes | no | NA | yes | Unclear | Negative | no | no | no | no | no | NA | NA | NA
CSF, cerebrospinal fluid
ID, identifier
ddPCR, digital droplet PCR
pos, positive
neg, negative
NA, not applicable or not assessed
DLBCL, diffuse large B-cell lymphoma
GBM, glioblastoma
HGBCL, high-grade B-cell lymphoma
mg, milligram
CP, cytopathology
FC, flow cytometry
ctDNA, circulating tumor DNA
LFU, loss of follow-up
R-DeVIC: Rituximab, Dexamethasone, Etoposide, Ifosfamide, Carboplatin
MARTA: High-dose Methotrexate, High-dose Cytarabine, Rituximab
MATRix: High-dose Methotrexate, High-dose Cytarabine, Rituximab, Thiotepa
R-MTX: Rituximab, High-dose methotrexate
R-DHAOx: Rituximab, Dexamethasone, Cytarabine, Oxaliplatin
R-CHOP: Rituximab, Cyclophosphamide, Doxorubicin, Vincristine, Prednisone
autoSCT: autologous stem cell transplantation
R: Rituximab
Pola-R-CHP: Polatuzumab Vedotin, Rituximab, Cyclophosphamide, Doxorubicin, Prednisone
CR, complete response
CRu, complete response unconfirmed
PR, partial response
PD, progressive disease