Critical aspects of droplet digital reverse transcription loop-mediated isothermal amplification (ddRT-LAMP) for viral pathogens detection

Kenia Chávez-Ramos1,2,3, Frida Trejo1,2,3, Prisciluis Caheri Salas-Navarrete1,2,3, Eva Ramón-Gallegos4, José Esteban Muñoz-Medina5, Luis Álvarez-Icaza1, Luis F. Olguin*2, Oscar Pilloni*1 & Laura Oropeza-Ramos*3

1Instituto de Ingeniería, Universidad Nacional Autónoma de México, CDMX 04510, México.

2Laboratorio de Biofisicoquímica, Facultad de Química, Universidad Nacional Autónoma de México, CDMX 04510, México.

3Facultad de Ingeniería, Universidad Nacional Autónoma de México, CDMX 04510, México.

4Laboratorio de Citopatología Ambiental, ENCB, Instituto Politécnico Nacional (IPN), Campus Zacatenco, Calle Wilfrido Massieu Esquina Cda. Manuel Stampa, Col. Zacatenco. Alcaldia Gustavo A. Madero, CDMX 07738, México.

5División de Laboratorios Especializados, Instituto Mexicano del Seguro Social, CDMX, México.

Corresponding authors: Luis F. Olguin (olguin.lf@quimica.unam.mx), Oscar Pilloni (opillonic@iingen.unam.mx) & Laura Oropeza-Ramos (loropeza@unam.mx)

ORCID

Kenia Chávez-Ramos: 0000-0002-9051-4667

Prisciluis Caheri Salas Navarrete: 0000-0003-2330-0746

Eva Ramón-Gallegos: 0000-0003-2262-520

José Esteban Muñoz-Medina: 0000-0002-1289-4457

Luis Álvarez-Icaza: 0000-0001-9516-3950

Luis F. Olguin: 0000-0002-7199-1740

Oscar Pilloni: 0000-0002-0508-6718

Laura Oropeza-Ramos: 0000-0002-7311-3204

Devices microfabrication

Approximately 3 mL of SU-8 3035 negative photoresist was added to a silicon wafer (diameter 75 mm, thickness 381 ± 25 µm; Silicon Valley Microelectronics) that had been previously cleaned with isopropanol and placed on a spin coater (WS-650MZ-23NPP, Laurell Technologies Corp.). A suitable centrifugation and acceleration program was applied to obtain a final SU-8 layer of around 100 µm for both microdevices (droplet generator and observation chamber devices). The wafer was then heated at 105 °C for 20 min on a hot plate (Corning PC-620D) to cure the photoresist, and the desired pattern was transferred using ultraviolet light with an intelligent micropatterning system maskless exposure machine (SF-100 XCEL, Scotech Ltd.). After that, the resulting silicon wafer was heated at 75°C for 1 min, followed by 5 min at 105°C. The silicon wafer was dipped into a propylene glycol monomethyl ether acetate container to remove the non-polymerized photoresist and finally washed with isopropanol. The last step was heating the silicon wafer at 200 °C for 20 min. Microdevice replicas were fabricated using soft lithography; the etched silicon wafer was placed into a petri dish, and a mixture of 10:1 w/w of PDMS: curing agent was poured. A vacuum desiccator was used for degassing, and the heat polymerization was at 70°C for around 2 hours in a heating incubator (DHP-9032, Prolab). The PDMS microdevices replicas were peeled off to perforate the inlets and outlets employing a 0.75 mm internal diameter retractable cutting cannula (Electron Microscopy Sciences). To complete the microfabrication process, microchannels were sealed on a microscope slide with corona treatment (BD-20AC, Electro-Technic Products).

The following design schematics illustrate the microdroplet generator used to form the droplets and the observation chamber microdevice employed for loading the microdroplets for fluorescent image acquisition. All dimensions are provided in millimeters.

Fig. S1. Schematic of the flow-focusing microdroplet generator device (a) with and (b) without dimensions. Detail A for the microdroplet generator device (c) with and (d) without dimensions.

Fig. S2. Schematic of the observation chamber microdevice (a) with and (b) without dimensions.

Surface treatment

Fig. S3 illustrates the importance of Novec EGC-1700 surface treatment, which is crucial for generating droplets using the flow-focusing method. This treatment creates a hydrophobic surface, essential for preventing droplet adhesion to the microchannel walls and facilitating smooth formation of smaller droplets. The enhanced hydrophobicity of the channels also minimizes droplet coalescence, ensuring uniform size and stability. Additionally, it reduces the adsorption of biomolecules, such as nucleic acids, onto the channel surface. These features are vital for achieving consistent and reproducible experimental results.

Fig. S3. Image of the formation and stability of flow-focusing water droplets using a PDMS/glass device and 3% w/w surfactant in HFE7500 oil (a) without and (b) with Novec EGC-1700 surface treatment. Q1=12.7 μL/min, Q2=25.3 μL/min (QT=38 μL/min).

Surfactant percentage

The addition of a low percentage of fluorinated surfactant to HFE7500 oil, at a minimum of 0.25% w/w, was crucial for generating water droplets. This reduced the interfacial tension between the aqueous and oil phases, facilitating the stable formation of smaller and more homogeneous droplets. Additionally, the surfactant prevents coalescence, ensuring that the droplets remain discrete and maintain their integrity throughout the experimental process (Fig. S4). To further enhance droplet stability during handling, a surfactant concentration of 3% w/w was selected.

Fig. S4. Images of droplet formation and stability using HFE7500 oil (a) without adding surfactant and (b) with 0.25% w/w surfactant. Q1=12.7 μL/min, Q2=25.3 μL/min, and QT=38 μL/min.

Droplet diameter size variation

The diameter of the generated droplets varied significantly depending on the height and width of the microchannel intersection, under consistent flow rate conditions (Q1=2.7 μL/min, Q2=5.3 μL/min, and QT=8 μL/min). For a height and width of 30 μm, droplets of around 50 μm in diameter were generated, whereas with a height of 100 μm and widths of 60 and 100 μm, droplets' diameters were 260 μm and 315 μm, respectively (Fig. S5). Additionally, the microchannel intersection size determined the maximum or minimum diameter of the droplets and the generation frequency when the QT and the Q2/Q1 ratio were modified. When the Q2/Q1 ratio increased while maintaining a fixed QT value of 38 μL/min, the droplet diameter and generation frequency decreased due to the higher oil inflow compared to water (Figs. S6(a-b)). Conversely, when QT was increased with a fixed Q2/Q1 ratio of 4, the droplet diameter decreased significantly until it no longer decreased further (Fig. S6(c)) due to the limitation imposed by the microchannel dimensions. Meanwhile, the generation frequency increased (Fig. S6(d)). As a result, both QT and Q2/Q1 ratio should be increased to form droplets with smaller diameters. In contrast, high QT values and low Q2/Q1 increase the droplet generation frequency per second.

Fig. S5. Images of water droplets formed with HFE7500 with 3% w/w of surfactant using three generation devices with different microchannel intersection widths and heights, identified as (a)–(c). Q1=2.7 μL/min, Q2=5.3 μL/min, and QT=8 μL/min.

Fig. S6. Graphs showing the modification of droplet diameter and generation frequency per second (a-b) while maintaining a fixed QT at 38 µL/min and varying the Q2/Q1 ratio between 2 and 10, and (c-d) while maintaining a fixed Q2/Q1 ratio at 4 and varying the QT value between 8 and 38 µL/min.

Master mix selection for RT-LAMP amplification

Fig. S7. Photograph of 2% agarose electrophoresis gel to evaluate the amplification of pDNA with RT-LAMP employing two master mix kits and ND primers. GelGreen® fluorescent dye concentration in the reaction was 2X. Lane (1) 5 µL of 1 kb plusTM DNA ladder, lane (2) NTC, lanes (3-5) pDNA concentration of 2.5x108 cp/µL, 2.5x106 cp/µL, and 2.5x104 cp/µL using SuperScript™ IV RT-LAMP Master Mix; lane (6) NTC, lanes (7-9) pDNA concentration of 2.5x108 cp/µL, 2.5x106 cp/µL, and 2.5x104 cp/µL employing WarmStart® Colorimetric LAMP 2X Master Mix.

Determination of the GelGreen® fluorescent dye concentration for RT-LAMP

Fig. S8. Photograph of 2% agarose electrophoresis gel to evaluate the amplification of pDNA with RT-LAMP under the addition of GelGreen® fluorescent dye employing SuperScript™ IV RT-LAMP Master Mix and ND primers. Lane (1) 5 µL of 1 kb plusTM DNA ladder, lane (2) NTC, lanes (3-5) pDNA concentration of 2.5x108 cp/µL, 2.5x106 cp/µL, and 2.5x104 cp/µL without GelGreen® addition; lane (6) NTC, lanes (7-9) pDNA concentration of 2.5x108 cp/µL, 2.5x106 cp/µL, and 2.5x104 cp/µL with the addition of GelGreen® 0.5X.

Fig. S9. Photograph of 2% agarose electrophoresis gel to evaluate the amplification time of pDNA with RT-LAMP by adding 0.5X GelGreen® fluorescent dye employing SuperScript™ IV RT-LAMP Master Mix and ND primers. Lane (1) 5 µL of 1 kb plusTM DNA ladder, lanes (2), (4), and (6) NTC at 8, 9, and 10 minutes of incubation time, respectively, lanes (3), (5), and (7) pDNA concentration of 2.5x104 cp/µL at 8, 9, and 10 minutes of incubation time, respectively.

Evaluation of primer sets for SARS-CoV-2 N gene amplification

Table S1. Different sets of six primers for gene N SARS-CoV-2 RT-LAMP amplification.

| Set name | Amplicon (B2-F2) | Primer | Sequence (5´- 3´) | Length (bp) |
| --- | --- | --- | --- | --- |
| NA | 170 bp | F3 | CCAGAATGGAGAACGCAGTG | 20 |
|  |  | B3 | CCGTCACCACCACGAATT | 18 |
|  |  | FIP | AGCGGTGAACCAAGACGCAGGGCGCGATCAAAACAACG | 38 |
|  |  | BIP | AATTCCCTCGAGGACAAGGCGAGCTCTTCGGTAGTAGCCAA | 41 |
|  |  | LF | TTATTGGGTAAACCTTGGGGC | 21 |
|  |  | LB | TTCCAATTAACACCAATAGCAGTCC | 25 |
| NB | 152 bp | F3 | AGATCACATTGGCACCCG | 18 |
|  |  | B3 | CCATTGCCAGCCATTCTAGC | 20 |
|  |  | FIP | TGCTCCCTTCTGCGTAGAAGCCAATGCTGCAATCGTGCTAC | 41 |
|  |  | BIP | GGCGGCAGTCAAGCCTCTTCCCTACTGCTGCCTGGAGTT | 39 |
|  |  | LF | GCAATGTTGTTCCTTGAGGAAGTT | 24 |
|  |  | LB | GTTCCTCATCACGTAGTCGCAACA | 24 |
| NC | 163 bp | F3 | TAGGAACTGGGCCAGAAGCTG | 21 |
|  |  | B3 | GAGGAACGAGAAGAGGCTTGAC | 22 |
|  |  | FIP | TGCCAATGTGATCTTGTGCTAACAAAGACGGCATCAT | 37 |
|  |  | BIP | AATCGTGCTACAACCCTCTGCTCCCTTCTGCGTA | 34 |
|  |  | LF | TATTCAAGGCTCCCTCAGTTGC | 22 |
|  |  | LB | TCAAGGAACAACATTGCCAAAAG | 23 |
| ND | 166 bp | F3 | CCCTCAGATTCAACTGGCAGTA | 22 |
|  |  | B3 | TTCATTTTACCGTCACCACCAC | 22 |
|  |  | FIP | GAGAGCGGTGAACGGGCGCGATCAAAACAACG | 32 |
|  |  | BIP | TCGAGGACAAGGCTTCGGTAGTAGCCAATTTGGTC | 35 |
|  |  | LF | GACGCAGTATTATTGGGTAAACCT | 24 |
|  |  | LB | CCAATTAACACCAATAGCAGTCCA | 24 |

Impact of incubation time and droplet size on ddRT-LAMP amplification

Fig. S10. ddRT-LAMP reaction with different incubation times at 65°C for a pDNA concentration of 2.5x106 cp/µL. Where %P is the percentage of positive droplets, N is the total number of droplets analyzed, and NP is the number of positive droplets.

Fig. S11. Images of ddRT-LAMP reactions using a pDNA concentration of 2.5x106 cp/µL with droplets of different diameters incubated at 65°C for 30 minutes. Where %P is the percentage of positive droplets, N is the total number of droplets analyzed, and NP is the number of positive droplets.

Intermediate precision and repeatability of ddRT-LAMP methodology

Fig. S12. Comparative images of ddRT-LAMP amplification with pDNA concentration of 2.5x105 cp/µL in duplicate over six days. R1: Replicate 1 and R2: Replicate 2.

Table S2. Percentage of positive droplets obtained after ddRT-LAMP amplification employing a pDNA concentration of 2.5x105 cp/µL in duplicate over six days.

| Day | %P1 | %P2 | %PID | SID%P | % CVID%P |
| --- | --- | --- | --- | --- | --- |
| 1 | 63.66 | 65.32 | 64.49 | 1.17 | 1.82 |
| 2 | 66.23 | 65.32 | 65.77 | 0.64 | 0.97 |
| 3 | 66.05 | 69.77 | 67.91 | 2.63 | 3.87 |
| 4 | 68.69 | 69.67 | 69.18 | 0.70 | 1.01 |
| 5 | 66.88 | 63.38 | 65.13 | 2.48 | 3.80 |
| 6 | 58.16 | 61.32 | 59.74 | 2.23 | 3.74 |
|  |  | %PIND | 65.37 |  |  |
|  |  | SIND%P | 3.28 |  |  |
|  |  | % CVIND%P | 5.00 |  |  |

Table S3. Results of the one-way analysis of variance (one-way ANOVA) for calculating the repeatability and intermediate precision of the ddRT-LAMP methodology.

| SUMMARY |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Groups | Account | Sum | Average | Variance |  |  |
| %P1 | 6 | 389.66 | 64.94 | 13.67 |  |  |
| %P2 | 6 | 394.77 | 65.80 | 11.43 |  |  |
| VARIANCE ANALYSIS |  |  |  |  |  |  |
| Origin of the variation | Sum of squares | Degrees of freedom | Mean square | F value | Probability | F-critical value |
| Inter-days | 2.18 | 1 | 2.18 | 0.17 | 0.69 | 4.96 |
| Intra-days | 125.50 | 10 | 12.55 |  |  |  |
| Total | 127.67 | 11 |  |  |  |  |

Determination of the dynamic range and limit of detection (LOD) using pDNA and SARS-CoV-2 synthetic RNA

Table S4. Results of positive droplet percentages for different concentrations of pDNA and synthetic RNA in duplicate with their respective statistical analysis.

| Template | Concentration (cp/µL) | %P1 | %P2 | %P | S%P | %CV%P |
| --- | --- | --- | --- | --- | --- | --- |
| pDNA | 2.5x102 | 0.68 | 0.42 | 0.55 | 0.18 | 33.43 |
|  | 2.5x103 | 1.75 | 1.37 | 1.56 | 0.27 | 17.22 |
|  | 2.5x104 | 7.17 | 8.86 | 8.02 | 1.20 | 14.91 |
|  | 2.5x105 | 61.22 | 66.05 | 63.64 | 3.42 | 5.37 |
| RNA | 1.6x102 | 0.42 | 0.66 | 0.54 | 0.17 | 31.43 |
|  | 1.6x103 | 2.21 | 1.66 | 1.94 | 0.39 | 20.10 |
|  | 1.6x104 | 12.69 | 12.55 | 12.62 | 0.10 | 0.78 |
|  | 1.6x105 | 74.26 | 73.40 | 73.83 | 0.61 | 0.82 |

ddPCR methodology

Reagents employed: 5x Taq Reaction Buffer (Accuris Instruments), 100 mM dNTPs (Accuris Instruments) diluted to 10 mM in nuclease-free water, 100 mM MgSO4 (New Englad BioLabs Inc.), 5 µM Primers F3 (CCCTCAGATTCAACTGGCAGTA) and B3 (TTCATTTTACCGTCACCACCAC ) for gene N amplification (T4OLIGO), Taq DNA Polymerase 5 000 U/mL (Accuris Instruments), GelGreen® nucleic acid stain 10 000X in water (Gold Biotehcnologies®) later diluted to 10X in free-nuclease water, 10% w/w NP-40 (Thermo Fisher Scientific) solution in nuclease-free water, mineral oil containing 2% w/w of span 80 (Sigma-Aldrich).  A volume of 12.5 µL of the PCR reaction mixture was prepared as follows: 3.4 µL of nuclease-free water, 2.5 µL of Buffer 5X, 0.25 µL of Mg2+ 100 mM, 0.25 µL of dNTPs 10 mM, 1.25 µL of primer F3 5 µM, 1.25 µL of primer B3 5 µM, 1.25 µL of GelGreen 10X, 0.25 µL of 10% w/w NP-40, 0.1 µL of Taq DNA polymerase 5 U/ µL, and 2 µL of pDNA or nuclease-free water depending on the preparation of positive or no-template controls. The PCR reaction mixture was encapsulated in droplets according to the established ddRT-LAMP with the flow rates of Q1 = 3.5 µL/min for the aqueous phase and Q2 = 34.5 µL/min for the 5% weight 008-FluoroSurfactant in HFE7500 oil (Q2/Q1 = 10). The droplets were generated with a diameter average of 93 ± 3.7 µm and collected on ice for three minutes. In this case, it was necessary to increase the percentage of surfactant from 3% to 5%, as well as add 0.2% of NP-40 solution to the reaction mixture and 70 µL of mineral oil containing 2% w/w of span 80 to keep the integrity of the droplets due to the high incubation time and temperature during ddPCR amplification.

The temperature program used for ddPCR is detailed below :

Fig. S13. Results obtained by applying the ddPCR methodology for the no-template control (NTC) and the positive pDNA controls with two different concentrations. Where %P is the percentage of positive droplets, N is the total number of droplets analyzed, and NP is the number of positive droplets.

Addition of 0.5% BSA Molecular Biology Grade

Fig. S14. ddRT-LAMP amplification results of synthetic RNA at 1.6x105 cp/µL (a) without BSA addition and (b) with 0.5 % BSA Molecular Biology Grade. Where %P is the percentage of positive droplets, N is the total number of droplets analyzed, and NP is the number of positive droplets.

Plasmid features

Fig. S15. pEZclone-NRS-Amp-Blunt-HC plasmid map with a 1260 bp sequence of gene N from SARS-CoV-2 and the cut site of the restriction enzyme XmnI for linearized the plasmid.  Gene N sequence from Wuhan-Hu-1: taatacgactcactatagggatgtctgataatggaccccaaaatcagcgaaatgcaccccgcattacgtttggtggaccctcagattcaactggcagtaaccagaatggagaacgcagtggggcgcgatcaaaacaacgtcggccccaaggtttacccaataatactgcgtcttggttcaccgctctcactcaacatggcaaggaagaccttaaattccctcgaggacaaggcgttccaattaacaccaatagcagtccagatgaccaaattggctactaccgaagagctaccagacgaattcgtggtggtgacggtaaaatgaaagatctcagtccaagatggtatttctactacctaggaactgggccagaagctggacttccctatggtgctaacaaagacggcatcatatgggttgcaactgagggagccttgaatacaccaaaagatcacattggcacccgcaatcctgctaacaatgctgcaatcgtgctacaacttcctcaaggaacaacattgccaaaaggcttctacgcagaagggagcagaggcggcagtcaagcctcttctcgttcctcatcacgtagtcgcaacagttcaagaaattcaactccaggcagcagtaggggaacttctcctgctagaatggctggcaatggcggtgatgctgctcttgctttgctgctgcttgacagattgaaccagcttgagagcaaaatgtctggtaaaggccaacaacaacaaggccaaactgtcactaagaaatctgctgctgaggcttctaagaagcctcggcaaaaacgtactgccactaaagcatacaatgtaacacaagctttcggcagacgtggtccagaacaaacccaaggaaattttggggaccaggaactaatcagacaaggaactgattacaaacattggccgcaaattgcacaatttgcccccagcgcttcagcgttcttcggaatgtcgcgcattggcatggaagtcacaccttcgggaacgtggttgacctacacaggtgccatcaaattggatgacaaagatccaaatttcaaagatcaagtcattttgctgaataagcatattgacgcatacaaaacattcccaccaacagagcctaaaaaggacaaaaagaagaaggctgatgaaactcaagccttaccgcagagacagaagaaacagcaaactgtgactcttcttcctgctgcagatttggatgatttctccaaacaattgcaacaatccatgagcagtgctgactcaactcaggcctaa

Generation of microfluidic droplets for ddRT-LAMP

Fig. S16. Example of filling 12.5 µL of aqueous phase (RT-LAMP reaction) into the syringe tubing connected to the microdevice inlet 2.